This journal © The Royal Society of Chemistry 2007

Electronic Supplementary Information

Table S1 MC-ICP-MS operating conditions for Pb analysis

- ICP conditions

Plasma power				1.2 kW

Torch 					Quartz glass torch with a sapphire injector

Plasma Ar gas flow rate		15 l min-1

Auxiliary Ar gas flow rate		0.80 l min-1

Nebulizer Ar gas flow rate		0.90 l min-1

2.  Interface

Sampling cone				Made of Ni

Skimmer cone				Made of Ni (X-skimmer)

3.  Data acquisition conditions

Washing time				960 sec after measurement

Uptake time				90 sec

Background data integration 		4 sec for 1 scan, 60 scans in one run

On-top zeroes

Sample data integration		4 sec for 1 scan, 20 scans in one run

4.  Desolvator conditions (Pb only)

Desolvator				ARIDUS (CETAC Technologies, USA)

Nebulizer				Micro-flow PFA nebulizer, PFA-20 (ESI, USA),

self-aspiration

Flow rate				60 l min-1 (observed)

Spray chamber temperature		110 ˚C

Desolvator temperature		160 ˚C

Sweep gas (Ar)			6.6 l min-1

Nitrogen gas 				4 ml min-1

5.  Cup configuration

L4    L3    L2   L1     C    H1   H2   H3

202  203  204  205  206  207  208  209