MACRO

H15 X

#___________________________________________________________

# AMC Software.  Minitab local macro

# Program to calculate Huber's H15 robust mean and standard deviation.

# Last modified:  05/02/01

# Input argument is a single column X, which is not changed.

# Executed in a Minitab session by using the command %H15 C

# where C is a column, assuming that the program is

# saved as a text file in the Minitab subdirectory <MACROS> under the

# filename <H15.MAC>.

# If the macro file is saved elsewhere, it must be called by using the session

# command %<pathname><filename>.

# The program will run under versions of Minitab from 10Xtra onwards.

# The robust technique is designed for medium sized data sets:

# Results from small sets should be treated with caution.

# Misleading results may be obtained if the program is used on data sets

# that are multimodal or strongly skewed or almost invariant.

# The algorithm is based on Analyst, 1989, 114, 1693.

# See also AMC Technical Brief no 6 on www.rsc.org/lap/rsccom/amc/amc_index.htm

#__________________________________________________________

#i/o arguments

MCOLUMN X 		#column of input data

#other declarations

MCONSTANTS M S 	#robust mean and sd

MCOLUMNS DIFF Y	#scratch columns

MCONSTANT SA 		#scratch constant

MCONSTANT MAD 	#median absolute difference

MCONSTANT TOL 		#terminates iteration if < 0.00001

MCONSTANT HIGH LOW  #limits for winsorisation

MCONSTANT KOUNT 	#loop counter

MCONSTANT MAX MIN 	#highest and lowest data values

MCONSTANT COLUMN	#column name

MCONSTANT ROBMEAN 	#output name

MCONSTANT ROBSD	#ditto

#_____________________________________________________________

# set up starting values

IF N(X)< 6

NOTE DATA SET TOO SMALL

LET M = '*'

LET S = '*'

EXIT

ELSE

LET M = MEDIAN(X)

LET DIFF = X - M

LET MAD= MEDIAN(ABSOLUTE(DIFF))

LET S = MAD/0.6745

ENDIF

# initialise variables for iteration

LET TOL = 0.1

LET KOUNT=0

LET MAX=MAXIMUM(X)*1.1

LET MIN=MINIMUM(X)*0.9

# iteration

WHILE (TOL > 0.00001 AND KOUNT<100) OR KOUNT<10

COPY X Y

LET SA = S

LET LOW = M-1.5*S

LET HIGH = M + 1.5*S

#winsorise raw data and update estimates

CODE (MIN:LOW) LOW  (HIGH:MAX) HIGH Y Y

LET M = MEAN(Y)

LET S = STDEV(Y)/0.882

#check for sufficient convergence

IF SA > 0

LET TOL = ABSOLUTE(SA-S)/SA

ELSE

NOTE PROBLEM WITH NEAR ZERO SD - TERMINATED

LET M = '*'

LET S = '*'

EXIT

ENDIF

LET KOUNT=KOUNT+1

ENDWHILE

#output

NOTE

NOTE ROBUST MEAN AND STANDARD DEVIATION

NOTE

NOTITLE

LET ROBMEAN = M

LET ROBSD = S

KKNAME COLUMN X

PRINT COLUMN ROBMEAN ROBSD

NOTE

TITLE

IF  KOUNT > 99

NOTE WARNING - ITERATION NOT CONVERGED

ENDIF

ENDMACRO

#____________________________________________________________