## Summary
Zr iRM calibration: Data and date treatment | [32 empty] | OPZ: User box | [14 empty] | Interference correction: user boxes | [24 empty] | Mass fractionation laws | [10 empty] | Mass-bias correction: user box |  |  |  | Width of data array | [15 empty] | FIGURES and TABLES | [63 empty] | Absolute ratios | [6 empty] | Mass-fractionation exponents (θ) | [8 empty] | x | -10 rel. Exp | -1 rel. Exp |  | +1 rel. Exp | +3 rel. Exp | +10 rel. Exp | x noise | Noise
The first 9 lines are for constants and user input on decisions such OPZ corrections, interference corrections, and mass fractionation laws. | [32 empty] | OPZ | (y=1,n=0) | [13 empty] | Lab 1: 95Mo correction |  |  | 0 | Isotope | 90Zr | 91Zr | 92Zr | 94Zr | 96Zr | 91/90 | 92/90 | 94/90 | 96/90 | [11 empty] | Exponential | [10 empty] | Lab | n for GPL | Ln92-91 | Ln96-91 | Ln96-92 | Smallest | Most similar | [77 empty] | Min | 0.21793 | 0.33334 | 0.05437 |  |  |  | n | -10 | -1 | 0 | 1 | 3 | 10 |  | Ln92 vs 91 | -1.5233 | -1.0983 | -1.0982 |  | -1.0982 | -1.0982 | -1.0982 | -1.5235 | -1.0985
Anything from line 10 and below is automated calculations. Do not modify. | [32 empty] | L1 | 1 | [13 empty] | Lab 1: 98Mo correction |  |  | 1 | Nat Ab | 51.446 | 11.2116 | 17.1512 | 17.3939 | 2.7972 | 0.21793 | 0.33338 | 0.3381 | 0.054372 | [11 empty] | Rmeas=Rtrue*(Mi/Mj)^beta | [10 empty] | L1 | 0 | 102.2522 | 161.2266 | 142.4159 | 1.21 | 1.21 | [77 empty] | Max | 0.21797 | 0.33338 | 0.054392 |  |  |  | 91/90 (4/0) | 0.29687 | 0.25839 | 0.25425 | 0.25013 | 0.242 | 0.21455 |  |  | -1.5236 | -1.0986 | -1.0986 | -1.0986 | -1.0986 | -1.0986 | -1.0986 | -1.5236 | -1.0986
[33 empty] | L2 | 1 | [13 empty] | Lab 2: 95Mo correction |  |  | 0 | Amu | 89.9047 | 90.9056 | 91.905 | 93.9063 | 95.9083 | [15 empty] | GPL | [10 empty] | L2 S1 | 0 | 46.1437 | 101.2102 | 98.8327 | 0.8 | 1.4 | [77 empty] | Range (ppm) | 197.8632 | 116.2207 | 416.3912 |  |  |  | 92/90 (4/0) | 0.5595 | 0.51077 | 0.50533 | 0.49988 | 0.489 | 0.45107 |  |  | -1.5237 | -1.0987 | -1.0987 |  | -1.0987 | -1.0987 | -1.0987 | -1.5236 | -1.0986
[33 empty] | L3 | 1 | [13 empty] | Lab 2: 98Mo correction |  |  | 0 | Isotope | 92Mo | 94Mo | 95Mo | 96Mo | 97Mo | 98Mo | 100Mo | 96Ru | 99Ru | 100Ru | 101Ru | [9 empty] | Rmeas=Rtrue*g^(Mi^n-Mj^n) | [10 empty] | L2 S2 | 0 | 24.1401 | 61.6522 | 61.2658 | -1.35 | 1.45 | [84 empty] | 96/90 (4/0) | 1.3485 | 1.469 | 1.4844 | 1.5003 | 1.5334 | 1.6652 |  | Ln96 vs 91 | -1.5233 | -2.9123 | -2.9124 |  | -2.9125 | -2.9125 | -2.9126 | -1.5235 | -2.9115
[48 empty] | Lab 2: 100Mo correction |  |  | 1 | Nat Ab | 14.649 | 9.187 | 15.873 | 16.673 | 9.582 | 24.292 | 9.744 | 5.52 | 12.7 | 12.6 | 17 | [20 empty] | L3 | 0 | 119.5419 | 279.0592 | 264.4988 | 0.56 | 0.78 | [44 empty] | Range extrenal reproducibility | [39 empty] | Slope triple iso log-log, when data fractionated by law with exponent k is corrected by law with exponent n. | [8 empty] | -1.5235 | -2.9117 | -2.9117 | -2.9117 | -2.9117 | -2.9117 | -2.9117 | -1.5236 | -2.9116
[48 empty] | Ru correction |  |  | 1 | Amu | 91.9068 | 93.9051 | 94.9058 | 95.9047 | 96.906 | 97.9054 | 99.9075 | 95.9076 | 98.9059 | 99.9042 | 100.9056 | [20 empty] | Norm 94/90Zr | 0.3381 | [49 empty] | Min | 6.3921 | 2.4244 | 11.5762 | [14 empty] | For plotting | [22 empty] | -10 rel. Exp | -1 rel. Exp |  | +1 rel. Exp | +1 rel. Exp | +10 rel. Exp | Noise - exp |  | -1.5237 | -2.9113 | -2.9112 |  | -2.9112 | -2.9112 | -2.9111 | -1.5236 | -2.9116
[48 empty] | beta corr. Interferences |  |  | 1 | [32 empty] | (Minster & Ricard 1981) | [50 empty] | Max | 21.9135 | 22.8406 | 58.9569 | [14 empty] | Starting | 0.8 | [20 empty] | Ln92 vs 91 | 1.2709 | 1.3136 |  | 1.3232 | 1.3328 | 1.3668 | 0.58629 | Ln96 vs 92 | -1.0986 | -2.9113 | -2.9113 |  | -2.9113 | -2.9113 | -2.9112 | -1.0986 | -2.9116
[48 empty] | Beta Mo /Beta Zr |  |  | 0.9 | [52 empty] | Fig 1 | Table 1. Community identified guidelines for selection of an iRM. | Table 2. Summary of the analytical setup and conditions used for Zr isotope analyses in each laboratory. |  |  |  | Fig 2 (Data) | [24 empty] | Avg | 12.1119 | 9.6517 | 31.3484 | [14 empty] | Offset 1 | 0.4 | Fig 3 | Table 4 and Figure 3 below | [10 empty] | Figure 4. Internally normalized ratios vs beta exponential law | [7 empty] | Ln96 vs 91 | -3.188 | -3.7254 |  | -3.8609 | -4.0028 | -4.5547 | 1.274 |  | -1.0985 | -2.9117 | -2.9117 | -2.9117 | -2.9117 | -2.9117 | -2.9117 | -1.0986 | -2.9116
Lab | Date | Seq | Meas. | Sample | OPZ | [13 empty] | SAMPLE | [13 empty] | OPZ corrected intensities (y=1, n=0) | [14 empty] | Interference correction (y=1, n=0) | [6 empty] | No beta corr formulaes | 6.7746 | 7.1362 | 1.1912 |  |  |  |  | To check for correlations between ratios not sharing any isotopes | [5 empty] | Raw | [12 empty] | Mass fractionation corrected ratios (Exp for n=0, GPL otherwise) | [6 empty] | To check for correlations between ratios not sharing any isotopes | [8 empty] | To check for correlations between ratios not sharing any isotopes | [5 empty] | Interlaboratory calibration protocol | 1) Must be demonstrably homogeneous. | Parameter | Lab 1 (Caltech) | Lab 2 (ETH) | Lab 3 (UCLA) | Sample-Std bracketing after internal normalization | [15 empty] | Arrange by sample | [14 empty] | Table 3. Differences in Zr isotope composition of the various RM 8299 aliquots and SRM 3169 lots. | [11 empty] | Offset 2 | 1.2 | Sample-Std bracketing after internal normalization | Comparison of absolute ratios | [10 empty] | Typical 2σ | 91/90Zr | 92/90Zr | 96/90Zr | Ln-Ln | 91/90Zr | 92/90Zr | 96/90Zr | Ln96 vs 92 | -2.5085 | -2.8361 |  | -2.9179 | -3.0032 | -3.3324 | 0.6325 |  | -1.0984 | -2.912 | -2.912 |  | -2.912 | -2.912 | -2.9121 | -1.0985 | -2.9116
|  | # | # |  | 177Hf++ | 49Ti40Ar | 90Zr | 91Zr | 92Zr | 93Nb | 94Zr | 95Mo | 96Zr | 97Mo | 98Mo | 99Ru | 100Ru | 101Ru | 177Hf++ | 49Ti40Ar | 90Zr | 91Zr | 92Zr | 93Nb | 94Zr | 95Mo | 96Zr | 97Mo | 98Mo | 99Ru | 100Ru | 101Ru | OPZ | 177Hf++ | 49Ti40Ar | 90Zr | 91Zr | 92Zr | 93Nb | 94Zr | 95Mo | 96Zr | 97Mo | 98Mo | 99Ru | 100Ru | 101Ru | 95Mo | 98 or 100 Mo | 100Mo | Ru | beta corr? | Beta | 90Zr | 91Zr | 92Zr | 94Zr | 96Zr | 91/90Zr | 92/90Zr | 94/90Zr | 96/90Zr | 92/91Zr | 94/91Zr | 96/91Zr | 94/92Zr | 96/92Zr | 96/94Zr | Ln(91/90) | Ln(92/90) | Ln(94/90) | Ln(96/90) | [9 empty] | Labels | n GPL | Beta (or g) | 91/90Zr | 92/90Zr | 94/90Zr | 96/90Zr | 92/91Zr | 94/91Zr | 96/91Zr | 94/92Zr | 96/92Zr | 96/94Zr | Ln(91/90Zr) | Ln(92/90Zr) | Ln(96/90Zr) | Ln(92/91Zr) | Ln(94/91Zr) | Ln(96/91Zr) | Ln(94/92Zr) | Ln(96/92Zr) | Ln(96/94Zr) |  | 2) Must be pure elements/chemical compounds easily dissolved into diluted acids. | Mass spectrometer | Neptune Plus MC-ICPMS | Neptune Plus MC-ICPMS | Neptune MC-ICPMS | STD | Lab | Session | Seq | Conc match | u91 | u92 | u96 | Sample | Lab | Session | Seq | Conc match | u91 | u92 | u96 | Sample | Lab | Session | Seq | Conc. Match | u91 | u92 | u96 | Rel. to |  |  |  |  | Lab | Session | Sample | Lot # | Session | Seq. | n | μ91Zr | 2SE | μ92Zr | 2SE | μ96Zr | 2SE | Rel. to | y-axis |  |  |  | 91/90Zr | 2SD | 92/90Zr | 2SD | 94/90Zr | 2SD | 96/90Zr | 2SD | x value | Ref. | ppm | 12 | 12 | 30
Lab 1 | Apr 4, 2022 | 1 | Zr_005 | iRM |  |  | 0.00023002 | 5.6712e-05 | 7.4629e-05 | 3.8819e-06 | 7.2605e-05 | -2.5875e-07 | 9.0466e-05 | -6.3369e-06 | 4.1976e-06 | [5 empty] | 19.5485 | 4.3444 | 6.7747 | 8.1816e-05 | 7.1364 | 0.00016165 | 1.1915 | 7.2234e-05 | 0.00018838 |  |  |  | 1 | 0 | 0 | 19.5483 | 4.3443 | 6.7747 | 7.7934e-05 | 7.1363 | 0.00016191 | 1.1914 | 7.8571e-05 | 0.00018418 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7617 | 19.5483 | 4.3443 | 6.7746 | 7.1362 | 1.1913 | 0.22224 | 0.34656 | 0.36506 | 0.060939 | 1.5594 | 1.6427 | 0.27421 | 1.0534 | 0.17584 | 0.16693 | -1.504 | -1.0597 | -1.0077 | -2.7979 | [9 empty] | L1 AprS1 | 0 | 1.7615 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014064 | -1.8133 | -1.8273 |  | 3) Isotopic composition within the range of natural variability (ideally representative of a major geological reservoir). | Sample introduction | Wet plasma (glass spray chamber) | Dry plasma (Aridus II) | Wet plasma (glass spray chamber) | [16 empty] | iRM | Lab 1 | Apr 4, 2022 | 1 | 98.51% | 8.5822 | -5.651 | -13.245 | iRM | Avg | -0.04467 | -2.3401 | -0.30986 |  |  | Laboratory 1 | [11 empty] | all | wtd avg |  | N+83 | 0.21819 | 0.00022 | 0.33339 | 0.00013 | 0.3381 |  | 0.054474 | 9e-05 | 0.26 | Nomura et al (1983) | Abs | 2.6154e-06 | 4.0004e-06 | 1.6314e-06 |  | 1.2e-05 | 1.2e-05 | 3e-05
Lab 1 | Apr 4, 2022 | 1 | Zr_010 | iRM_1 |  |  | 0.00040412 | 8.3994e-05 | 0.00014014 | 1.4003e-05 | 0.00011701 | 9.7932e-06 | 8.9972e-05 | -1.4306e-06 | 4.7201e-06 | [5 empty] | 18.8222 | 4.1829 | 6.5228 | 7.2572e-05 | 6.871 | 4.5899e-05 | 1.147 | -8.3632e-08 | 8.6134e-06 |  |  |  | 1 | 0 | 0 | 18.8218 | 4.1829 | 6.5227 | 5.8569e-05 | 6.8709 | 3.6106e-05 | 1.1469 | 1.347e-06 | 3.8933e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.761 | 18.8218 | 4.1829 | 6.5227 | 6.8709 | 1.1469 | 0.22223 | 0.34655 | 0.36505 | 0.060936 | 1.5594 | 1.6426 | 0.2742 | 1.0534 | 0.17584 | 0.16693 | -1.504 | -1.0597 | -1.0077 | -2.7979 | Calculate GPL exponent based on raw ratios | [7 empty] | Slope noise |  | 0 | 1.761 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014073 | -1.8133 | -1.8274 |  | 4) Must be free of isotopic anomalies. | Nebulizer type / flow | PFA nebulizer, 50 μl/min | PFA nebulizer, 50 μl/min | PFA nebulizer, 50 μl/min | [8 empty] | iRM_1 | Lab 1 | Apr 4, 2022 | 1 | 97.45% | -3.6088 | -4.9876 | -15.8405 | iRM | Lab 1 | Apr 4, 2022 | 1 | 100% | -4.1295 | 4.3402 | 47.7162 | iRM | 2SD | 11.2446 | 7.7542 | 47.0646 | Lab 1 | Apr 4, 2022 | Bottle 1 |  | Apr | 1 | 9 | -5.4788 | 3.7323 | -0.3949 | 3.3436 | 5.6246 | 9.1627 | Bottle 12 | 0.8 |  |  | M&R81 | 0.21799 | 5e-05 | 0.33338 | 6e-05 | 0.3381 |  | 0.05439 | 7e-05 | 0.12 | Minster & Ricard (1981) | [8 empty] | Figure 5. Log-log plots of internally normalized ratios
Lab 1 | Apr 4, 2022 | 1 | Zr_015 | iRM |  |  | 0.00037328 | 8.6251e-05 | 0.00012659 | 5.0764e-06 | 0.00010035 | -6.2877e-07 | 8.0107e-05 | -9.0873e-06 | 4.4407e-06 | [5 empty] | 19.0788 | 4.24 | 6.6119 | 7.7438e-05 | 6.9649 | 0.00015956 | 1.1628 | 7.4882e-05 | 0.00018351 |  |  |  | 1 | 0 | 0 | 19.0785 | 4.2399 | 6.6118 | 7.2361e-05 | 6.9648 | 0.00016019 | 1.1627 | 8.3969e-05 | 0.00017907 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7617 | 19.0785 | 4.2399 | 6.6117 | 6.9647 | 1.1626 | 0.22224 | 0.34655 | 0.36506 | 0.060939 | 1.5594 | 1.6426 | 0.27421 | 1.0534 | 0.17584 | 0.16693 | -1.504 | -1.0597 | -1.0077 | -2.7979 |  | (y vs x) | Lin-Lin slope | Noise slope | Ln-Ln slope | GPL slope | n for GPL | Ln-Ln noise | after B corr |  | 0 | 1.7615 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42503 | 0.4391 | -1.3882 | 0.014072 | -1.8133 | -1.8273 |  | 5) Community-led effort to choose the iRM against which to report data. | Cones | General H sampler and H skimmer | General H sampler and H skimmer | General H sampler and H skimmer | iRM | Lab 1 | Apr 4, 2022 | 1 | 98.51% | 8.5822 | -5.651 | -13.245 | [8 empty] | iRM | Lab 1 | Apr 4, 2022 | 1 | 99.96% | -2.5573 | -3.384 | -45.2331 | iRM | n | 12 | 12 | 12 | Lab 1 | Jun 20, 2022 | Bottle 1 |  | June | 1 | 10 | 0.77891 | 3.879 | -2.1173 | 2.2172 | -0.95373 | 8.0748 | Bottle 12 | 0.8 |  |  | S+04 | 0.21793 | 2.3e-05 | 0.33338 | 2.5e-05 | 0.3381 |  | 0.054371 | 8e-06 | 0.4 | Schonbachler et al (2004)
Lab 1 | Apr 4, 2022 | 1 | Zr_020 | iRM_2 |  |  | 0.00036594 | 8.5974e-05 | 0.00012587 | 7.5588e-06 | 0.0001057 | -7.1812e-07 | 8.3511e-05 | -7.7917e-06 | 7.0114e-06 | [5 empty] | 18.8949 | 4.199 | 6.5478 | 7.2362e-05 | 6.8971 | 4.5304e-05 | 1.1514 | 6.7295e-06 | 4.892e-06 |  |  |  | 1 | 0 | 0 | 18.8946 | 4.1989 | 6.5477 | 6.4803e-05 | 6.897 | 4.6022e-05 | 1.1513 | 1.4521e-05 | -2.1194e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7596 | 18.8946 | 4.1989 | 6.5477 | 6.897 | 1.1513 | 0.22223 | 0.34654 | 0.36503 | 0.060932 | 1.5594 | 1.6426 | 0.27419 | 1.0534 | 0.17583 | 0.16693 | -1.504 | -1.0598 | -1.0078 | -2.798 | Lab 1 | 92 vs 91 | 3.1853 | 1.3107 | 2.0426 | 2.0434 | 5.0037 | 0.84054 | 1.2021 |  | 0 | 1.7596 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43912 | -1.3882 | 0.014072 | -1.8133 | -1.8273 |  | 6) Must be widely available (i.e., conflict-of-interest free distribution). | Resolution | Low resolution | Low resolution | Low resolution | [8 empty] | iRM_2 | Lab 1 | Apr 4, 2022 | 1 | 98.76% | -8.4493 | -1.2719 | -7.3659 | iRM | Lab 1 | Apr 4, 2022 | 1 | 100.1% | -5.8138 | 0.97256 | 4.5991 | iRM | 2SE | 3.246 | 2.2384 | 13.5864 |  |  | Bottle 1 avg |  |  |  |  | -2.4705 | 2.6895 | -1.5912 | 1.8478 | 1.9219 | 6.0581 |  | 0.8 | 0.8 |  | Q+15 | 0.21797 | 0.00045 | 0.33342 | 0.00079 | 0.3381 |  | 0.054341 | 0.0005 | 0.54 | Quemet et al. (2015)
Lab 1 | Apr 4, 2022 | 1 | Zr_025 | iRM |  |  | 0.00040394 | 9.6806e-05 | 0.00013231 | 1.6281e-05 | 0.00012717 | -2.0454e-06 | 8.1198e-05 | -7.5456e-06 | 4.4056e-06 | [5 empty] | 19.1864 | 4.2638 | 6.649 | 7.3816e-05 | 7.0037 | 0.0001608 | 1.1693 | 6.6459e-05 | 0.00018685 |  |  |  | 1 | 0 | 0 | 19.186 | 4.2637 | 6.6488 | 5.7535e-05 | 7.0036 | 0.00016285 | 1.1692 | 7.4005e-05 | 0.00018245 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7602 | 19.186 | 4.2637 | 6.6487 | 7.0035 | 1.1691 | 0.22223 | 0.34654 | 0.36503 | 0.060935 | 1.5594 | 1.6426 | 0.2742 | 1.0534 | 0.17584 | 0.16693 | -1.504 | -1.0598 | -1.0078 | -2.7979 |  | 94 vs 91 | 6.7103 | 1.3544 | 4.0852 | 4.0857 | 2.3181 | 0.82459 | 3.2606 |  | 0 | 1.76 | 0.21794 | 0.33338 | 0.3381 | 0.054382 | 1.5297 | 1.5513 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43911 | -1.3882 | 0.01407 | -1.8132 | -1.8273 |  | 7) Stock must be stable and sufficient to last decades. | Cup Configuration | L4 / L3 / L2 / L1 / C / H1 / H2 / H3 / H4 | L4 / L3 / L2 / L1 / C / H1 / H2 / H3 / H4 | L3 / L2 / L1 / C / H1 / H2 / H4 | iRM | Lab 1 | Apr 4, 2022 | 1 | 100% | -4.1295 | 4.3402 | 47.7162 | [8 empty] | iRM | Lab 1 | Apr 4, 2022 | 1 | 100.1% | 8.54 | -0.99325 | 14.7092 | iRM |  |  |  |  | Lab 1 | Apr 4, 2022 | Bottle 2 |  | Apr | 1 | 8 | -0.94856 | 5.161 | -2.3907 | 3.3646 | -12.749 | 7.6063 | Bottle 12 | 2.4 |  |  | ZIRC-1 norm | 0.21792 | 1.4323e-05 | 0.33332 | 1.5505e-05 | 0.3381 |  | 0.054375 | 6.4009e-06 | 0.68 | He et al (2021)
Lab 1 | Apr 4, 2022 | 1 | Zr_030 | iRM_10 |  |  | 0.0003689 | 8.262e-05 | 0.00011468 | 7.5108e-06 | 0.00010213 | -4.3971e-06 | 8.4972e-05 | -1.0308e-05 | 6.3218e-06 | [5 empty] | 18.9221 | 4.205 | 6.557 | 7.4246e-05 | 6.9066 | 4.5745e-05 | 1.1529 | -2.7107e-06 | 3.5951e-06 |  |  |  | 1 | 0 | 0 | 18.9217 | 4.2049 | 6.5569 | 6.6735e-05 | 6.9065 | 5.0142e-05 | 1.1528 | 7.597e-06 | -2.7268e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7581 | 18.9217 | 4.2049 | 6.5569 | 6.9065 | 1.1528 | 0.22223 | 0.34653 | 0.365 | 0.060926 | 1.5593 | 1.6425 | 0.27416 | 1.0533 | 0.17582 | 0.16692 | -1.5041 | -1.0598 | -1.0079 | -2.7981 |  | 96 vs 91 | 1.6781 | 0.48784 | 6.12 | 6.1255 | 1.7722 | 1.7793 | 4.3407 |  | 0 | 1.7581 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43911 | -1.3882 | 0.01407 | -1.8133 | -1.8273 |  | Recommendations from Teng et al. (2017) | Isotopes monitored | 90Zr, 91Zr, 92Zr, 93Nb, 94Zr, 95Mo, 96Zr, 97Mo, 98Mo | 90Zr, 91Zr, 92Zr, 94Zr, 95Mo, 96Zr, 99Ru, 100Mo/Ru, 101Ru | 177Hf++, 49Ti40Ar, 90Zr, 91Zr, 92Zr, 94Zr, 96Zr | [8 empty] | iRM_10 | Lab 1 | Apr 4, 2022 | 1 | 98.37% | 10.8212 | 2.7856 | 11.362 | iRM | Lab 1 | Apr 5, 2022 | 2 | 100.4% | -4.7411 | -8.7339 | 15.3966 | iRM_12 |  |  |  |  | Lab 1 | Jun 20, 2022 | Bottle 2 |  | June | 1 | 9 | -1.1346 | 2.7875 | -0.91402 | 2.0049 | -0.32674 | 5.7019 | Bottle 12 | 2.4 |  |  | Lab 1 | 0.21794 | 8.5216e-06 | 0.33338 | 7.4263e-06 | 0.3381 |  | 0.05438 | 3.2228e-06 | 0.86 | This work
Lab 1 | Apr 4, 2022 | 1 | Zr_035 | iRM |  |  | 0.00040911 | 8.2403e-05 | 0.00013216 | 1.1854e-05 | 0.00012957 | 1.31e-06 | 9.8574e-05 | 3.4476e-07 | 5.1537e-06 | [5 empty] | 19.284 | 4.2853 | 6.6823 | 7.7667e-05 | 7.0385 | 0.00016051 | 1.175 | 7.1313e-05 | 0.00018888 |  |  |  | 1 | 0 | 0 | 19.2836 | 4.2853 | 6.6822 | 6.5813e-05 | 7.0383 | 0.0001592 | 1.1749 | 7.0969e-05 | 0.00018373 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7574 | 19.2836 | 4.2853 | 6.6821 | 7.0383 | 1.1748 | 0.22222 | 0.34652 | 0.36499 | 0.06092 | 1.5593 | 1.6424 | 0.27414 | 1.0533 | 0.17581 | 0.16691 | -1.5041 | -1.0598 | -1.0079 | -2.7982 |  | 94 vs 92 | 2.1067 | 1.0333 | 2 | 1.9995 | 0.9534 | 0.98102 | 1.019 |  | 0 | 1.7572 | 0.21794 | 0.33337 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014076 | -1.8133 | -1.8274 |  |  | Amplifiers | 1012 Ω on H1 & H4; 1011 Ω on all other cups | 1012 Ω on H2 & H4; 1011 Ω on all other cups | 1011 Ω on all cups | iRM | Lab 1 | Apr 4, 2022 | 1 | 99.96% | -2.5573 | -3.384 | -45.2331 | [8 empty] | iRM | Lab 1 | Apr 5, 2022 | 2 | 99.17% | -1.513 | -6.4571 | -8.0459 | iRM_12 | [6 empty] | Bottle 2 avg |  |  |  |  | -1.0925 | 2.4526 | -1.301 | 1.7223 | -4.796 | 4.5623 |  | 2.4 | 2.4 |  | Lab 2 | 0.21793 | 2.9873e-06 | 0.33335 | 5.6492e-06 | 0.3381 |  | 0.05439 | 1.8296e-06 | 0.96 | This work
Lab 1 | Apr 4, 2022 | 1 | Zr_040 | iRM_11 |  |  | 0.00039854 | 8.2016e-05 | 0.00012723 | 7.8169e-06 | 0.0001249 | 2.3882e-06 | 9.9333e-05 | -1.4291e-05 | 5.5183e-06 | [5 empty] | 19.0779 | 4.2396 | 6.6108 | 6.665e-05 | 6.9632 | 4.4724e-05 | 1.1624 | 3.529e-06 | 3.6551e-06 |  |  |  | 1 | 0 | 0 | 19.0775 | 4.2395 | 6.6107 | 5.8834e-05 | 6.9631 | 4.2336e-05 | 1.1623 | 1.782e-05 | -1.8632e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7575 | 19.0775 | 4.2395 | 6.6107 | 6.9631 | 1.1623 | 0.22222 | 0.34652 | 0.36499 | 0.060924 | 1.5593 | 1.6424 | 0.27415 | 1.0533 | 0.17582 | 0.16692 | -1.5041 | -1.0598 | -1.0079 | -2.7981 |  | 96 vs 92 | 0.52684 | 0.37219 | 2.9963 | 2.9978 | 0.94616 | 2.1169 | 0.87936 |  | 0 | 1.7575 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43911 | -1.3882 | 0.014077 | -1.8132 | -1.8273 |  |  | Gain calibration | Daily | Daily | Weekly | [8 empty] | iRM_11 | Lab 1 | Apr 4, 2022 | 1 | 98.64% | 5.6088 | -1.2676 | 27.1619 | iRM | Lab 1 | Apr 5, 2022 | 3 | 98.83% | 4.6161 | -4.9887 | 2.5602 | iRM_12 |  |  |  |  | Lab 1 | Apr 4, 2022 | Bottle 10 |  | Apr | 1 | 8 | -0.31231 | 5.0476 | -2.8536 | 3.2369 | -3.143 | 12.6778 | Bottle 12 | 10 |  |  | Lab 3 | 0.21796 | 9.8616e-06 | 0.33336 | 8.01e-06 | 0.3381 |  | 0.054377 | 7.5927e-06 | 1.06 | This work
Lab 1 | Apr 4, 2022 | 1 | Zr_045 | iRM |  |  | 0.00044124 | 0.00010106 | 0.00015172 | 1.1614e-05 | 0.00014035 | 5.2046e-08 | 9.1135e-05 | 7.9156e-07 | 3.9598e-06 | [5 empty] | 19.3968 | 4.3103 | 6.7211 | 7.4178e-05 | 7.0789 | 0.00016132 | 1.1817 | 7.6098e-05 | 0.00019257 |  |  |  | 1 | 0 | 0 | 19.3963 | 4.3102 | 6.721 | 6.2564e-05 | 7.0788 | 0.00016127 | 1.1816 | 7.5307e-05 | 0.00018861 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7552 | 19.3963 | 4.3102 | 6.7209 | 7.0787 | 1.1815 | 0.22222 | 0.3465 | 0.36495 | 0.060913 | 1.5593 | 1.6423 | 0.27412 | 1.0533 | 0.17579 | 0.16691 | -1.5041 | -1.0599 | -1.008 | -2.7983 |  | 96 vs 94 | 0.25008 | 0.36019 | 1.4981 | 1.4993 | 0.93935 | 2.1579 | -0.65971 |  | 0 | 1.755 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8273 |  |  | Acid matrix | 0.58 M HNO3 + 0.01 M HF | 0.5 M HNO3 + 0.005 M HF | 0.3 M HNO3 + 0.005 M HF | iRM | Lab 1 | Apr 4, 2022 | 1 | 100.1% | -5.8138 | 0.97256 | 4.5991 | [8 empty] | iRM | Lab 1 | Apr 5, 2022 | 3 | 99.87% | 4.1551 | -2.6105 | 1.9655 | iRM_12 |  |  |  |  | Lab 1 | Jun 20, 2022 | Bottle 10 |  | June | 1 | 9 | -0.92231 | 3.4539 | -1.0418 | 4.1196 | 2.1531 | 11.6946 | Bottle 12 | 10
Lab 1 | Apr 4, 2022 | 1 | Zr_050 | iRM_12 |  |  | 0.00042458 | 0.00010266 | 0.00014309 | 8.7026e-06 | 0.00013091 | 1.1186e-05 | 9.3021e-05 | -3.176e-06 | 7.7575e-06 | [5 empty] | 19.2318 | 4.2736 | 6.6637 | 7.2221e-05 | 7.0184 | 5.8047e-05 | 1.1715 | 3.2075e-06 | 4.8403e-06 |  |  |  | 1 | 0 | 0 | 19.2313 | 4.2735 | 6.6636 | 6.3518e-05 | 7.0183 | 4.6861e-05 | 1.1714 | 6.3835e-06 | -2.9172e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7543 | 19.2313 | 4.2735 | 6.6636 | 7.0183 | 1.1714 | 0.22221 | 0.3465 | 0.36494 | 0.060911 | 1.5593 | 1.6423 | 0.27411 | 1.0532 | 0.17579 | 0.16691 | -1.5041 | -1.0599 | -1.008 | -2.7983 |  |  | Correlation not from CS. |  |  | Best estimate for n: | 0.9463 |  |  |  | 0 | 1.7543 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43912 | -1.3882 | 0.014077 | -1.8132 | -1.8273 |  |  | Solution [Zr] | 1 μg/g | 350 ng/g | 2 μg/g | [8 empty] | iRM_12 | Lab 1 | Apr 4, 2022 | 1 | 98.91% | -9.2161 | -0.63121 | 8.0797 | iRM | Lab 1 | Apr 5, 2022 | 3 | 101% | 1.9212 | -1.8048 | -26.3623 | iRM_12 | [6 empty] | Bottle 10 avg |  |  |  |  | -0.72778 | 2.8505 | -2.162 | 2.5452 | -0.28166 | 8.596 |  | 10 | 10 |  | ZIRC-1 | 0.21786 | 0.00034 | 0.33305 | 0.00104 | 0.33714 | 0.00208 | 0.05408 | 0.0005 | -0.065295 | He et al (2021)
Lab 1 | Apr 4, 2022 | 1 | Zr_055 | iRM |  |  | 0.00049816 | 0.00010974 | 0.0001632 | 1.2241e-05 | 0.00015712 | -6.3098e-07 | 8.6715e-05 | -1.7464e-05 | 5.2474e-06 | [5 empty] | 19.4893 | 4.3309 | 6.7531 | 8.1326e-05 | 7.1126 | 0.00015983 | 1.1873 | 6.8945e-05 | 0.00019265 |  |  |  | 1 | 0 | 0 | 19.4888 | 4.3308 | 6.7529 | 6.9085e-05 | 7.1124 | 0.00016046 | 1.1872 | 8.6408e-05 | 0.0001874 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7549 | 19.4888 | 4.3308 | 6.7528 | 7.1124 | 1.1871 | 0.22222 | 0.3465 | 0.36495 | 0.060913 | 1.5593 | 1.6423 | 0.27411 | 1.0532 | 0.1758 | 0.16691 | -1.5041 | -1.0599 | -1.008 | -2.7983 | [10 empty] | 0 | 1.7547 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42503 | 0.43911 | -1.3882 | 0.014077 | -1.8132 | -1.8273 |  |  | Intensity on 90Zr | 22-23 V | 42-47 V | 11-29 V | iRM | Lab 1 | Apr 4, 2022 | 1 | 100.1% | 8.54 | -0.99325 | 14.7092 | [8 empty] | iRM | Lab 1 | Apr 5, 2022 | 3 | 100.9% | -8.715 | 3.1855 | 14.2855 | iRM_12 |  |  |  |  | Lab 1 | Apr 4, 2022 | Bottle 11 |  | Apr | 1 | 8 | -0.13677 | 3.3184 | -2.0315 | 1.4721 | 8.7566 | 8.6822 | Bottle 12 | 11.6 |  |  | N+83 | 0.21814 | 0.00022 | 0.33324 | 0.00013 | 0.33779 | 0.00021 | 0.0544 | 9e-05 | -0.021065
Lab 1 | Apr 4, 2022 | 1 | Zr_060 | iRM_1 |  |  | 0.00043281 | 9.6234e-05 | 0.00014454 | 9.651e-06 | 0.00013287 | -4.0126e-06 | 8.2797e-05 | -1.5385e-05 | 7.3024e-06 | [5 empty] | 19.5617 | 4.3468 | 6.7776 | 7.0882e-05 | 7.138 | 4.3195e-05 | 1.1914 | -3.709e-07 | 4.4053e-06 |  |  |  | 1 | 0 | 0 | 19.5613 | 4.3467 | 6.7775 | 6.1231e-05 | 7.1379 | 4.7207e-05 | 1.1913 | 1.5014e-05 | -2.8971e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7516 | 19.5613 | 4.3467 | 6.7775 | 7.1379 | 1.1913 | 0.22221 | 0.34648 | 0.3649 | 0.0609 | 1.5592 | 1.6421 | 0.27407 | 1.0532 | 0.17577 | 0.1669 | -1.5041 | -1.0599 | -1.0081 | -2.7985 | Lab 2 S1 | 92 vs 91 | 3.1197 |  | 2.0017 | 2.0028 | 1.3934 |  | 1.1611 |  | 0 | 1.7516 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43911 | -1.3882 | 0.014076 | -1.8133 | -1.8273 |  |  | Instrument sensitivity | 43-45 V / ppm of Zr | 233-261 V / ppm of Zr | 11-28 V / ppm of Zr | [8 empty] | iRM_1 | Lab 1 | Apr 4, 2022 | 1 | 100.2% | -6.8602 | 0.20517 | -6.0613 | iRM | Lab 1 | Apr 5, 2022 | 3 | 100.7% | -0.88083 | -1.9562 | -12.0644 | iRM_12 |  |  |  |  | Lab 1 | Jun 20, 2022 | Bottle 11 |  | June | 1 | 9 | 0.78423 | 2.7866 | -2.1024 | 3.5069 | 2.7517 | 6.7115 | Bottle 12 | 11.6 |  |  | Q+15 | 0.21785 | 0.00045 | 0.33304 | 0.00079 | 0.33734 | 0.00157 | 0.05416 | 0.0005 | -0.051677 | Quemet et al (2015)
Lab 1 | Apr 4, 2022 | 1 | Zr_065 | iRM |  |  | 0.00045116 | 9.9357e-05 | 0.0001445 | 9.9242e-06 | 0.00014462 | -1.376e-06 | 9.5573e-05 | -8.0225e-06 | 7.8342e-06 | [5 empty] | 19.5464 | 4.3435 | 6.7726 | 7.6234e-05 | 7.1328 | 0.0001609 | 1.1906 | 7.2736e-05 | 0.00019138 |  |  |  | 1 | 0 | 0 | 19.5459 | 4.3434 | 6.7724 | 6.631e-05 | 7.1326 | 0.00016228 | 1.1905 | 8.0759e-05 | 0.00018354 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7527 | 19.5459 | 4.3434 | 6.7723 | 7.1326 | 1.1904 | 0.22221 | 0.34648 | 0.36491 | 0.060904 | 1.5592 | 1.6422 | 0.27408 | 1.0532 | 0.17578 | 0.1669 | -1.5041 | -1.0599 | -1.0081 | -2.7985 |  | 94 vs 91 | 6.5707 |  | 4.0063 | 4.007 | 1.1399 |  | 3.1818 |  | 0 | 1.7525 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42503 | 0.43911 | -1.3882 | 0.014076 | -1.8133 | -1.8273 |  |  | Uptake time | 55 s | 55 s | 100 s | [31 empty] | Bottle 11 avg |  |  |  |  | 0.40335 | 2.134 | -2.0421 | 1.3574 | 4.9978 | 5.31 |  | 11.6 | 11.6
Lab 1 | Apr 4, 2022 | 1 | Zr_070 | iRM_2 |  |  | 0.00039698 | 8.9929e-05 | 0.00013501 | 1.6696e-05 | 0.00011126 | -3.0476e-06 | 7.9972e-05 | -2.5165e-06 | 8.1446e-06 | [5 empty] | 19.7381 | 4.3859 | 6.8387 | 7.1646e-05 | 7.2022 | 4.2874e-05 | 1.202 | 1.8843e-07 | 5.8437e-06 |  |  |  | 1 | 0 | 0 | 19.7377 | 4.3858 | 6.8385 | 5.4951e-05 | 7.2021 | 4.5921e-05 | 1.202 | 2.7049e-06 | -2.3009e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.751 | 19.7377 | 4.3858 | 6.8385 | 7.2021 | 1.202 | 0.22221 | 0.34647 | 0.36489 | 0.060896 | 1.5592 | 1.6421 | 0.27405 | 1.0532 | 0.17576 | 0.16689 | -1.5041 | -1.06 | -1.0082 | -2.7986 |  | 96 vs 91 | 1.6348 |  | 5.9765 | 5.9839 | 0.91382 |  | 4.1972 |  | 0 | 1.751 | 0.21794 | 0.33337 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014078 | -1.8133 | -1.8274 |  |  | Measurement | 1 block of 50 cycles | 3 blocks of 20 cycles | 1 block of 20 cycles | [8 empty] | iRM_2 | Lab 1 | Apr 4, 2022 | 1 | 101% | -11.8787 | -2.203 | -20.6067 | iRM | Lab 1 | Jun 20, 2022 | 1 | 102.9% | 5.3687 | -2.5492 | -15.0364 | iRM_12 | Avg | 1.3509 | -4.395 | -5.6578 | Lab 1 | Apr 4, 2022 | Bottle 12 |  | Apr | 1 | 34 | -0.40694 | 2.4031 | -0.099125 | 1.6099 | 0.85021 | 4.3445 | Bottle 12 | 13.2
Lab 1 | Apr 5, 2022 | 2 | Zr_005 | iRM_12 |  |  | 0.00029617 | 7.5819e-05 | 9.543e-05 | 2.2003e-05 | 8.1716e-05 | -2.4697e-06 | 0.00010947 | -1.7605e-06 | 3.9521e-06 | [5 empty] | 21.6426 | 4.8063 | 7.4894 | 6.8328e-05 | 7.8778 | 3.9949e-05 | 1.3132 | -1.0308e-07 | 3.2481e-06 |  |  |  | 1 | 0 | 0 | 21.6423 | 4.8062 | 7.4893 | 4.6325e-05 | 7.8777 | 4.2418e-05 | 1.3131 | 1.6574e-06 | -7.04e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6948 | 21.6423 | 4.8062 | 7.4893 | 7.8777 | 1.3131 | 0.22207 | 0.34605 | 0.364 | 0.060674 | 1.5583 | 1.6391 | 0.27322 | 1.0519 | 0.17534 | 0.16669 | -1.5047 | -1.0612 | -1.0106 | -2.8022 |  | 94 vs 92 | 2.1072 |  | 2.0024 | 2.0017 | 1.0544 |  | 1.0214 | L1 AprS2 | 0 | 1.6948 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42503 | 0.43909 | -1.3883 | 0.01406 | -1.8133 | -1.8274 |  |  | Cycle duration | 4.194 s | 4.194 s | 8.388 s | [16 empty] | iRM | Lab 1 | Jun 20, 2022 | 1 | 104.1% | 3.4708 | -4.7714 | -8.5831 | iRM_12 | 2SD | 9.2506 | 8.2109 | 17.5027 | Lab 1 | Jun 20, 2022 | Bottle 12 |  | June | 1 | 151 | -0.17555 | 1.1286 | 0.050546 | 0.82695 | 0.21684 | 2.2079 | Bottle 12 | 13.2
Lab 1 | Apr 5, 2022 | 2 | Zr_010 | iRM_1 |  |  | 0.00036435 | 8.6935e-05 | 0.00011487 | 1.3037e-05 | 0.0001105 | -4.2908e-06 | 0.00010565 | 4.8368e-06 | 1.4997e-06 | [5 empty] | 21.9698 | 4.8795 | 7.6044 | 7.1351e-05 | 8.0008 | 4.0538e-05 | 1.3341 | 5.9605e-07 | 3.2414e-06 |  |  |  | 1 | 0 | 0 | 21.9694 | 4.8794 | 7.6043 | 5.8314e-05 | 8.0007 | 4.4829e-05 | 1.334 | -4.2408e-06 | 1.7417e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.706 | 21.9694 | 4.8794 | 7.6043 | 8.0007 | 1.334 | 0.2221 | 0.34613 | 0.36417 | 0.060719 | 1.5585 | 1.6397 | 0.27338 | 1.0521 | 0.17542 | 0.16673 | -1.5046 | -1.0609 | -1.0101 | -2.8015 |  | 96 vs 92 | 0.52457 |  | 2.9889 | 2.9909 | 0.83887 |  | 0.87196 |  | 0 | 1.7059 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014061 | -1.8133 | -1.8274 |  |  | On peak zero | 10 cycles | 3 blocks of 20 cycles | 1 block of 20 cycles | [8 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 2 | 104.1% | 1.5479 | 6.5112 | -4.3196 | iRM | Lab 1 | Jun 20, 2022 | 1 | 104.1% | 0.45672 | -5.3748 | 1.1419 | iRM_12 | n | 9 | 9 | 9 |  |  | Bottle 12 avg |  |  |  |  | -0.21736 | 1.0216 | 0.019299 | 0.73558 | 0.34685 | 1.9683 |  | 13.2 | 13.2
Lab 1 | Apr 5, 2022 | 2 | Zr_015 | iRM_12 |  |  | 0.00040017 | 9.5753e-05 | 0.00013374 | 1.0993e-05 | 0.00011097 | -3.2809e-06 | 9.9868e-05 | -1.4524e-05 | 2.6999e-06 | [5 empty] | 20.5615 | 4.567 | 7.1179 | 6.3433e-05 | 7.49 | 3.4615e-05 | 1.2491 | 1.3691e-06 | 1.0864e-06 |  |  |  | 1 | 0 | 0 | 20.5611 | 4.5669 | 7.1177 | 5.244e-05 | 7.4898 | 3.7896e-05 | 1.249 | 1.5894e-05 | -1.6135e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7122 | 20.5611 | 4.5669 | 7.1177 | 7.4898 | 1.249 | 0.22211 | 0.34618 | 0.36427 | 0.060744 | 1.5586 | 1.64 | 0.27348 | 1.0523 | 0.17547 | 0.16675 | -1.5046 | -1.0608 | -1.0099 | -2.8011 |  | 96 vs 94 | 0.24903 |  | 1.4932 | 1.4947 | 0.6529 |  | -0.66469 |  | 0 | 1.7122 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014074 | -1.8133 | -1.8274 |  |  | Cycle duration | 4.194 s | 4.194 s | 8.388 s | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 98.55% | -7.5159 | -8.6024 | 9.85 | [8 empty] | iRM | Lab 1 | Jun 20, 2022 | 1 | 105% | 1.3611 | -6.9592 | 1.6414 | iRM_12 | 2SE | 3.0835 | 2.737 | 5.8342 | Lab 1 | Jun 21, 2022 | SRM 3169 | 130920 | June | 3 | 21 | 1.8116 | 2.8289 | -0.38309 | 2.1369 | 1.5634 | 4.9859 | Bottle 12 | 14.4 | 14.4
Lab 1 | Apr 5, 2022 | 2 | Zr_020 | iRM_2 |  |  | 0.00037837 | 9.4239e-05 | 0.00012194 | 1.0683e-05 | 0.00011076 | -1.3159e-06 | 9.884e-05 | -5.9024e-06 | 4.859e-06 | [5 empty] | 20.4101 | 4.5336 | 7.0661 | 6.855e-05 | 7.4362 | 3.8366e-05 | 1.2402 | -4.5549e-06 | 8.9813e-07 |  |  |  | 1 | 0 | 0 | 20.4097 | 4.5335 | 7.066 | 5.7867e-05 | 7.436 | 3.9681e-05 | 1.2401 | 1.3475e-06 | -3.9609e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7163 | 20.4097 | 4.5335 | 7.066 | 7.436 | 1.2401 | 0.22212 | 0.34621 | 0.36434 | 0.06076 | 1.5586 | 1.6402 | 0.27354 | 1.0524 | 0.1755 | 0.16677 | -1.5045 | -1.0607 | -1.0097 | -2.8008 |  |  | Correlation not from CS. |  |  | Best estimate for n: | 0.99886 |  |  |  | 0 | 1.7163 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3883 | 0.014069 | -1.8133 | -1.8274 |  |  | Rinse duration | 568 s | 310 s | 150 s | [8 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 2 | 100.4% | 10.077 | 3.23 | -3.9028 | iRM | Lab 1 | Jun 20, 2022 | 2 | 104.9% | 3.7475 | -7.8556 | -16.8316 | iRM_12
Lab 1 | Apr 5, 2022 | 2 | Zr_025 | iRM_12 |  |  | 0.00038545 | 8.8126e-05 | 0.00011817 | 1.1764e-05 | 0.00011359 | -7.7789e-06 | 9.1455e-05 | -1.0962e-05 | 1.8269e-06 | [5 empty] | 20.0837 | 4.4611 | 6.9534 | 6.2545e-05 | 7.3179 | 3.4645e-05 | 1.2205 | -6.1853e-06 | -6.7307e-07 |  |  |  | 1 | 0 | 0 | 20.0833 | 4.461 | 6.9533 | 5.0781e-05 | 7.3178 | 4.2424e-05 | 1.2204 | 4.7765e-06 | -2.5e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7184 | 20.0833 | 4.461 | 6.9533 | 7.3178 | 1.2204 | 0.22213 | 0.34622 | 0.36437 | 0.060768 | 1.5587 | 1.6404 | 0.27357 | 1.0524 | 0.17552 | 0.16678 | -1.5045 | -1.0607 | -1.0096 | -2.8007 | [10 empty] | 0 | 1.7184 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8274 |  |  | Bold font in the "Cup configuration" and "Isotopes monitored" rows denotes the central cup. |  |  |  | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.25% | -4.8912 | 0.45562 | 6.9415 | [8 empty] | iRM | Lab 1 | Jun 20, 2022 | 2 | 106% | -1.4952 | -3.3985 | -13.2823 | iRM_12 | [6 empty] | Laboratory 2
Lab 1 | Apr 5, 2022 | 2 | Zr_030 | iRM_10 |  |  | 0.00038335 | 9.2361e-05 | 0.0001209 | 3.5967e-06 | 0.00010293 | -2.7672e-06 | 8.0788e-05 | -1.2878e-05 | -7.5879e-07 | [5 empty] | 20.9873 | 4.6621 | 7.2671 | 6.6903e-05 | 7.6489 | 3.7522e-05 | 1.2759 | -1.2023e-06 | 2.3882e-07 |  |  |  | 1 | 0 | 0 | 20.9869 | 4.6621 | 7.267 | 6.3306e-05 | 7.6488 | 4.0289e-05 | 1.2758 | 1.1676e-05 | 9.9761e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7237 | 20.9869 | 4.6621 | 7.267 | 7.6488 | 1.2758 | 0.22214 | 0.34626 | 0.36446 | 0.06079 | 1.5588 | 1.6406 | 0.27365 | 1.0525 | 0.17556 | 0.1668 | -1.5044 | -1.0606 | -1.0094 | -2.8003 | Lab 2 S2 | 92 vs 91 | 3.0432 |  | 1.954 | 1.9528 | -3.2529 |  | 1.1134 |  | 0 | 1.7237 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014074 | -1.8133 | -1.8273 | [14 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 2 | 105% | 8.3578 | -4.1183 | 18.2493 | iRM | Lab 1 | Jun 20, 2022 | 2 | 107.4% | -8.6617 | -11.1932 | 2.1736 | iRM_12 |  |  |  |  | Lab 2 | May 28, 2020 | Bottle 1 |  | May | 1 | 6 | -2.4926 | 2.9394 | -0.016189 | 2.7527 | 5.2006 | 10.0528 | SRM 3169 | 1.2
Lab 1 | Apr 5, 2022 | 2 | Zr_035 | iRM_12 |  |  | 0.000373 | 8.1831e-05 | 0.00011762 | 1.6921e-05 | 0.0001044 | -1.7399e-07 | 9.7905e-05 | -5.9713e-06 | 2.6984e-06 | [5 empty] | 19.91 | 4.4228 | 6.8941 | 6.1824e-05 | 7.2562 | 3.4157e-05 | 1.2103 | 1.8349e-07 | -1.9726e-06 |  |  |  | 1 | 0 | 0 | 19.9096 | 4.4227 | 6.8939 | 4.4903e-05 | 7.2561 | 3.4331e-05 | 1.2102 | 6.1548e-06 | -4.671e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7234 | 19.9096 | 4.4227 | 6.894 | 7.2561 | 1.2102 | 0.22214 | 0.34626 | 0.36445 | 0.060787 | 1.5588 | 1.6406 | 0.27364 | 1.0525 | 0.17555 | 0.16679 | -1.5045 | -1.0606 | -1.0094 | -2.8004 |  | 94 vs 91 | 6.2663 |  | 3.8283 | 3.8276 | -1.6888 |  | 3.0037 |  | 0 | 1.7234 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 100.4% | 1.6111 | 1.5565 | -17.5867 | [8 empty] | iRM | Lab 1 | Jun 20, 2022 | 2 | 106.9% | 0.63437 | 1.6652 | -9.1426 | iRM_12 |  |  |  |  | Lab 2 | June 3, 2020 | Bottle 1 |  | June | 2 | 6 | 0.4227 | 3.3717 | 2.2713 | 2.4214 | -2.6492 | 6.4568 | SRM 3169 | 1.2
Lab 1 | Apr 5, 2022 | 2 | Zr_040 | iRM_11 |  |  | 0.00038344 | 9.8706e-05 | 0.00012214 | 1.5741e-05 | 0.00011561 | -2.2657e-06 | 7.8197e-05 | -4.9126e-06 | -9.4579e-07 | [5 empty] | 20.6252 | 4.5819 | 7.1422 | 6.4667e-05 | 7.5179 | 3.5454e-05 | 1.2541 | 7.4234e-07 | 3.0735e-07 |  |  |  | 1 | 0 | 0 | 20.6248 | 4.5818 | 7.1421 | 4.8926e-05 | 7.5178 | 3.7719e-05 | 1.254 | 5.655e-06 | 1.2531e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7266 | 20.6248 | 4.5818 | 7.1421 | 7.5178 | 1.254 | 0.22215 | 0.34629 | 0.3645 | 0.060801 | 1.5588 | 1.6408 | 0.27369 | 1.0526 | 0.17558 | 0.1668 | -1.5044 | -1.0605 | -1.0092 | -2.8002 |  | 96 vs 91 | 1.5245 |  | 5.5915 | 5.5831 | -1.6876 |  | 3.8122 |  | 0 | 1.7266 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014072 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 2 | 104.4% | 0.44544 | -3.9419 | 9.5767 | iRM | Lab 1 | Jun 20, 2022 | 2 | 108.7% | 7.2755 | 0.88182 | 6.9992 | iRM_12 | [6 empty] | Bottle 1 avg |  |  |  |  | -1.2337 | 2.2156 | 1.2734 | 1.8181 | -0.35669 | 5.4327 |  | 1.2 | 1.2
Lab 1 | Apr 5, 2022 | 2 | Zr_045 | iRM_12 |  |  | 0.00038169 | 8.1411e-05 | 0.00012734 | 9.537e-06 | 0.00010458 | -4.1373e-06 | 8.3304e-05 | -1.2701e-05 | -2.6905e-07 | [5 empty] | 19.5922 | 4.3523 | 6.7843 | 5.9408e-05 | 7.1409 | 3.4234e-05 | 1.1912 | -2.1239e-06 | -1.3461e-06 |  |  |  | 1 | 0 | 0 | 19.5918 | 4.3522 | 6.7841 | 4.9871e-05 | 7.1407 | 3.8371e-05 | 1.1911 | 1.0577e-05 | -1.0771e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.725 | 19.5918 | 4.3522 | 6.7841 | 7.1407 | 1.1911 | 0.22214 | 0.34627 | 0.36448 | 0.060794 | 1.5588 | 1.6407 | 0.27367 | 1.0526 | 0.17557 | 0.1668 | -1.5044 | -1.0605 | -1.0093 | -2.8003 |  | 94 vs 92 | 2.059 |  | 1.9592 | 1.96 | -0.89555 |  | 0.97816 |  | 0 | 1.725 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.39% | -4.885 | -0.55477 | 9.0958 | [21 empty] | Lab 2 | May 28, 2020 | Bottle 2 |  | May | 1 | 6 | 2.3114 | 5.2323 | 1.6696 | 1.9558 | -6.6929 | 7.0329 | SRM 3169 | 2.8
Lab 1 | Apr 5, 2022 | 2 | Zr_050 | iRM |  |  | 0.00038013 | 8.8286e-05 | 0.00012368 | 9.7229e-06 | 0.00010956 | 2.1093e-06 | 7.5863e-05 | -1.6429e-05 | 1.9805e-06 | [5 empty] | 19.6281 | 4.3604 | 6.797 | 5.7811e-05 | 7.1546 | 3.5037e-05 | 1.1935 | -4.0507e-06 | 8.1589e-07 |  |  |  | 1 | 0 | 0 | 19.6277 | 4.3603 | 6.7968 | 4.8088e-05 | 7.1545 | 3.2927e-05 | 1.1934 | 1.2378e-05 | -1.1646e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7272 | 19.6277 | 4.3603 | 6.7968 | 7.1545 | 1.1934 | 0.22215 | 0.34629 | 0.36451 | 0.060804 | 1.5588 | 1.6408 | 0.27371 | 1.0526 | 0.17559 | 0.16681 | -1.5044 | -1.0605 | -1.0092 | -2.8001 |  | 96 vs 92 | 0.50103 |  | 2.8621 | 2.8596 | -1.2725 |  | 0.74515 |  | 0 | 1.7272 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014077 | -1.8133 | -1.8273 | [14 empty] | iRM | Lab 1 | Apr 5, 2022 | 2 | 100.4% | -4.7411 | -8.7339 | 15.3966 | iRM_1 | Lab 1 | Apr 4, 2022 | 1 | 97.45% | -3.6088 | -4.9876 | -15.8405 | iRM | Avg | -5.4788 | -0.3949 | 5.6246 | Lab 2 | June 3, 2020 | Bottle 2 |  | June | 2 | 6 | 0.045558 | 2.8934 | 3.6719 | 2.0129 | 0.88254 | 14.396 | SRM 3169 | 2.8
Lab 1 | Apr 5, 2022 | 2 | Zr_055 | iRM_12 |  |  | 0.00037366 | 8.4271e-05 | 0.00011938 | 1.1509e-05 | 9.9364e-05 | -1.0564e-06 | 8.9165e-05 | -7.0181e-06 | 8.3845e-07 | [5 empty] | 19.5147 | 4.3353 | 6.7579 | 5.7986e-05 | 7.1136 | 3.2945e-05 | 1.1867 | 2.418e-06 | -9.8162e-08 |  |  |  | 1 | 0 | 0 | 19.5143 | 4.3352 | 6.7578 | 4.6477e-05 | 7.1135 | 3.4001e-05 | 1.1866 | 9.4361e-06 | -9.3661e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7281 | 19.5143 | 4.3352 | 6.7578 | 7.1135 | 1.1866 | 0.22215 | 0.3463 | 0.36452 | 0.060806 | 1.5588 | 1.6409 | 0.27371 | 1.0526 | 0.17559 | 0.16681 | -1.5044 | -1.0605 | -1.0092 | -2.8001 |  | 96 vs 94 | 0.24338 |  | 1.4611 | 1.4593 | -1.639 |  | -0.69677 |  | 0 | 1.7281 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.55% | 9.831 | 3.5015 | -2.1286 | [8 empty] | iRM_1 | Lab 1 | Apr 4, 2022 | 1 | 100.2% | -6.8602 | 0.20517 | -6.0613 | iRM | 2SD | 11.197 | 10.0308 | 27.4881 |  |  | Bottle 2 avg |  |  |  |  | 0.57617 | 2.532 | 2.6419 | 1.4027 | -5.2333 | 6.3191 |  | 2.8 | 2.8
Lab 1 | Apr 5, 2022 | 2 | Zr_060 | iRM_1 |  |  | 0.00036879 | 9.5254e-05 | 0.00012018 | 1.9553e-05 | 0.00010521 | -2.6101e-06 | 7.776e-05 | -1.9374e-05 | -7.3377e-07 | [5 empty] | 20.4643 | 4.5462 | 7.0871 | 6.3669e-05 | 7.4604 | 3.5628e-05 | 1.2446 | -6.4793e-06 | -5.2821e-07 |  |  |  | 1 | 0 | 0 | 20.4639 | 4.5462 | 7.087 | 4.4116e-05 | 7.4603 | 3.8238e-05 | 1.2445 | 1.2895e-05 | 2.0556e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7302 | 20.4639 | 4.5462 | 7.087 | 7.4603 | 1.2445 | 0.22215 | 0.34632 | 0.36456 | 0.060816 | 1.5589 | 1.641 | 0.27375 | 1.0527 | 0.17561 | 0.16682 | -1.5044 | -1.0604 | -1.0091 | -2.7999 |  |  | Correlation not from CS. |  |  | Best estimate for n: | -1.4367 |  |  |  | 0 | 1.7302 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014066 | -1.8133 | -1.8273 | [14 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 2 | 104.6% | -11.898 | 5.4843 | 12.2471 | iRM_1 | Lab 1 | Apr 5, 2022 | 2 | 104.1% | 1.5479 | 6.5112 | -4.3196 | iRM_12 | n | 9 | 9 | 9 | Lab 2 | May 28, 2020 | Bottle 3 |  | May | 1 | 6 | -1.2489 | 2.6096 | -0.86351 | 0.98976 | -7.2463 | 8.4435 | SRM 3169 | 4
Lab 1 | Apr 5, 2022 | 2 | Zr_065 | iRM_12 |  |  | 0.00037768 | 8.0843e-05 | 0.00012166 | 1.5156e-05 | 0.00010312 | -7.95e-07 | 7.6395e-05 | -7.5519e-06 | 2.6745e-06 | [5 empty] | 19.6127 | 4.3571 | 6.7921 | 6.1484e-05 | 7.1498 | 3.2603e-05 | 1.1928 | 2.5191e-06 | 1.5284e-06 |  |  |  | 1 | 0 | 0 | 19.6124 | 4.357 | 6.7919 | 4.6328e-05 | 7.1497 | 3.3398e-05 | 1.1927 | 1.0071e-05 | -1.146e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7298 | 19.6124 | 4.357 | 6.7919 | 7.1497 | 1.1927 | 0.22216 | 0.34631 | 0.36455 | 0.060813 | 1.5589 | 1.641 | 0.27374 | 1.0527 | 0.1756 | 0.16682 | -1.5044 | -1.0604 | -1.0091 | -2.7999 | [10 empty] | 0 | 1.7298 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 100.9% | -8.6841 | -5.4091 | 0.09599 | [8 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 2 | 104.6% | -11.898 | 5.4843 | 12.2471 | iRM_12 | 2SE | 3.7323 | 3.3436 | 9.1627 | Lab 2 | June 3, 2020 | Bottle 3 |  | June | 2 | 6 | -2.3408 | 3.728 | -0.46757 | 1.7044 | 2.1217 | 14.9728 | SRM 3169 | 4
Lab 1 | Apr 5, 2022 | 2 | Zr_070 | iRM_2 |  |  | 0.00041412 | 8.7688e-05 | 0.00012978 | 1.9418e-05 | 0.00012574 | 8.8149e-07 | 8.7815e-05 | -1.639e-05 | -4.7432e-07 | [5 empty] | 19.7206 | 4.3811 | 6.8296 | 6.4313e-05 | 7.1895 | 3.4713e-05 | 1.1994 | -8.5449e-06 | 1.1465e-07 |  |  |  | 1 | 0 | 0 | 19.7202 | 4.381 | 6.8294 | 4.4895e-05 | 7.1893 | 3.3832e-05 | 1.1993 | 7.8448e-06 | 5.8897e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7307 | 19.7202 | 4.381 | 6.8294 | 7.1893 | 1.1993 | 0.22216 | 0.34632 | 0.36457 | 0.060815 | 1.5589 | 1.641 | 0.27375 | 1.0527 | 0.17561 | 0.16682 | -1.5044 | -1.0604 | -1.009 | -2.7999 | Lab 3 | 92 vs 91 | 3.1401 |  | 2.0074 | 2.0074 | 1.8068 |  | 1.1668 |  | 0 | 1.7307 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014073 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 2 | 101.2% | -7.3462 | 0.12981 | -32.2272 | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 104.6% | -12.7717 | -4.0036 | -5.3054 | iRM_12 | [6 empty] | Bottle 3 avg |  |  |  |  | -1.608 | 2.1378 | -0.76366 | 0.85591 | -4.986 | 7.3547 |  | 4 | 4
Lab 1 | Apr 5, 2022 | 2 | Zr_075 | iRM_12 |  |  | 0.00040642 | 8.3289e-05 | 0.00013438 | 1.2196e-05 | 0.00011512 | -1.763e-06 | 8.3254e-05 | -9.1784e-07 | 3.0279e-06 | [5 empty] | 19.3774 | 4.3048 | 6.7106 | 6.2005e-05 | 7.0641 | 3.1635e-05 | 1.1785 | -3.8383e-06 | 6.0761e-08 |  |  |  | 1 | 0 | 0 | 19.377 | 4.3048 | 6.7105 | 4.9808e-05 | 7.064 | 3.3398e-05 | 1.1784 | -2.9205e-06 | -2.9672e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.73 | 19.377 | 4.3048 | 6.7105 | 7.064 | 1.1784 | 0.22216 | 0.34631 | 0.36456 | 0.060815 | 1.5589 | 1.641 | 0.27375 | 1.0527 | 0.17561 | 0.16682 | -1.5044 | -1.0604 | -1.0091 | -2.7999 |  | 94 vs 91 | 6.6008 |  | 3.9798 | 3.9798 | 0.7231 |  | 3.1552 |  | 0 | 1.73 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.3% | 5.2606 | 2.3574 | 0.83388 | [8 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 105.7% | -6.5697 | -4.9003 | 25.3421 | iRM_12 |  |  |  |  | Lab 2 | May 28, 2020 | Bottle 4 |  | May | 1 | 6 | 1.1912 | 3.1224 | 2.339 | 2.3098 | 1.4715 | 4.726 | SRM 3169 | 5.2
Lab 1 | Apr 5, 2022 | 2 | Zr_080 | iRM_10 |  |  | 0.00042064 | 0.0001011 | 0.00014354 | 1.3995e-06 | 0.00012094 | -4.1983e-06 | 8.4147e-05 | -1.1219e-05 | 9.1272e-07 | [5 empty] | 20.3468 | 4.5203 | 7.0467 | 6.009e-05 | 7.4182 | 3.3625e-05 | 1.2376 | -4.8563e-06 | -3.4566e-07 |  |  |  | 1 | 0 | 0 | 20.3464 | 4.5202 | 7.0465 | 5.869e-05 | 7.4181 | 3.7823e-05 | 1.2375 | 6.3631e-06 | -1.2584e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7322 | 20.3464 | 4.5202 | 7.0465 | 7.4181 | 1.2375 | 0.22216 | 0.34633 | 0.36459 | 0.060823 | 1.5589 | 1.6411 | 0.27378 | 1.0527 | 0.17562 | 0.16683 | -1.5044 | -1.0604 | -1.009 | -2.7998 |  | 96 vs 91 | 1.6554 |  | 5.942 | 5.942 | 0.65373 |  | 4.1626 |  | 0 | 1.7322 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [14 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 2 | 104.9% | -10.6246 | 0.098244 | -5.1414 | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 108% | 4.3646 | 5.8517 | 16.4011 | iRM_12 |  |  |  |  | Lab 2 | June 3, 2020 | Bottle 4 |  | June | 2 | 6 | -4.5834 | 5.2312 | -0.95722 | 4.0738 | 5.7156 | 16.3024 | SRM 3169 | 5.2
Lab 1 | Apr 5, 2022 | 2 | Zr_085 | iRM_12 |  |  | 0.00039701 | 7.7254e-05 | 0.0001281 | 8.4833e-06 | 0.0001124 | -3.9055e-06 | 7.5348e-05 | -2.2853e-05 | 3.5086e-06 | [5 empty] | 19.4159 | 4.3134 | 6.7241 | 5.6314e-05 | 7.0785 | 3.1096e-05 | 1.1809 | -8.6797e-06 | -4.527e-07 |  |  |  | 1 | 0 | 0 | 19.4155 | 4.3133 | 6.724 | 4.7831e-05 | 7.0784 | 3.5002e-05 | 1.1808 | 1.4173e-05 | -3.9613e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.731 | 19.4155 | 4.3133 | 6.724 | 7.0784 | 1.1808 | 0.22216 | 0.34632 | 0.36457 | 0.060819 | 1.5589 | 1.641 | 0.27376 | 1.0527 | 0.17561 | 0.16682 | -1.5044 | -1.0604 | -1.009 | -2.7999 |  | 94 vs 92 | 2.1021 |  | 1.9826 | 1.9826 | 0.1723 |  | 1.0016 |  | 0 | 1.7311 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.96% | -3.1594 | -1.1881 | 6.2284 | [8 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 108.5% | -6.422 | -3.0818 | 15.4585 | iRM_12 | [6 empty] | Bottle 4 avg |  |  |  |  | -0.32574 | 2.6812 | 1.5372 | 2.0093 | 1.8005 | 4.5391 |  | 5.2 | 5.2
Lab 1 | Apr 5, 2022 | 2 | Zr_090 | iRM_11 |  |  | 0.00037677 | 7.7983e-05 | 0.00012234 | 1.0828e-05 | 0.00010825 | -8.1134e-08 | 6.9192e-05 | -7.9746e-06 | -4.4618e-07 | [5 empty] | 20.358 | 4.5228 | 7.0506 | 6.2126e-05 | 7.4225 | 3.3447e-05 | 1.2383 | 2.1307e-06 | -1.3079e-06 |  |  |  | 1 | 0 | 0 | 20.3576 | 4.5227 | 7.0505 | 5.1299e-05 | 7.4224 | 3.3529e-05 | 1.2383 | 1.0105e-05 | -8.6174e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7328 | 20.3576 | 4.5227 | 7.0505 | 7.4224 | 1.2383 | 0.22216 | 0.34633 | 0.3646 | 0.060826 | 1.5589 | 1.6411 | 0.27379 | 1.0527 | 0.17563 | 0.16683 | -1.5043 | -1.0604 | -1.009 | -2.7997 |  | 96 vs 92 | 0.52719 |  | 2.9601 | 2.9601 | 0.35863 |  | 0.84322 |  | 0 | 1.7328 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014072 | -1.8133 | -1.8273 | [14 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 2 | 104.7% | -4.2944 | -2.4135 | 6.3647 | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 107.7% | -7.0913 | -4.6331 | 12.6998 | iRM_12 |  |  |  |  | Lab 2 | May 28, 2020 | Bottle 5 |  | May | 1 | 6 | 1.2768 | 4.4285 | 1.4146 | 1.3264 | 5.448 | 5.763 | SRM 3169 | 6.8
Lab 1 | Apr 5, 2022 | 2 | Zr_095 | iRM_12 |  |  | 0.00040082 | 9.2274e-05 | 0.00012712 | 3.5727e-06 | 0.00012687 | -2.9506e-06 | 9.1808e-05 | -2.2436e-06 | -2.4174e-06 | [5 empty] | 19.4696 | 4.3254 | 6.7428 | 5.6375e-05 | 7.0982 | 3.2278e-05 | 1.1842 | -1.2047e-06 | -1.2879e-06 |  |  |  | 1 | 0 | 0 | 19.4692 | 4.3253 | 6.7427 | 5.2802e-05 | 7.0981 | 3.5229e-05 | 1.1841 | 1.0389e-06 | 1.1295e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7317 | 19.4692 | 4.3253 | 6.7427 | 7.0981 | 1.1841 | 0.22216 | 0.34633 | 0.36458 | 0.060821 | 1.5589 | 1.6411 | 0.27377 | 1.0527 | 0.17562 | 0.16682 | -1.5043 | -1.0604 | -1.009 | -2.7998 |  | 96 vs 94 | 0.25079 |  | 1.4931 | 1.4931 | 0.5486 |  | -0.66479 |  | 0 | 1.7317 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014069 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 100.3% | 8.242 | 2.0289 | -17.0452 | [21 empty] | Lab 2 | June 3, 2020 | Bottle 5 |  | June | 2 | 6 | -4.754 | 3.8665 | -0.76649 | 4.2236 | 16.2937 | 13.4868 | SRM 3169 | 6.8
Lab 1 | Apr 5, 2022 | 2 | Zr_100 | iRM |  |  | 0.0003776 | 8.1236e-05 | 0.00011514 | 1.3994e-05 | 0.00012196 | -4.8289e-06 | 8.5352e-05 | 1.2024e-06 | 1.0416e-06 | [5 empty] | 19.2758 | 4.2824 | 6.6758 | 5.5408e-05 | 7.0279 | 3.0531e-05 | 1.1725 | -1.9388e-07 | -3.2061e-07 |  |  |  | 1 | 0 | 0 | 19.2754 | 4.2823 | 6.6757 | 4.1414e-05 | 7.0278 | 3.5359e-05 | 1.1724 | -1.3963e-06 | -1.3623e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7326 | 19.2754 | 4.2823 | 6.6757 | 7.0278 | 1.1724 | 0.22216 | 0.34633 | 0.3646 | 0.060825 | 1.5589 | 1.6411 | 0.27378 | 1.0527 | 0.17563 | 0.16683 | -1.5043 | -1.0604 | -1.009 | -2.7998 |  |  | Correlation not from CS. |  |  | Best estimate for n: | 0.71053 |  |  |  | 0 | 1.7326 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [14 empty] | iRM | Lab 1 | Apr 5, 2022 | 2 | 99.17% | -1.513 | -6.4571 | -8.0459 | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 99.91% | 8.741 | -6.0315 | -5.0627 | iRM_12 | Avg | 0.77891 | -2.1173 | -0.95373 |  |  | Bottle 5 avg |  |  |  |  | -2.1453 | 2.9126 | 1.2188 | 1.2654 | 7.1225 | 5.2995 |  | 6.8 | 6.8
Lab 1 | Apr 5, 2022 | 2 | Zr_105 | iRM_12 |  |  | 0.00043773 | 0.00010667 | 0.00014681 | 9.1497e-06 | 0.00014092 | 1.6166e-07 | 7.9073e-05 | -6.5859e-06 | 2.4742e-06 | [5 empty] | 19.406 | 4.3113 | 6.7209 | 5.8322e-05 | 7.0753 | 3.2364e-05 | 1.1804 | 1.5371e-06 | 3.5583e-07 |  |  |  | 1 | 0 | 0 | 19.4055 | 4.3112 | 6.7207 | 4.9172e-05 | 7.0751 | 3.2202e-05 | 1.1803 | 8.123e-06 | -2.1183e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7324 | 19.4055 | 4.3112 | 6.7207 | 7.0751 | 1.1803 | 0.22216 | 0.34633 | 0.36459 | 0.060825 | 1.5589 | 1.6411 | 0.27379 | 1.0527 | 0.17563 | 0.16683 | -1.5044 | -1.0604 | -1.009 | -2.7998 | [10 empty] | 0 | 1.7324 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [22 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 100.5% | 4.1243 | 2.0526 | 13.0788 | iRM_12 | 2SD | 12.2664 | 7.0113 | 25.5348 | Lab 2 | May 28, 2020 | Bottle 6 |  | May | 1 | 6 | -1.6342 | 5.4225 | -0.45386 | 5.0947 | 8.38 | 16.078 | SRM 3169 | 8.4
Lab 1 | Apr 5, 2022 | 3 | Zr_005 | iRM_12 |  |  | 0.00029079 | 6.5521e-05 | 9.624e-05 | 2.595e-05 | 8.2758e-05 | -3.2077e-06 | 8.1141e-05 | -1.3737e-06 | 2.7013e-06 | [5 empty] | 19.5117 | 4.3347 | 6.7574 | 5.5824e-05 | 7.1136 | 3.2048e-05 | 1.1868 | 7.553e-06 | -1.1689e-06 |  |  |  | 1 | 0 | 0 | 19.5114 | 4.3347 | 6.7573 | 2.9874e-05 | 7.1135 | 3.5255e-05 | 1.1867 | 8.9267e-06 | -3.8702e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7317 | 19.5114 | 4.3347 | 6.7573 | 7.1135 | 1.1867 | 0.22216 | 0.34632 | 0.36458 | 0.060821 | 1.5589 | 1.6411 | 0.27377 | 1.0527 | 0.17562 | 0.16682 | -1.5044 | -1.0604 | -1.009 | -2.7998 | [9 empty] | L1 AprS3 | 0 | 1.7317 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014072 | -1.8133 | -1.8273 | [22 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 101.9% | 4.6207 | 0.9203 | -26.4963 | iRM_12 | n | 10 | 10 | 10 | Lab 2 | June 3, 2020 | Bottle 6 |  | June | 2 | 6 | -3.1847 | 3.134 | 1.623 | 2.7823 | 1.7262 | 8.9469 | SRM 3169 | 8.4
Lab 1 | Apr 5, 2022 | 3 | Zr_010 | iRM_1 |  |  | 0.00036762 | 8.2453e-05 | 0.00011175 | 4.7523e-06 | 0.00010968 | -1.757e-06 | 8.4518e-05 | -2.7446e-06 | 1.8875e-06 | [5 empty] | 20.4207 | 4.5368 | 7.0726 | 6.4213e-05 | 7.4459 | 3.2731e-05 | 1.2423 | -2.9546e-06 | -2.7191e-07 |  |  |  | 1 | 0 | 0 | 20.4203 | 4.5367 | 7.0725 | 5.9461e-05 | 7.4458 | 3.4488e-05 | 1.2422 | -2.1007e-07 | -2.1594e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7345 | 20.4203 | 4.5367 | 7.0725 | 7.4458 | 1.2422 | 0.22217 | 0.34635 | 0.36463 | 0.060832 | 1.559 | 1.6412 | 0.27381 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7996 | [10 empty] | 0 | 1.7345 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [14 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 104.6% | -12.7717 | -4.0036 | -5.3054 | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 102.4% | 0.42284 | -4.3221 | 9.4976 | iRM_12 | 2SE | 3.879 | 2.2172 | 8.0748 |  |  | Bottle 6 avg |  |  |  |  | -2.7964 | 2.7134 | 1.1459 | 2.4419 | 3.2994 | 7.8179 |  | 8.4 | 8.4
Lab 1 | Apr 5, 2022 | 3 | Zr_015 | iRM_12 |  |  | 0.00037941 | 8.5896e-05 | 0.00012299 | 1.675e-05 | 0.00011633 | -2.6239e-06 | 8.1951e-05 | -1.0818e-05 | -2.2694e-06 | [5 empty] | 19.5461 | 4.3425 | 6.7695 | 5.8685e-05 | 7.1265 | 3.6806e-05 | 1.189 | -5.6199e-06 | -1.1037e-06 |  |  |  | 1 | 0 | 0 | 19.5457 | 4.3424 | 6.7694 | 4.1935e-05 | 7.1264 | 3.943e-05 | 1.1889 | 5.1978e-06 | 1.1657e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.733 | 19.5457 | 4.3424 | 6.7694 | 7.1264 | 1.1889 | 0.22217 | 0.34634 | 0.3646 | 0.060827 | 1.5589 | 1.6411 | 0.27379 | 1.0527 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.733 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43911 | -1.3882 | 0.014061 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.99% | 11.6984 | 10.4048 | 22.7649 | [8 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 103.2% | 8.3736 | 1.0795 | -2.2693 | iRM_12 |  |  |  |  | Lab 2 | May 28, 2020 | SRM 3169 | 071226 | May | 1 | 37 | 0.067358 | 1.3844 | 0.031281 | 1.0895 | 0.0080002 | 4.3115 | SRM 3169 | 14.8
Lab 1 | Apr 5, 2022 | 3 | Zr_020 | iRM_2 |  |  | 0.00037316 | 8.6161e-05 | 0.00011193 | 1.7446e-05 | 0.00010548 | 3.8391e-07 | 0.00010278 | -4.7804e-06 | 2.0006e-06 | [5 empty] | 19.9063 | 4.4225 | 6.8943 | 6.6374e-05 | 7.2582 | 2.9557e-05 | 1.2109 | -1.0361e-06 | 1.3555e-06 |  |  |  | 1 | 0 | 0 | 19.9059 | 4.4224 | 6.8942 | 4.8928e-05 | 7.2581 | 2.9173e-05 | 1.2108 | 3.7444e-06 | -6.4515e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.734 | 19.9059 | 4.4224 | 6.8942 | 7.2581 | 1.2108 | 0.22217 | 0.34634 | 0.36462 | 0.060828 | 1.5589 | 1.6412 | 0.2738 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.734 | 0.21794 | 0.33337 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3883 | 0.014077 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 101.7% | 3.5542 | -10.6569 | -27.0886 | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 103.7% | 2.283 | -0.5952 | -0.082763 | iRM_12 |  |  |  |  | Lab 2 | June 3, 2020 | SRM 3169 | 071226 | June | 2 | 7 | 0.46094 | 2.9916 | -0.14036 | 3.6108 | 0.2085 | 7.0333 | SRM 3169 | 14.8
Lab 1 | Apr 5, 2022 | 3 | Zr_025 | iRM_12 |  |  | 0.00039503 | 9.6087e-05 | 0.00012798 | 1.8191e-05 | 0.00011156 | -8.1977e-07 | 9.0893e-05 | -9.4025e-06 | 2.2119e-06 | [5 empty] | 19.5851 | 4.3511 | 6.783 | 5.8544e-05 | 7.1408 | 3.2063e-05 | 1.1913 | -1.4974e-06 | -1.3391e-07 |  |  |  | 1 | 0 | 0 | 19.5847 | 4.351 | 6.7829 | 4.0353e-05 | 7.1407 | 3.2882e-05 | 1.1912 | 7.9051e-06 | -2.3458e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.733 | 19.5847 | 4.351 | 6.7829 | 7.1407 | 1.1912 | 0.22216 | 0.34634 | 0.3646 | 0.060825 | 1.5589 | 1.6412 | 0.27379 | 1.0527 | 0.17563 | 0.16683 | -1.5044 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.733 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.3% | -14.6271 | -5.09 | -20.0219 | [8 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 104.3% | -0.38187 | -6.5961 | -12.7188 | iRM_12 | [6 empty] | SRM 3169 avg |  |  |  |  | 0.13678 | 1.2564 | 0.016957 | 1.0431 | 0.062763 | 3.6758 |  | 14.8 | 14.8
Lab 1 | Apr 5, 2022 | 3 | Zr_030 | iRM_10 |  |  | 0.00042288 | 9.23e-05 | 0.00013372 | 2.1802e-05 | 0.0001273 | -1.2598e-06 | 7.6795e-05 | -1.1817e-05 | 3.5809e-06 | [5 empty] | 20.5426 | 4.5639 | 7.115 | 6.5068e-05 | 7.4907 | 3.3706e-05 | 1.2498 | 1.257e-06 | 1.3592e-06 |  |  |  | 1 | 0 | 0 | 20.5422 | 4.5638 | 7.1148 | 4.3266e-05 | 7.4906 | 3.4966e-05 | 1.2497 | 1.3074e-05 | -2.2217e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7355 | 20.5422 | 4.5638 | 7.1148 | 7.4906 | 1.2497 | 0.22217 | 0.34635 | 0.36464 | 0.060836 | 1.559 | 1.6413 | 0.27383 | 1.0528 | 0.17565 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7996 | [10 empty] | 0 | 1.7355 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014077 | -1.8133 | -1.8273 | [14 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 105% | -2.7296 | -5.6711 | 16.6301 | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 106.5% | -3.5822 | -2.7739 | -5.2892 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_035 | iRM_12 |  |  | 0.00037883 | 7.6459e-05 | 0.00011089 | 2.0727e-05 | 0.00010793 | -2.1635e-06 | 0.00012902 | -1.4477e-05 | 3.563e-06 | [5 empty] | 19.5254 | 4.3378 | 6.7624 | 6.1487e-05 | 7.1191 | 2.9769e-05 | 1.1878 | -2.5595e-06 | -2.5524e-08 |  |  |  | 1 | 0 | 0 | 19.525 | 4.3378 | 6.7623 | 4.076e-05 | 7.119 | 3.1933e-05 | 1.1877 | 1.1917e-05 | -3.5885e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7335 | 19.525 | 4.3378 | 6.7623 | 7.119 | 1.1877 | 0.22216 | 0.34634 | 0.36461 | 0.060828 | 1.5589 | 1.6412 | 0.2738 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7335 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.57% | 7.5591 | 0.42347 | -3.2582 | [8 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 105.6% | -8.2727 | -6.3618 | 16.8375 | iRM_12 | [6 empty] | Laboratory 3
Lab 1 | Apr 5, 2022 | 3 | Zr_040 | iRM_11 |  |  | 0.00039484 | 8.6021e-05 | 0.00012097 | 1.6059e-05 | 0.00011291 | 1.0931e-05 | 7.9204e-05 | -8.9038e-06 | 9.7652e-07 | [5 empty] | 20.5835 | 4.573 | 7.129 | 6.4859e-05 | 7.5054 | 4.4094e-05 | 1.2522 | 1.0116e-07 | -2.2088e-07 |  |  |  | 1 | 0 | 0 | 20.5831 | 4.5729 | 7.1289 | 4.88e-05 | 7.5053 | 3.3164e-05 | 1.2521 | 9.005e-06 | -1.1974e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7349 | 20.5831 | 4.5729 | 7.1289 | 7.5053 | 1.2521 | 0.22217 | 0.34635 | 0.36463 | 0.060833 | 1.559 | 1.6413 | 0.27382 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7996 | [10 empty] | 0 | 1.7349 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014076 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 105.1% | -5.1932 | -4.9331 | -11.9378 | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 106.8% | -8.5396 | 1.4557 | 2.9679 | iRM_12 |  |  |  |  | Lab 3 | Feb 4, 2020 | Bottle 4 |  | Feb | 1 | 18 | -1.9867 | 3.5636 | -1.0713 | 1.7335 | -0.37857 | 6.2539 | SRM 3169 | 5.6
Lab 1 | Apr 5, 2022 | 3 | Zr_045 | iRM_12 |  |  | 0.00037443 | 8.695e-05 | 0.00012229 | 1.2554e-05 | 0.00010632 | 9.9738e-06 | 6.8902e-05 | -9.7356e-06 | 3.774e-06 | [5 empty] | 19.6325 | 4.3616 | 6.7995 | 6.1105e-05 | 7.1582 | 4.1968e-05 | 1.1943 | -7.3734e-06 | -9.6716e-07 |  |  |  | 1 | 0 | 0 | 19.6321 | 4.3616 | 6.7994 | 4.8552e-05 | 7.1581 | 3.1994e-05 | 1.1942 | 2.3622e-06 | -4.7411e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7337 | 19.6321 | 4.3616 | 6.7994 | 7.1581 | 1.1942 | 0.22216 | 0.34634 | 0.36461 | 0.06083 | 1.5589 | 1.6412 | 0.27381 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7337 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.3% | -2.1986 | 0.6795 | 31.3167 | [21 empty] | Lab 3 | Jul 20, 2020 | Bottle 4 |  | Jul 1 | 1 | 18 | 1.4727 | 4.9678 | 0.65789 | 3.3328 | -11.7489 | 9.7838 | SRM 3169 | 5.6
Lab 1 | Apr 5, 2022 | 3 | Zr_050 | iRM |  |  | 0.00043985 | 9.2159e-05 | 0.000135 | 3.3926e-06 | 0.00013134 | -2.5436e-06 | 8.5388e-05 | -5.5722e-06 | 2.4546e-06 | [5 empty] | 19.4031 | 4.3107 | 6.72 | 5.7829e-05 | 7.0746 | 2.9554e-05 | 1.1803 | -1.2151e-05 | 2.8727e-08 |  |  |  | 1 | 0 | 0 | 19.4026 | 4.3106 | 6.7199 | 5.4436e-05 | 7.0745 | 3.2098e-05 | 1.1802 | -6.5785e-06 | -2.4258e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7337 | 19.4026 | 4.3106 | 6.7199 | 7.0745 | 1.1802 | 0.22217 | 0.34634 | 0.36461 | 0.060829 | 1.5589 | 1.6412 | 0.2738 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7337 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014077 | -1.8133 | -1.8273 | [14 empty] | iRM | Lab 1 | Apr 5, 2022 | 3 | 98.83% | 4.6161 | -4.9887 | 2.5602 | iRM_2 | Lab 1 | Apr 4, 2022 | 1 | 98.76% | -8.4493 | -1.2719 | -7.3659 | iRM | Avg | -0.94856 | -2.3907 | -12.749 | Lab 3 | Jul 30, 2020 | Bottle 4 |  | Jul 2 | 5 | 5 | -2.9774 | 4.6533 | 0.87742 | 2.9456 | 0.72423 | 22.9481 | SRM 3169 | 5.6
Lab 1 | Apr 5, 2022 | 3 | Zr_055 | iRM_12 |  |  | 0.00037969 | 8.7319e-05 | 0.00011815 | 8.8916e-06 | 9.7275e-05 | -5.2085e-06 | 7.7282e-05 | -1.1205e-05 | -1.1895e-06 | [5 empty] | 19.6326 | 4.3617 | 6.7995 | 5.4418e-05 | 7.1582 | 2.9443e-05 | 1.1942 | -5.3127e-06 | -6.8058e-07 |  |  |  | 1 | 0 | 0 | 19.6322 | 4.3616 | 6.7994 | 4.5527e-05 | 7.1581 | 3.4652e-05 | 1.1942 | 5.8922e-06 | 5.0889e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7335 | 19.6322 | 4.3616 | 6.7994 | 7.1581 | 1.1942 | 0.22216 | 0.34634 | 0.36461 | 0.060826 | 1.5589 | 1.6412 | 0.27379 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7335 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3883 | 0.014073 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.84% | -1.3618 | -4.2492 | -28.2706 | [8 empty] | iRM_2 | Lab 1 | Apr 4, 2022 | 1 | 101% | -11.8787 | -2.203 | -20.6067 | iRM | 2SD | 14.5976 | 9.5166 | 21.5138 |  |  | Bottle 4 avg |  |  |  |  | -1.416 | 2.4585 | -0.36459 | 1.3633 | -3.4563 | 5.1357 |  | 5.6 | 5.6
Lab 1 | Apr 5, 2022 | 3 | Zr_060 | iRM_1 |  |  | 0.00037784 | 8.7899e-05 | 0.00011971 | 9.4288e-06 | 0.00011798 | -3.8858e-07 | 6.8014e-05 | -3.9402e-08 | 6.7644e-07 | [5 empty] | 20.7865 | 4.6181 | 7.1994 | 6.7121e-05 | 7.5795 | 3.2311e-05 | 1.2646 | -2.2489e-06 | 1.5313e-06 |  |  |  | 1 | 0 | 0 | 20.7861 | 4.618 | 7.1993 | 5.7692e-05 | 7.5794 | 3.2699e-05 | 1.2645 | -2.2095e-06 | 8.549e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7352 | 20.7861 | 4.618 | 7.1993 | 7.5794 | 1.2645 | 0.22217 | 0.34635 | 0.36464 | 0.060835 | 1.559 | 1.6413 | 0.27383 | 1.0528 | 0.17565 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7996 | [10 empty] | 0 | 1.7352 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [14 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 105.7% | -6.5697 | -4.9003 | 25.3421 | iRM_2 | Lab 1 | Apr 5, 2022 | 2 | 100.4% | 10.077 | 3.23 | -3.9028 | iRM_12 | n | 8 | 8 | 8 | Lab 3 | Feb 4, 2020 | Bottle 5 |  | Feb | 1 | 18 | -0.64855 | 2.6938 | -0.60179 | 2.347 | -2.3274 | 8.49 | SRM 3169 | 7.2
Lab 1 | Apr 5, 2022 | 3 | Zr_065 | iRM_12 |  |  | 0.00039515 | 9.2709e-05 | 0.00012476 | 1.3796e-05 | 0.00011762 | -9.1886e-07 | 6.9948e-05 | -1.0407e-06 | 2.84e-06 | [5 empty] | 19.6944 | 4.3754 | 6.8208 | 5.6943e-05 | 7.1804 | 2.8953e-05 | 1.1979 | -1.5051e-06 | 2.1972e-07 |  |  |  | 1 | 0 | 0 | 19.694 | 4.3753 | 6.8206 | 4.3147e-05 | 7.1802 | 2.9872e-05 | 1.1978 | -4.6439e-07 | -2.6203e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7321 | 19.694 | 4.3753 | 6.8206 | 7.1802 | 1.1978 | 0.22216 | 0.34633 | 0.36459 | 0.060822 | 1.5589 | 1.6411 | 0.27377 | 1.0527 | 0.17562 | 0.16682 | -1.5043 | -1.0604 | -1.009 | -2.7998 | [10 empty] | 0 | 1.7321 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014066 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.4% | 1.7773 | 5.6029 | 0.85659 | [8 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 2 | 101.2% | -7.3462 | 0.12981 | -32.2272 | iRM_12 | 2SE | 5.161 | 3.3646 | 7.6063 | Lab 3 | Jul 20, 2020 | Bottle 5 |  | Jul 1 | 1 | 16 | -0.019457 | 5.2236 | 0.8487 | 3.9771 | 10.6151 | 5.2605 | SRM 3169 | 7.2
Lab 1 | Apr 5, 2022 | 3 | Zr_070 | iRM_2 |  |  | 0.00038841 | 8.3283e-05 | 0.00011837 | 1.7521e-05 | 0.00011791 | -4.598e-06 | 8.456e-05 | -4.1027e-06 | 3.1387e-06 | [5 empty] | 20.3093 | 4.512 | 7.034 | 6.3027e-05 | 7.4051 | 3.0259e-05 | 1.2354 | 1.0673e-06 | -1.0476e-06 |  |  |  | 1 | 0 | 0 | 20.3089 | 4.512 | 7.0339 | 4.5505e-05 | 7.405 | 3.4857e-05 | 1.2354 | 5.17e-06 | -4.1863e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7339 | 20.3089 | 4.512 | 7.0339 | 7.405 | 1.2354 | 0.22217 | 0.34634 | 0.36462 | 0.060829 | 1.5589 | 1.6412 | 0.2738 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.734 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014067 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 103.4% | -0.20099 | 1.8576 | -11.5534 | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 101.7% | 3.5542 | -10.6569 | -27.0886 | iRM_12 | Not included as no bracketing standard after this analysis |  |  |  | Lab 3 | Jul 30, 2020 | Bottle 5 |  | Jul 2 | 5 | 5 | 2.6242 | 4.2693 | -0.77758 | 3.5138 | 0.98632 | 19.4181 | SRM 3169 | 7.2
Lab 1 | Apr 5, 2022 | 3 | Zr_075 | iRM_12 |  |  | 0.00038118 | 7.367e-05 | 0.0001225 | 1.2412e-05 | 9.7637e-05 | -6.0195e-07 | 8.5594e-05 | -8.4363e-06 | 3.3419e-06 | [5 empty] | 19.5831 | 4.3506 | 6.7823 | 5.4289e-05 | 7.1399 | 2.8597e-05 | 1.1912 | -6.888e-06 | -2.4618e-07 |  |  |  | 1 | 0 | 0 | 19.5827 | 4.3506 | 6.7821 | 4.1877e-05 | 7.1398 | 2.9199e-05 | 1.1911 | 1.5484e-06 | -3.5881e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7328 | 19.5827 | 4.3506 | 6.7821 | 7.1398 | 1.1911 | 0.22216 | 0.34633 | 0.3646 | 0.060825 | 1.5589 | 1.6411 | 0.27379 | 1.0527 | 0.17563 | 0.16683 | -1.5043 | -1.0604 | -1.009 | -2.7998 | [10 empty] | 0 | 1.7328 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.21% | 4.1596 | -3.8485 | 8.886 | [8 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 103.4% | -0.20099 | 1.8576 | -11.5534 | iRM_12 | [6 empty] | Bottle 5 avg |  |  |  |  | 0.23497 | 2.0882 | -0.36398 | 1.7521 | 6.7206 | 4.3577 |  | 7.2 | 7.2
Lab 1 | Apr 5, 2022 | 3 | Zr_080 | iRM_10 |  |  | 0.00047304 | 9.7038e-05 | 0.00014647 | 1.2265e-05 | 0.00013243 | 9.7347e-07 | 9.0851e-05 | -9.6481e-06 | -2.224e-06 | [5 empty] | 20.873 | 4.6372 | 7.2292 | 5.839e-05 | 7.6106 | 3.0866e-05 | 1.2697 | 2.3065e-07 | 2.5628e-07 |  |  |  | 1 | 0 | 0 | 20.8725 | 4.6371 | 7.229 | 4.6125e-05 | 7.6105 | 2.9892e-05 | 1.2696 | 9.8788e-06 | 2.4803e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.734 | 20.8725 | 4.6371 | 7.229 | 7.6105 | 1.2696 | 0.22217 | 0.34634 | 0.36462 | 0.060828 | 1.5589 | 1.6412 | 0.2738 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7339 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014072 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 106% | -0.73227 | -2.2709 | -16.6974 | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 104% | 4.8817 | -1.8488 | -5.3646 | iRM_12 |  |  |  |  | Lab 3 | Feb 4, 2020 | Bottle 6 |  | Feb | 1 | 18 | -1.0744 | 2.2877 | -1.2983 | 1.8659 | -0.69774 | 5.7077 | SRM 3169 | 8.8
Lab 1 | Apr 5, 2022 | 3 | Zr_085 | iRM_12 |  |  | 0.00041218 | 9.1001e-05 | 0.00014027 | 9.5488e-06 | 0.00011193 | -4.7177e-06 | 7.7366e-05 | -1.8723e-05 | -1.5797e-06 | [5 empty] | 19.7826 | 4.395 | 6.8515 | 5.4165e-05 | 7.2128 | 3.042e-05 | 1.2034 | -5.4382e-06 | 2.9133e-07 |  |  |  | 1 | 0 | 0 | 19.7822 | 4.3949 | 6.8513 | 4.4616e-05 | 7.2127 | 3.5138e-05 | 1.2033 | 1.3285e-05 | 1.871e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7332 | 19.7822 | 4.3949 | 6.8513 | 7.2127 | 1.2033 | 0.22216 | 0.34634 | 0.36461 | 0.060826 | 1.5589 | 1.6412 | 0.27379 | 1.0527 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7332 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.5% | -3.7754 | 1.784 | 0.87042 | [8 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 103.7% | 0.67668 | -2.4634 | -6.4825 | iRM_12 |  |  |  |  | Lab 3 | Jul 21, 2020 | Bottle 6 |  | Jul 1 | 2 | 15 | -5.407 | 4.9302 | 2.0403 | 4.6417 | 1.884 | 13.0087 | SRM 3169 | 8.8
Lab 1 | Apr 5, 2022 | 3 | Zr_090 | iRM_11 |  |  | 0.0003941 | 6.9236e-05 | 0.00012247 | 1.0602e-05 | 0.00010051 | -3.097e-06 | 8.349e-05 | -6.6253e-06 | 5.0614e-06 | [5 empty] | 20.8592 | 4.6342 | 7.2246 | 6.1856e-05 | 7.606 | 2.9985e-05 | 1.269 | -9.0306e-07 | -1.7068e-07 |  |  |  | 1 | 0 | 0 | 20.8588 | 4.6342 | 7.2245 | 5.1254e-05 | 7.6059 | 3.3082e-05 | 1.269 | 5.7222e-06 | -5.2321e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7353 | 20.8588 | 4.6342 | 7.2245 | 7.6059 | 1.269 | 0.22217 | 0.34635 | 0.36464 | 0.060836 | 1.559 | 1.6413 | 0.27383 | 1.0528 | 0.17565 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7996 | [10 empty] | 0 | 1.7353 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014066 | -1.8133 | -1.8273 | [14 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 105.4% | 6.2903 | 2.0221 | 22.3022 | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 104.7% | -10.7815 | -8.1017 | -8.007 | iRM_12 |  |  |  |  | Lab 3 | Jul 29, 2020 | Bottle 6 |  | Jul 2 | 4 | 9 | 0.58319 | 4.7603 | -2.3577 | 4.6863 | -3.7591 | 11.1621 | SRM 3169 | 8.8
Lab 1 | Apr 5, 2022 | 3 | Zr_095 | iRM_12 |  |  | 0.00047295 | 9.9351e-05 | 0.0001411 | 1.6732e-05 | 0.0001418 | -1.9187e-06 | 8.3681e-05 | -5.3504e-06 | 4.9801e-07 | [5 empty] | 19.7961 | 4.398 | 6.8561 | 5.765e-05 | 7.2178 | 2.9e-05 | 1.2042 | -7.3864e-06 | -1.6843e-08 |  |  |  | 1 | 0 | 0 | 19.7957 | 4.3979 | 6.856 | 4.0918e-05 | 7.2177 | 3.0918e-05 | 1.2041 | -2.036e-06 | -5.1485e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7334 | 19.7957 | 4.3979 | 6.856 | 7.2177 | 1.2041 | 0.22216 | 0.34634 | 0.36461 | 0.060826 | 1.5589 | 1.6412 | 0.27379 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7334 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.4% | -3.1975 | 0.91127 | -11.3724 | [23 empty] | Bottle 6 avg |  |  |  |  | -1.4547 | 1.9023 | -1.0169 | 1.624 | -0.90645 | 4.7335 |  | 8.8 | 8.8
Lab 1 | Apr 5, 2022 | 3 | Zr_100 | iRM |  |  | 0.00045118 | 9.6329e-05 | 0.00013939 | 1.557e-05 | 0.00013682 | 2.8928e-06 | 7.6069e-05 | -1.3811e-05 | 5.658e-06 | [5 empty] | 19.6986 | 4.3764 | 6.8226 | 5.8477e-05 | 7.1827 | 2.6183e-05 | 1.1984 | -9.0977e-06 | 1.8479e-06 |  |  |  | 1 | 0 | 0 | 19.6982 | 4.3763 | 6.8224 | 4.2907e-05 | 7.1826 | 2.329e-05 | 1.1983 | 4.7129e-06 | -3.8101e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7349 | 19.6982 | 4.3763 | 6.8224 | 7.1826 | 1.1983 | 0.22217 | 0.34635 | 0.36463 | 0.060833 | 1.559 | 1.6412 | 0.27381 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7996 | [10 empty] | 0 | 1.7349 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014072 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Apr 5, 2022 | 3 | 99.87% | 4.1551 | -2.6105 | 1.9655 | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 100.6% | -8.2098 | -4.438 | -7.7383 | iRM_12 | Avg | -1.1346 | -0.91402 | -0.32674 | Lab 3 | Jul 21, 2020 | Bottle 10 |  | Jul 1 | 2 | 14 | -0.011466 | 2.9643 | 2.742 | 6.1044 | 0.18722 | 9.749 | SRM 3169 | 10.4
Lab 1 | Apr 5, 2022 | 3 | Zr_105 | iRM_12 |  |  | 0.00058357 | 0.0001323 | 0.00019074 | 1.2529e-05 | 0.00017566 | -1.4505e-06 | 8.7384e-05 | -1.0191e-05 | 3.111e-06 | [5 empty] | 19.6544 | 4.3665 | 6.807 | 5.4781e-05 | 7.1661 | 2.8191e-05 | 1.1956 | -8.6855e-06 | -4.4678e-07 |  |  |  | 1 | 0 | 0 | 19.6538 | 4.3664 | 6.8068 | 4.2251e-05 | 7.1659 | 2.9642e-05 | 1.1955 | 1.5051e-06 | -3.5578e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7333 | 19.6538 | 4.3664 | 6.8068 | 7.1659 | 1.1955 | 0.22216 | 0.34634 | 0.36461 | 0.060827 | 1.5589 | 1.6412 | 0.27379 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7333 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.01407 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.46% | 7.0653 | 0.56432 | 3.4118 | [8 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 102.2% | -0.0045312 | 2.2875 | -0.99852 | iRM_12 | 2SD | 8.3626 | 6.0148 | 17.1058 | Lab 3 | Jul 30, 2020 | Bottle 10 |  | Jul 2 | 5 | 9 | -0.13569 | 5.2178 | 0.69467 | 3.8305 | -4.7363 | 16.0499 | SRM 3169 | 10.4
Lab 1 | Apr 5, 2022 | 3 | Zr_110 | iRM_1 |  |  | 0.00041866 | 9.9961e-05 | 0.00014102 | 9.2306e-06 | 0.00013326 | -1.5481e-06 | 7.9942e-05 | -5.8635e-06 | -5.0069e-07 | [5 empty] | 21.2575 | 4.7228 | 7.3628 | 6.3121e-05 | 7.7516 | 3.0641e-05 | 1.2934 | -2.2247e-07 | -3.0142e-07 |  |  |  | 1 | 0 | 0 | 21.2571 | 4.7227 | 7.3626 | 5.389e-05 | 7.7515 | 3.2189e-05 | 1.2933 | 5.641e-06 | 1.9927e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7363 | 21.2571 | 4.7227 | 7.3626 | 7.7515 | 1.2933 | 0.22217 | 0.34636 | 0.36466 | 0.06084 | 1.559 | 1.6413 | 0.27384 | 1.0528 | 0.17565 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7995 | [10 empty] | 0 | 1.7363 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014066 | -1.8133 | -1.8273 | [14 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 108% | 4.3646 | 5.8517 | 16.4011 | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 103.1% | 1.7339 | 2.0491 | 10.329 | iRM_12 | n | 9 | 9 | 9 |  |  | Bottle 10 avg |  |  |  |  | -0.041776 | 2.5774 | 1.2731 | 3.2446 | -1.1397 | 8.3323 |  | 10.4 | 10.4
Lab 1 | Apr 5, 2022 | 3 | Zr_115 | iRM_12 |  |  | 0.00039817 | 8.5941e-05 | 0.00011876 | 1.2493e-05 | 0.00011716 | -2.6294e-06 | 7.7306e-05 | -9.2552e-06 | 3.2749e-06 | [5 empty] | 19.7264 | 4.3825 | 6.8321 | 5.1847e-05 | 7.1927 | 2.6203e-05 | 1.2 | 2.883e-07 | -1.0067e-06 |  |  |  | 1 | 0 | 0 | 19.726 | 4.3824 | 6.832 | 3.9353e-05 | 7.1925 | 2.8832e-05 | 1.2 | 9.5435e-06 | -4.2816e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7343 | 19.726 | 4.3824 | 6.832 | 7.1925 | 1.2 | 0.22216 | 0.34634 | 0.36462 | 0.060831 | 1.559 | 1.6412 | 0.27381 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7343 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014074 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100% | -6.2409 | -6.1191 | -1.2318 | [8 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 103.3% | 0.30847 | -0.8 | 8.1277 | iRM_12 | 2SE | 2.7875 | 2.0049 | 5.7019 | Lab 3 | Jul 21, 2020 | Bottle 11 |  | Jul 1 | 2 | 16 | -3.1857 | 3.6302 | 1.839 | 4.6204 | -7.5664 | 11.1538 | SRM 3169 | 12
Lab 1 | Apr 5, 2022 | 3 | Zr_120 | iRM_2 |  |  | 0.00040041 | 0.00010053 | 0.00013813 | 9.7709e-06 | 0.00011632 | -3.779e-06 | 7.8911e-05 | -4.3398e-06 | 7.7311e-07 | [5 empty] | 20.5469 | 4.5648 | 7.1163 | 5.7926e-05 | 7.4919 | 2.6651e-05 | 1.25 | 8.0721e-08 | 1.4555e-07 |  |  |  | 1 | 0 | 0 | 20.5465 | 4.5647 | 7.1162 | 4.8155e-05 | 7.4918 | 3.043e-05 | 1.2499 | 4.4205e-06 | -6.2756e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7345 | 20.5465 | 4.5647 | 7.1162 | 7.4918 | 1.2499 | 0.22217 | 0.34635 | 0.36463 | 0.060832 | 1.559 | 1.6412 | 0.27381 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7996 | [10 empty] | 0 | 1.7345 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [14 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 104% | 4.8817 | -1.8488 | -5.3646 | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 104.9% | 1.8058 | -5.8402 | 6.7952 | iRM_12 |  |  |  |  | Lab 3 | Jul 29, 2020 | Bottle 11 |  | Jul 2 | 4 | 9 | 2.5994 | 5.5941 | -0.82938 | 3.9451 | 3.5411 | 10.8555 | SRM 3169 | 12
Lab 1 | Apr 5, 2022 | 3 | Zr_125 | iRM_12 |  |  | 0.00045226 | 9.3352e-05 | 0.00013723 | 7.7959e-06 | 0.00013233 | -5.3688e-08 | 8.2046e-05 | -1.1322e-05 | -4.7326e-07 | [5 empty] | 19.7827 | 4.395 | 6.8517 | 6.167e-05 | 7.2133 | 2.7456e-05 | 1.2035 | -5.0347e-06 | 1.4015e-06 |  |  |  | 1 | 0 | 0 | 19.7822 | 4.3949 | 6.8516 | 5.3874e-05 | 7.2132 | 2.751e-05 | 1.2034 | 6.2868e-06 | 1.8748e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7347 | 19.7822 | 4.3949 | 6.8516 | 7.2132 | 1.2034 | 0.22217 | 0.34635 | 0.36463 | 0.060833 | 1.559 | 1.6412 | 0.27382 | 1.0528 | 0.17564 | 0.16684 | -1.5043 | -1.0603 | -1.0089 | -2.7996 | [10 empty] | 0 | 1.7347 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014065 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.56% | -2.8011 | 7.3851 | 14.011 | [8 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 106.4% | 3.0505 | 1.7192 | -14.1487 | iRM_12 | [6 empty] | Bottle 11 avg |  |  |  |  | -1.4714 | 3.0452 | 0.29578 | 3.0002 | -1.8622 | 7.7793 |  | 12 | 12
Lab 1 | Apr 5, 2022 | 3 | Zr_130 | iRM_10 |  |  | 0.0069936 | 0.0015687 | 0.0024117 | 2.2579e-05 | 0.0025216 | -2.8767e-06 | 0.0005222 | -1.1516e-05 | 2.3051e-06 | [5 empty] | 21.0276 | 4.6717 | 7.2832 | 6.4192e-05 | 7.6679 | 2.8536e-05 | 1.2794 | -3.0357e-06 | -6.8479e-07 |  |  |  | 1 | 0 | 0 | 21.0207 | 4.6702 | 7.2808 | 4.1613e-05 | 7.6654 | 3.1413e-05 | 1.2789 | 8.4803e-06 | -2.9899e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7367 | 21.0207 | 4.6702 | 7.2808 | 7.6654 | 1.2789 | 0.22217 | 0.34636 | 0.36466 | 0.060839 | 1.559 | 1.6414 | 0.27384 | 1.0528 | 0.17565 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7995 | [10 empty] | 0 | 1.7367 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.01407 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 105.6% | -6.2028 | -2.1299 | -33.8211 | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 106.9% | 0.43074 | -0.8313 | -7.8614 | iRM_12 |  |  |  |  | Lab 3 | Feb 4, 2020 | SRM 3169 | 130920 | Feb | 1 | 52 | 0.023857 | 1.2723 | 0.11918 | 1.7311 | 0.54157 | 6.0718 | SRM 3169 | 15.2
Lab 1 | Apr 5, 2022 | 3 | Zr_135 | iRM_12 |  |  | 0.00041199 | 8.6161e-05 | 0.00014111 | 1.4802e-05 | 0.00012821 | 4.0282e-07 | 8.3866e-05 | -1.2971e-05 | -1.0965e-06 | [5 empty] | 20.0121 | 4.446 | 6.9311 | 5.9451e-05 | 7.297 | 2.8255e-05 | 1.2174 | -8.4538e-07 | -1.1338e-06 |  |  |  | 1 | 0 | 0 | 20.0117 | 4.446 | 6.931 | 4.465e-05 | 7.2968 | 2.7853e-05 | 1.2173 | 1.2126e-05 | -3.7316e-08 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7346 | 20.0117 | 4.446 | 6.931 | 7.2968 | 1.2173 | 0.22217 | 0.34635 | 0.36463 | 0.060832 | 1.5589 | 1.6412 | 0.27381 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7996 | [10 empty] | 0 | 1.7346 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 101.1% | 11.1635 | -3.3328 | -12.2254 | [8 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 105.9% | -1.1484 | 0.9976 | 5.8119 | iRM_12 |  |  |  |  | Lab 3 | Jul 20, 2020 | SRM 3169 | 130920 | Jul 1 | 1 | 77 | 0.21049 | 2.4973 | -0.11461 | 1.7354 | -0.85194 | 6.071 | SRM 3169 | 15.2
Lab 1 | Apr 5, 2022 | 3 | Zr_140 | iRM_11 |  |  | 0.00040977 | 8.9727e-05 | 0.00013423 | 1.8977e-05 | 0.00012025 | -2.0211e-06 | 7.7453e-05 | 6.9035e-06 | -2.4582e-06 | [5 empty] | 21.1169 | 4.6916 | 7.3142 | 5.6333e-05 | 7.7008 | 2.9296e-05 | 1.2849 | 3.1151e-06 | -2.1333e-06 |  |  |  | 1 | 0 | 0 | 21.1165 | 4.6915 | 7.3141 | 3.7356e-05 | 7.7006 | 3.1317e-05 | 1.2848 | -3.7884e-06 | 3.2492e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7374 | 21.1165 | 4.6915 | 7.3141 | 7.7006 | 1.2848 | 0.22217 | 0.34637 | 0.36467 | 0.060844 | 1.559 | 1.6414 | 0.27386 | 1.0529 | 0.17566 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7994 | [10 empty] | 0 | 1.7374 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014073 | -1.8133 | -1.8273 | [14 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 106% | -4.3284 | -2.7454 | 5.4767 | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 108.1% | -8.1778 | -3.37 | -3.2575 | iRM_12 |  |  |  |  | Lab 3 | Jul 29, 2020 | SRM 3169 | 130920 | Jul 2 | 4 | 35 | 0.52064 | 3.5518 | 0.14497 | 2.2192 | 0.35594 | 9.9655 | SRM 3169 | 15.2
Lab 1 | Apr 5, 2022 | 3 | Zr_145 | iRM_12 |  |  | 0.00041476 | 9.7409e-05 | 0.00014136 | 9.6239e-06 | 0.00011995 | -2.8471e-06 | 7.5836e-05 | -8.1394e-06 | -9.9993e-07 | [5 empty] | 19.8255 | 4.4046 | 6.8665 | 5.0134e-05 | 7.229 | 2.6179e-05 | 1.2061 | -3.9084e-06 | -4.0524e-07 |  |  |  | 1 | 0 | 0 | 19.8251 | 4.4045 | 6.8664 | 4.051e-05 | 7.2289 | 2.9027e-05 | 1.206 | 4.2309e-06 | 5.9469e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7348 | 19.8251 | 4.4045 | 6.8664 | 7.2289 | 1.206 | 0.22217 | 0.34635 | 0.36463 | 0.060833 | 1.559 | 1.6413 | 0.27382 | 1.0528 | 0.17564 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7996 | [10 empty] | 0 | 1.7348 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.01407 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.71% | -8.2804 | 0.2123 | 2.2107 | [23 empty] | SRM 3169 avg |  |  |  |  | 0.10469 | 1.08 | 0.035851 | 1.0729 | -0.075257 | 3.9428 |  | 15.2 | 15.2
Lab 1 | Apr 5, 2022 | 3 | Zr_150 | iRM |  |  | 0.00041003 | 9.1159e-05 | 0.00012756 | 2.1261e-05 | 0.0001095 | 1.5051e-06 | 7.3224e-05 | -3.2393e-06 | 1.6187e-06 | [5 empty] | 19.9881 | 4.4408 | 6.9232 | 5.9935e-05 | 7.289 | 2.8468e-05 | 1.2162 | -2.4971e-06 | 4.2144e-07 |  |  |  | 1 | 0 | 0 | 19.9877 | 4.4407 | 6.9231 | 3.8674e-05 | 7.2889 | 2.6963e-05 | 1.2161 | 7.4228e-07 | -1.1973e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7373 | 19.9877 | 4.4407 | 6.9231 | 7.2889 | 1.2161 | 0.22217 | 0.34637 | 0.36467 | 0.060841 | 1.559 | 1.6414 | 0.27385 | 1.0528 | 0.17566 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7995 | [10 empty] | 0 | 1.7373 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014072 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Apr 5, 2022 | 3 | 101% | 1.9212 | -1.8048 | -26.3623 | iRM_10 | Lab 1 | Apr 4, 2022 | 1 | 98.37% | 10.8212 | 2.7856 | 11.362 | iRM | Avg | -0.31231 | -2.8536 | -3.143
Lab 1 | Apr 5, 2022 | 3 | Zr_155 | iRM_12 |  |  | 0.00041271 | 8.8218e-05 | 0.00012785 | 1.4355e-05 | 0.00011998 | 1.24e-06 | 7.2946e-05 | -2.4384e-06 | 6.2083e-06 | [5 empty] | 19.7554 | 4.3891 | 6.8426 | 5.8204e-05 | 7.2041 | 2.6504e-05 | 1.202 | 2.3708e-06 | 4.2766e-07 |  |  |  | 1 | 0 | 0 | 19.755 | 4.389 | 6.8424 | 4.3849e-05 | 7.2039 | 2.5264e-05 | 1.2019 | 4.8092e-06 | -5.7806e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7369 | 19.755 | 4.389 | 6.8424 | 7.204 | 1.2019 | 0.22217 | 0.34637 | 0.36467 | 0.060842 | 1.559 | 1.6414 | 0.27385 | 1.0528 | 0.17566 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7995 | [10 empty] | 0 | 1.7369 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.01407 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.27% | -0.043363 | -1.0138 | -2.9812 | [8 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 2 | 105% | 8.3578 | -4.1183 | 18.2493 | iRM_12 | 2SD | 14.2769 | 9.1554 | 35.8584 |  |  | Grand average and 2SD |  |  |  | Avg | -0.72283 |  | -0.027914 |  | 0.24959
Lab 1 | Apr 5, 2022 | 3 | Zr_160 | iRM_1 |  |  | 0.00041405 | 8.6099e-05 | 0.000141 | 1.4838e-05 | 0.0001217 | -3.3323e-06 | 7.2922e-05 | 4.5403e-06 | 5.629e-06 | [5 empty] | 21.5628 | 4.7907 | 7.4689 | 6.334e-05 | 7.8638 | 2.9148e-05 | 1.3121 | 1.0843e-06 | -1.0118e-06 |  |  |  | 1 | 0 | 0 | 21.5624 | 4.7906 | 7.4688 | 4.8502e-05 | 7.8637 | 3.248e-05 | 1.3121 | -3.456e-06 | -6.6408e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7388 | 21.5624 | 4.7906 | 7.4688 | 7.8637 | 1.3121 | 0.22218 | 0.34638 | 0.3647 | 0.06085 | 1.559 | 1.6415 | 0.27388 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7993 | [10 empty] | 0 | 1.7389 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014072 | -1.8133 | -1.8273 | [14 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 108.5% | -6.422 | -3.0818 | 15.4585 | iRM_10 | Lab 1 | Apr 5, 2022 | 2 | 104.9% | -10.6246 | 0.098244 | -5.1414 | iRM_12 | n | 8 | 8 | 8 | [6 empty] | 2SD | 2.3444 |  | 2.6058 |  | 7.1969
Lab 1 | Apr 5, 2022 | 3 | Zr_165 | iRM_12 |  |  | 0.00047889 | 0.00010423 | 0.00014571 | 1.31e-05 | 0.00014152 | -5.6768e-07 | 7.9536e-05 | -5.5814e-06 | 9.8092e-07 | [5 empty] | 19.9746 | 4.4379 | 6.9187 | 5.8506e-05 | 7.2845 | 2.6841e-05 | 1.2155 | 1.684e-06 | -1.2974e-06 |  |  |  | 1 | 0 | 0 | 19.9742 | 4.4378 | 6.9186 | 4.5406e-05 | 7.2843 | 2.7409e-05 | 1.2154 | 7.2654e-06 | -2.2783e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7383 | 19.9742 | 4.4378 | 6.9186 | 7.2843 | 1.2154 | 0.22218 | 0.34638 | 0.36469 | 0.060848 | 1.559 | 1.6414 | 0.27387 | 1.0529 | 0.17567 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7994 | [10 empty] | 0 | 1.7383 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.1% | 6.8239 | 2.8429 | 10.3542 | [8 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 105% | -2.7296 | -5.6711 | 16.6301 | iRM_12 | 2SE | 5.0476 | 3.2369 | 12.6778
Lab 1 | Apr 5, 2022 | 3 | Zr_170 | iRM_2 |  |  | 0.00048429 | 0.00010535 | 0.00016393 | 1.1686e-05 | 0.00014921 | 7.0957e-07 | 8.7644e-05 | -4.4115e-06 | 1.6866e-06 | [5 empty] | 20.8194 | 4.6255 | 7.2113 | 6.2449e-05 | 7.5925 | 2.7628e-05 | 1.2668 | 2.9363e-07 | -1.8093e-07 |  |  |  | 1 | 0 | 0 | 20.8189 | 4.6254 | 7.2111 | 5.0763e-05 | 7.5923 | 2.6918e-05 | 1.2667 | 4.7051e-06 | -1.8675e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7382 | 20.8189 | 4.6254 | 7.2111 | 7.5924 | 1.2667 | 0.22217 | 0.34637 | 0.36469 | 0.060846 | 1.559 | 1.6414 | 0.27387 | 1.0529 | 0.17567 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7994 | [10 empty] | 0 | 1.7382 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014072 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 103.7% | 0.67668 | -2.4634 | -6.4825 | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 106% | -0.73227 | -2.2709 | -16.6974 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_175 | iRM_12 |  |  | 0.00047347 | 0.00010324 | 0.00016116 | 1.5684e-05 | 0.00015092 | -2.0642e-06 | 8.1756e-05 | -1.546e-05 | -3.1888e-08 | [5 empty] | 20.1703 | 4.4813 | 6.9863 | 5.7467e-05 | 7.3555 | 2.6086e-05 | 1.2273 | -3.9943e-06 | 1.9579e-06 |  |  |  | 1 | 0 | 0 | 20.1698 | 4.4812 | 6.9862 | 4.1783e-05 | 7.3554 | 2.815e-05 | 1.2272 | 1.1466e-05 | 1.9898e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7374 | 20.1698 | 4.4812 | 6.9862 | 7.3554 | 1.2272 | 0.22217 | 0.34637 | 0.36467 | 0.060843 | 1.559 | 1.6414 | 0.27386 | 1.0528 | 0.17566 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7995 | [10 empty] | 0 | 1.7374 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 101.1% | -7.0288 | -2.8042 | -10.6751 | [8 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 105.6% | -6.2028 | -2.1299 | -33.8211 | iRM_12 |  |  |  |  | Digest of small Zr metal fragment from NIST. Not plotted to keep focus on RM 8299 and SRM 3169.
Lab 1 | Apr 5, 2022 | 3 | Zr_180 | iRM_10 |  |  | 0.00055567 | 0.00013775 | 0.00018043 | 1.0534e-05 | 0.00017239 | -1.2362e-06 | 8.994e-05 | -1.0862e-05 | 1.861e-06 | [5 empty] | 21.4153 | 4.7581 | 7.4182 | 6.102e-05 | 7.8108 | 2.9631e-05 | 1.3033 | -1.0233e-06 | 5.823e-07 |  |  |  | 1 | 0 | 0 | 21.4148 | 4.758 | 7.418 | 5.0486e-05 | 7.8107 | 3.0868e-05 | 1.3033 | 9.8387e-06 | -1.2787e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7411 | 21.4148 | 4.758 | 7.418 | 7.8107 | 1.3033 | 0.22218 | 0.3464 | 0.36473 | 0.060858 | 1.5591 | 1.6416 | 0.27391 | 1.0529 | 0.17569 | 0.16686 | -1.5043 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7411 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 106.9% | 1.5668 | 0.51845 | -4.8814 | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 106.9% | 1.5668 | 0.51845 | -4.8814 | iRM_12 |  |  |  |  | Lab 1 | Apr 4, 2022 | iRM |  | Apr | 1 | 12 | -0.04467 | 3.246 | -2.3401 | 2.2384 | -0.30986 | 13.5864 | iRM/iRM12
Lab 1 | Apr 5, 2022 | 3 | Zr_185 | iRM_12 |  |  | 0.00047029 | 9.9906e-05 | 0.00015128 | 1.1956e-05 | 0.00013547 | -4.5502e-06 | 7.0937e-05 | -8.1036e-06 | 4.2348e-06 | [5 empty] | 19.9091 | 4.4233 | 6.896 | 5.4542e-05 | 7.2604 | 2.4685e-05 | 1.2114 | -1.6936e-06 | 1.9187e-07 |  |  |  | 1 | 0 | 0 | 19.9086 | 4.4232 | 6.8958 | 4.2586e-05 | 7.2603 | 2.9235e-05 | 1.2113 | 6.41e-06 | -4.043e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7379 | 19.9086 | 4.4232 | 6.8958 | 7.2603 | 1.2113 | 0.22217 | 0.34637 | 0.36468 | 0.060845 | 1.559 | 1.6414 | 0.27386 | 1.0529 | 0.17566 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7994 | [10 empty] | 0 | 1.7379 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.32% | 4.2492 | 6.2924 | 11.2456 | [8 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 108.3% | -2.9551 | -12.0407 | -10.8439 | iRM_12 |  |  |  |  | Lab 1 | Jun 20, 2022 | iRM |  | Jun | 1 | 9 | 1.3509 | 3.0835 | -4.395 | 2.737 | -5.6578 | 5.8342 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_190 | iRM_11 |  |  | 0.0004744 | 0.00010536 | 0.00015036 | 1.0158e-05 | 0.00013193 | -1.4002e-06 | 7.2877e-05 | -6.892e-06 | 2.2016e-06 | [5 empty] | 21.5922 | 4.7974 | 7.4795 | 6.4004e-05 | 7.8756 | 2.9511e-05 | 1.3142 | -4.9572e-07 | 9.0862e-07 |  |  |  | 1 | 0 | 0 | 21.5917 | 4.7973 | 7.4794 | 5.3846e-05 | 7.8754 | 3.0911e-05 | 1.3141 | 6.3963e-06 | -1.2929e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7418 | 21.5917 | 4.7973 | 7.4794 | 7.8754 | 1.3141 | 0.22218 | 0.3464 | 0.36474 | 0.06086 | 1.5591 | 1.6416 | 0.27392 | 1.053 | 0.17569 | 0.16686 | -1.5043 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7418 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 108.4% | 3.2601 | -1.5368 | -0.24085
Lab 1 | Apr 5, 2022 | 3 | Zr_195 | iRM_12 |  |  | 0.00046849 | 0.00012107 | 0.00014373 | 1.8839e-05 | 0.00013951 | 9.8013e-07 | 8.2046e-05 | -1.4108e-05 | 3.8412e-06 | [5 empty] | 19.9197 | 4.4257 | 6.8998 | 5.6926e-05 | 7.2648 | 2.8115e-05 | 1.2122 | -5.3907e-06 | 5.1593e-07 |  |  |  | 1 | 0 | 0 | 19.9192 | 4.4256 | 6.8997 | 3.8087e-05 | 7.2647 | 2.7135e-05 | 1.2121 | 8.7176e-06 | -3.3253e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7395 | 19.9192 | 4.4256 | 6.8997 | 7.2647 | 1.2121 | 0.22218 | 0.34638 | 0.36471 | 0.060851 | 1.559 | 1.6415 | 0.27389 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7993 | [10 empty] | 0 | 1.7395 | 0.21794 | 0.33337 | 0.3381 | 0.054379 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014079 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.67% | -9.6411 | -10.4282 | -9.7697 | [8 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 100.6% | -8.1597 | -4.5443 | 2.9292 | iRM_12 | Avg | -0.92231 | -1.0418 | 2.1531
Lab 1 | Apr 5, 2022 | 3 | Zr_200 | iRM |  |  | 0.00047854 | 0.00011325 | 0.0001521 | 1.627e-05 | 0.00015761 | -7.7856e-08 | 7.0952e-05 | -7.9867e-06 | 1.7677e-06 | [5 empty] | 20.1775 | 4.4829 | 6.9891 | 5.7398e-05 | 7.3587 | 2.7932e-05 | 1.2278 | -1.2599e-06 | 2.2016e-06 |  |  |  | 1 | 0 | 0 | 20.177 | 4.4828 | 6.9889 | 4.1128e-05 | 7.3586 | 2.801e-05 | 1.2278 | 6.7268e-06 | 4.3389e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7391 | 20.177 | 4.4828 | 6.9889 | 7.3586 | 1.2278 | 0.22218 | 0.34638 | 0.3647 | 0.06085 | 1.559 | 1.6415 | 0.27388 | 1.0529 | 0.17567 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7993 | [10 empty] | 0 | 1.7391 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.01407 | -1.8133 | -1.8273 | [14 empty] | iRM | Lab 1 | Apr 5, 2022 | 3 | 100.9% | -8.715 | 3.1855 | 14.2855 | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 101% | -2.9173 | 0.54313 | 7.2901 | iRM_12 | 2SD | 10.3618 | 12.3588 | 35.0839
Lab 1 | Apr 5, 2022 | 3 | Zr_205 | iRM_12 |  |  | 0.0004825 | 0.0001079 | 0.0001541 | 2.1216e-05 | 0.00014635 | 1.1987e-05 | 8.5178e-05 | -1.7604e-06 | 4.4619e-06 | [5 empty] | 20.0613 | 4.4573 | 6.9491 | 5.3897e-05 | 7.3168 | 3.6871e-05 | 1.2209 | 1.7597e-06 | -2.8194e-07 |  |  |  | 1 | 0 | 0 | 20.0608 | 4.4572 | 6.9489 | 3.268e-05 | 7.3166 | 2.4884e-05 | 1.2208 | 3.5201e-06 | -4.7439e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7406 | 20.0608 | 4.4572 | 6.9489 | 7.3167 | 1.2208 | 0.22218 | 0.34639 | 0.36472 | 0.060856 | 1.559 | 1.6415 | 0.2739 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0086 | -2.7993 | [10 empty] | 0 | 1.7406 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.94% | 11.4185 | 6.0451 | -4.5778 | [8 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 101.5% | 1.5505 | 1.5702 | 38.3952 | iRM_12 | n | 9 | 9 | 9
Lab 1 | Apr 5, 2022 | 3 | Zr_210 | iRM_1 |  |  | 0.0004756 | 0.00010839 | 0.00015476 | 1.0143e-05 | 0.00015114 | -5.7928e-06 | 7.6974e-05 | -8.5773e-06 | 2.5886e-06 | [5 empty] | 21.6884 | 4.8188 | 7.5128 | 6.1971e-05 | 7.9105 | 3.0326e-05 | 1.32 | -5.1021e-06 | 8.7922e-07 |  |  |  | 1 | 0 | 0 | 21.6879 | 4.8187 | 7.5127 | 5.1828e-05 | 7.9104 | 3.6119e-05 | 1.3199 | 3.4752e-06 | -1.7093e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7414 | 21.6879 | 4.8187 | 7.5127 | 7.9104 | 1.3199 | 0.22218 | 0.3464 | 0.36474 | 0.06086 | 1.5591 | 1.6416 | 0.27392 | 1.0529 | 0.17569 | 0.16686 | -1.5043 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7414 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014074 | -1.8133 | -1.8273 | [14 empty] | iRM_1 | Lab 1 | Apr 5, 2022 | 3 | 107.7% | -7.0913 | -4.6331 | 12.6998 | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 102.5% | 3.9712 | 1.9916 | -26.0589 | iRM_12 | 2SE | 3.4539 | 4.1196 | 11.6946
Lab 1 | Apr 5, 2022 | 3 | Zr_215 | iRM_12 |  |  | 0.00048891 | 0.00011034 | 0.00014649 | 1.0717e-05 | 0.00014489 | -2.9286e-06 | 7.4204e-05 | -1.1579e-05 | 1.7048e-06 | [5 empty] | 20.2288 | 4.4944 | 7.0069 | 5.7357e-05 | 7.3776 | 2.5929e-05 | 1.231 | -5.402e-06 | 4.3078e-07 |  |  |  | 1 | 0 | 0 | 20.2283 | 4.4943 | 7.0068 | 4.664e-05 | 7.3774 | 2.8858e-05 | 1.2309 | 6.1772e-06 | -1.274e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7396 | 20.2283 | 4.4943 | 7.0068 | 7.3774 | 1.2309 | 0.22218 | 0.34639 | 0.36471 | 0.060852 | 1.559 | 1.6415 | 0.27389 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7993 | [10 empty] | 0 | 1.7396 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.01407 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 101% | -4.9752 | -0.1426 | 11.9754 | [8 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 104.5% | -0.062223 | -10.9723 | 4.0616 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_220 | iRM_2 |  |  | 0.0087965 | 0.0019549 | 0.0030433 | 1.3049e-05 | 0.0031847 | -4.3392e-06 | 0.00058648 | -3.0081e-06 | 1.8719e-06 | [5 empty] | 21.0687 | 4.6811 | 7.2982 | 5.9422e-05 | 7.6847 | 2.7166e-05 | 1.2823 | -3.2512e-07 | -4.2679e-07 |  |  |  | 1 | 0 | 0 | 21.0599 | 4.6791 | 7.2952 | 4.6373e-05 | 7.6815 | 3.1505e-05 | 1.2817 | 2.683e-06 | -2.2987e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.742 | 21.0599 | 4.6791 | 7.2952 | 7.6815 | 1.2817 | 0.22218 | 0.3464 | 0.36475 | 0.060861 | 1.5591 | 1.6417 | 0.27392 | 1.053 | 0.1757 | 0.16686 | -1.5043 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.742 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014078 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Apr 5, 2022 | 3 | 104.7% | -10.7815 | -8.1017 | -8.007 | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 103.9% | 7.3238 | -1.8952 | 10.3072 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_225 | iRM_12 |  |  | 0.00046409 | 0.00010681 | 0.00014286 | 1.8752e-05 | 0.00014462 | -6.3724e-07 | 7.6989e-05 | -5.3922e-06 | 2.7367e-06 | [5 empty] | 19.9838 | 4.44 | 6.9222 | 5.9652e-05 | 7.2885 | 2.4519e-05 | 1.2162 | 3.2278e-06 | -2.7408e-08 |  |  |  | 1 | 0 | 0 | 19.9834 | 4.4399 | 6.9221 | 4.09e-05 | 7.2884 | 2.5156e-05 | 1.2161 | 8.62e-06 | -2.7641e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7405 | 19.9834 | 4.4399 | 6.9221 | 7.2884 | 1.2161 | 0.22218 | 0.34639 | 0.36472 | 0.060856 | 1.5591 | 1.6416 | 0.2739 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0086 | -2.7993 | [10 empty] | 0 | 1.7405 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.09% | 3.6893 | -3.1906 | -1.2966 | [8 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 104.6% | -1.0247 | -7.6783 | -11.5869 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_230 | iRM_10 |  |  | 0.00048961 | 0.00011202 | 0.00016228 | 7.4327e-06 | 0.00015628 | -6.0262e-07 | 8.9727e-05 | -8.9161e-06 | 6.493e-07 | [5 empty] | 21.6993 | 4.8214 | 7.517 | 6.4017e-05 | 7.9155 | 2.9122e-05 | 1.3209 | 8.3181e-06 | -1.5548e-07 |  |  |  | 1 | 0 | 0 | 21.6988 | 4.8212 | 7.5168 | 5.6584e-05 | 7.9153 | 2.9724e-05 | 1.3208 | 1.7234e-05 | -8.0478e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7442 | 21.6988 | 4.8212 | 7.5168 | 7.9153 | 1.3208 | 0.22219 | 0.34642 | 0.36478 | 0.060869 | 1.5591 | 1.6418 | 0.27395 | 1.053 | 0.17571 | 0.16686 | -1.5042 | -1.0601 | -1.0085 | -2.799 | [10 empty] | 0 | 1.7442 | 0.21794 | 0.33337 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.01408 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Apr 5, 2022 | 3 | 108.3% | -2.9551 | -12.0407 | -10.8439 | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 106.2% | -0.36255 | 10.1373 | -5.6021 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_235 | iRM_12 |  |  | 0.00046355 | 0.00011392 | 0.000149 | 1.1872e-05 | 0.00014892 | -1.3006e-06 | 8.0872e-05 | -7.5064e-06 | 3.3923e-06 | [5 empty] | 20.1059 | 4.4671 | 6.9645 | 5.7099e-05 | 7.3329 | 2.5062e-05 | 1.2236 | 4.2768e-06 | 7.8894e-07 |  |  |  | 1 | 0 | 0 | 20.1054 | 4.467 | 6.9643 | 4.5227e-05 | 7.3327 | 2.6363e-05 | 1.2235 | 1.1783e-05 | -2.6033e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7401 | 20.1054 | 4.467 | 6.9643 | 7.3327 | 1.2235 | 0.22218 | 0.34639 | 0.36472 | 0.060853 | 1.5591 | 1.6415 | 0.27389 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0086 | -2.7993 | [10 empty] | 0 | 1.7401 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014066 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.77% | -0.3178 | 6.3687 | -8.0645 | [8 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 106.4% | -8.6197 | 1.4718 | -0.35788 | iRM_12
Lab 1 | Apr 5, 2022 | 3 | Zr_240 | iRM_11 |  |  | 0.00048508 | 0.00011022 | 0.00015006 | 7.7328e-06 | 0.00014955 | -1.5229e-06 | 8.3493e-05 | 2.246e-06 | 8.0786e-07 | [5 empty] | 21.7954 | 4.8426 | 7.55 | 5.5663e-05 | 7.9498 | 2.7068e-05 | 1.3266 | -4.9206e-07 | -8.3491e-07 |  |  |  | 1 | 0 | 0 | 21.795 | 4.8425 | 7.5498 | 4.793e-05 | 7.9496 | 2.8591e-05 | 1.3265 | -2.7381e-06 | -1.6428e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7419 | 21.795 | 4.8425 | 7.5498 | 7.9496 | 1.3265 | 0.22218 | 0.3464 | 0.36475 | 0.060862 | 1.5591 | 1.6416 | 0.27393 | 1.0529 | 0.1757 | 0.16686 | -1.5043 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7419 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [14 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 107.8% | -2.8828 | -1.436 | 11.3489
Lab 1 | Apr 5, 2022 | 3 | Zr_245 | iRM_12 |  |  | 0.00056573 | 0.00012202 | 0.00018567 | 1.5462e-05 | 0.00018009 | 6.3264e-07 | 8.6792e-05 | -9.9149e-06 | 2.9405e-07 | [5 empty] | 20.3189 | 4.5145 | 7.0385 | 5.531e-05 | 7.4111 | 2.5606e-05 | 1.2367 | 9.2766e-07 | -8.694e-07 |  |  |  | 1 | 0 | 0 | 20.3184 | 4.5144 | 7.0383 | 3.9848e-05 | 7.4109 | 2.4973e-05 | 1.2366 | 1.0843e-05 | -1.1634e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7417 | 20.3184 | 4.5144 | 7.0383 | 7.4109 | 1.2366 | 0.22218 | 0.3464 | 0.36474 | 0.06086 | 1.5591 | 1.6416 | 0.27392 | 1.0529 | 0.17569 | 0.16686 | -1.5043 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7417 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014073 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.5% | -5.5439 | -6.1824 | 20.3843 | [8 empty] | iRM_11 | Lab 1 | Apr 4, 2022 | 1 | 98.64% | 5.6088 | -1.2676 | 27.1619 | iRM | Avg | -0.13677 | -2.0315 | 8.7566
Lab 1 | Apr 5, 2022 | 3 | Zr_250 | iRM |  |  | 0.00048716 | 0.00010282 | 0.00016031 | 1.9715e-05 | 0.0001499 | -2.8792e-06 | 8.5217e-05 | -5.3052e-06 | -4.4619e-06 | [5 empty] | 20.4644 | 4.5468 | 7.0888 | 5.6839e-05 | 7.464 | 2.6072e-05 | 1.2455 | 3.5667e-06 | 5.2541e-07 |  |  |  | 1 | 0 | 0 | 20.4639 | 4.5467 | 7.0886 | 3.7124e-05 | 7.4639 | 2.8952e-05 | 1.2454 | 8.8719e-06 | 4.9873e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7413 | 20.4639 | 4.5467 | 7.0886 | 7.4639 | 1.2454 | 0.22218 | 0.3464 | 0.36473 | 0.060857 | 1.5591 | 1.6416 | 0.27391 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7413 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5514 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3883 | 0.014073 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Apr 5, 2022 | 3 | 100.7% | -0.88083 | -1.9562 | -12.0644 | iRM_11 | Lab 1 | Apr 5, 2022 | 2 | 104.4% | 0.44544 | -3.9419 | 9.5767 | iRM_12 | 2SD | 9.3857 | 4.1637 | 24.5569
Lab 1 | Apr 5, 2022 | 3 | Zr_255 | iRM_12 |  |  | 0.00049125 | 0.00011408 | 0.00016506 | 8.1502e-06 | 0.00014895 | -1.1466e-06 | 8.9031e-05 | -3.8838e-06 | 2.3523e-06 | [5 empty] | 20.3213 | 4.5149 | 7.0389 | 5.2248e-05 | 7.4111 | 2.346e-05 | 1.2366 | -1.3159e-07 | 3.6831e-07 |  |  |  | 1 | 0 | 0 | 20.3208 | 4.5148 | 7.0388 | 4.4097e-05 | 7.411 | 2.4606e-05 | 1.2365 | 3.7522e-06 | -1.984e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.739 | 20.3208 | 4.5148 | 7.0388 | 7.411 | 1.2365 | 0.22218 | 0.34638 | 0.3647 | 0.060848 | 1.559 | 1.6415 | 0.27387 | 1.0529 | 0.17567 | 0.16684 | -1.5043 | -1.0602 | -1.0087 | -2.7994 | [10 empty] | 0 | 1.739 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014069 | -1.8133 | -1.8274 | [22 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 2 | 104.7% | -4.2944 | -2.4135 | 6.3647 | iRM_12 | n | 8 | 8 | 8
Lab 1 | Jun 20, 2022 | 1 | Zr_005 | iRM_12 |  |  | 0.00023357 | 6.3928e-05 | 8.413e-05 | 6.3377e-05 | 8.7225e-05 | 2.1446e-06 | 0.00014844 | 8.2739e-07 | 7.3652e-07 | [5 empty] | 23.3358 | 5.1808 | 8.0708 | 0.00013421 | 8.4846 | 4.054e-05 | 1.4137 | 2.0007e-05 | 3.4302e-06 |  |  |  | 1 | 0 | 0 | 23.3356 | 5.1807 | 8.0707 | 7.0833e-05 | 8.4845 | 3.8396e-05 | 1.4135 | 1.918e-05 | 2.6937e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.669 | 23.3356 | 5.1807 | 8.0707 | 8.4845 | 1.4135 | 0.22201 | 0.34586 | 0.36359 | 0.060573 | 1.5578 | 1.6377 | 0.27284 | 1.0513 | 0.17514 | 0.1666 | -1.505 | -1.0617 | -1.0117 | -2.8039 | [9 empty] | L1 JunS1 | 0 | 1.669 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0141 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014048 | -1.8133 | -1.8274 | [22 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 105.1% | -5.1932 | -4.9331 | -11.9378 | iRM_12 | 2SE | 3.3184 | 1.4721 | 8.6822
Lab 1 | Jun 20, 2022 | 1 | Zr_010 | iRM_1 |  |  | 0.0002925 | 7.577e-05 | 0.0001008 | 6.6247e-05 | 9.9952e-05 | 4.7438e-06 | 0.00015455 | 1.1246e-05 | 5.4653e-07 | [5 empty] | 23.1672 | 5.1435 | 8.0127 | 0.00013538 | 8.4238 | 4.0021e-05 | 1.4036 | 1.7245e-05 | 5.7284e-06 |  |  |  | 1 | 0 | 0 | 23.1669 | 5.1434 | 8.0126 | 6.9134e-05 | 8.4237 | 3.5277e-05 | 1.4034 | 5.9991e-06 | 5.1819e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6703 | 23.1669 | 5.1434 | 8.0126 | 8.4237 | 1.4034 | 0.22202 | 0.34586 | 0.36361 | 0.060578 | 1.5578 | 1.6378 | 0.27285 | 1.0513 | 0.17515 | 0.1666 | -1.505 | -1.0617 | -1.0117 | -2.8038 | [10 empty] | 0 | 1.6703 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5296 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42503 | 0.43909 | -1.3883 | 0.014056 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 99.91% | 8.741 | -6.0315 | -5.0627 | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 105.4% | 6.2903 | 2.0221 | 22.3022 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_015 | iRM_12 |  |  | 0.0002977 | 8.1133e-05 | 0.00010288 | 4.9861e-05 | 0.0001004 | 3.2279e-06 | 0.00014999 | 8.6576e-06 | 4.1724e-06 | [5 empty] | 23.039 | 5.115 | 7.9684 | 0.00012626 | 8.3773 | 3.8249e-05 | 1.3958 | 2.0202e-05 | 2.5442e-06 |  |  |  | 1 | 0 | 0 | 23.0387 | 5.1149 | 7.9683 | 7.6397e-05 | 8.3772 | 3.5022e-05 | 1.3957 | 1.1544e-05 | -1.6282e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6707 | 23.0387 | 5.1149 | 7.9683 | 8.3772 | 1.3957 | 0.22202 | 0.34587 | 0.36362 | 0.060579 | 1.5579 | 1.6378 | 0.27286 | 1.0513 | 0.17515 | 0.1666 | -1.505 | -1.0617 | -1.0117 | -2.8038 | [10 empty] | 0 | 1.6707 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43909 | -1.3883 | 0.014051 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.51% | -3.7238 | -0.81394 | -10.2838 | [8 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 106% | -4.3284 | -2.7454 | 5.4767 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_020 | iRM_2 |  |  | 0.0002951 | 7.0996e-05 | 9.8881e-05 | 5.0973e-05 | 0.00011112 | 7.8774e-06 | 0.00015291 | 1.2721e-05 | 5.279e-07 | [5 empty] | 23.1483 | 5.1394 | 8.0067 | 0.00012089 | 8.4182 | 3.8289e-05 | 1.4027 | 1.4899e-05 | 4.2388e-06 |  |  |  | 1 | 0 | 0 | 23.148 | 5.1394 | 8.0066 | 6.9915e-05 | 8.4181 | 3.0411e-05 | 1.4026 | 2.1782e-06 | 3.7109e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6737 | 23.148 | 5.1394 | 8.0066 | 8.418 | 1.4026 | 0.22202 | 0.34589 | 0.36366 | 0.060591 | 1.5579 | 1.638 | 0.27291 | 1.0514 | 0.17517 | 0.16661 | -1.505 | -1.0616 | -1.0115 | -2.8036 | [10 empty] | 0 | 1.6737 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014056 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 100.6% | -8.2098 | -4.438 | -7.7383 | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 108.4% | 3.2601 | -1.5368 | -0.24085 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_025 | iRM_12 |  |  | 0.00029226 | 6.6517e-05 | 9.7886e-05 | 4.7633e-05 | 0.00010131 | 4.4152e-06 | 0.00015153 | 2.4468e-06 | 1.727e-06 | [5 empty] | 22.971 | 5.1002 | 7.9455 | 0.0001194 | 8.3539 | 5.1933e-05 | 1.3921 | 1.7171e-05 | 4.2451e-06 |  |  |  | 1 | 0 | 0 | 22.9707 | 5.1001 | 7.9454 | 7.1764e-05 | 8.3538 | 4.7518e-05 | 1.3919 | 1.4724e-05 | 2.5181e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6744 | 22.9707 | 5.1001 | 7.9454 | 8.3538 | 1.3919 | 0.22203 | 0.3459 | 0.36367 | 0.060595 | 1.5579 | 1.638 | 0.27292 | 1.0514 | 0.17518 | 0.16662 | -1.505 | -1.0616 | -1.0115 | -2.8035 | [10 empty] | 0 | 1.6744 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5296 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42503 | 0.43908 | -1.3883 | 0.014052 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.96% | 11.5767 | -1.3464 | 8.4289 | [8 empty] | iRM_11 | Lab 1 | Apr 5, 2022 | 3 | 107.8% | -2.8828 | -1.436 | 11.3489 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_030 | iRM_10 |  |  | 0.00029005 | 7.3907e-05 | 9.3938e-05 | 4.1449e-05 | 0.00010256 | 1.8189e-05 | 0.00015548 | -3.3366e-07 | 7.6162e-07 | [5 empty] | 23.0796 | 5.1244 | 7.9836 | 0.00011116 | 8.3945 | 5.2463e-05 | 1.3989 | 1.0617e-05 | 3.4622e-06 |  |  |  | 1 | 0 | 0 | 23.0793 | 5.1243 | 7.9835 | 6.9706e-05 | 8.3944 | 3.4274e-05 | 1.3987 | 1.0951e-05 | 2.7006e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6772 | 23.0793 | 5.1243 | 7.9835 | 8.3944 | 1.3987 | 0.22203 | 0.34592 | 0.36372 | 0.060606 | 1.558 | 1.6382 | 0.27296 | 1.0515 | 0.1752 | 0.16663 | -1.5049 | -1.0616 | -1.0114 | -2.8034 | [10 empty] | 0 | 1.6772 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014056 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 100.6% | -8.1597 | -4.5443 | 2.9292
Lab 1 | Jun 20, 2022 | 1 | Zr_035 | iRM_12 |  |  | 0.00028353 | 7.6725e-05 | 9.9085e-05 | 4.7844e-05 | 9.2266e-05 | 2.0547e-05 | 0.00014846 | -5.1319e-07 | -2.3998e-06 | [5 empty] | 22.9226 | 5.0895 | 7.9294 | 0.00011233 | 8.3375 | 4.2118e-05 | 1.3894 | 1.4472e-05 | 3.9951e-06 |  |  |  | 1 | 0 | 0 | 22.9223 | 5.0895 | 7.9293 | 6.4485e-05 | 8.3374 | 2.1571e-05 | 1.3893 | 1.4985e-05 | 6.3949e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6777 | 22.9223 | 5.0895 | 7.9293 | 8.3374 | 1.3893 | 0.22203 | 0.34592 | 0.36373 | 0.060607 | 1.558 | 1.6382 | 0.27297 | 1.0515 | 0.1752 | 0.16663 | -1.5049 | -1.0615 | -1.0114 | -2.8033 | [10 empty] | 0 | 1.6777 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014051 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.9% | -5.7185 | 1.3635 | -3.3013 | [8 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 99.99% | -1.3069 | -2.5165 | -2.7574 | iRM_12 | Avg | 0.78423 | -2.1024 | 2.7517
Lab 1 | Jun 20, 2022 | 1 | Zr_040 | iRM_11 |  |  | 0.00028451 | 6.799e-05 | 0.00010267 | 5.2461e-05 | 0.00010098 | 6.1126e-06 | 0.00014612 | 1.2524e-05 | -7.0634e-07 | [5 empty] | 22.919 | 5.0888 | 7.9283 | 0.00011036 | 8.3365 | 4.4054e-05 | 1.3893 | 1.4515e-05 | 3.9223e-06 |  |  |  | 1 | 0 | 0 | 22.9187 | 5.0887 | 7.9282 | 5.7895e-05 | 8.3364 | 3.7941e-05 | 1.3891 | 1.9912e-06 | 4.6286e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6786 | 22.9187 | 5.0887 | 7.9282 | 8.3364 | 1.3891 | 0.22203 | 0.34593 | 0.36374 | 0.06061 | 1.558 | 1.6382 | 0.27298 | 1.0515 | 0.17521 | 0.16663 | -1.5049 | -1.0615 | -1.0113 | -2.8033 | [10 empty] | 0 | 1.6786 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014054 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 99.99% | -1.3069 | -2.5165 | -2.7574 | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 101.9% | -7.0174 | -8.3555 | -11.2734 | iRM_12 | 2SD | 8.3598 | 10.5208 | 20.1346
Lab 1 | Jun 20, 2022 | 1 | Zr_045 | iRM_12 |  |  | 0.00033904 | 8.2275e-05 | 0.00010562 | 5.1289e-05 | 0.00010988 | 2.2233e-06 | 0.00014806 | 1.6551e-05 | 3.9292e-06 | [5 empty] | 22.9194 | 5.0889 | 7.9286 | 0.0001118 | 8.3371 | 3.7878e-05 | 1.3894 | 1.5536e-05 | 4.8232e-06 |  |  |  | 1 | 0 | 0 | 22.919 | 5.0888 | 7.9285 | 6.0513e-05 | 8.337 | 3.5655e-05 | 1.3892 | -1.0146e-06 | 8.9404e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6796 | 22.919 | 5.0888 | 7.9285 | 8.337 | 1.3892 | 0.22204 | 0.34593 | 0.36376 | 0.060615 | 1.558 | 1.6383 | 0.273 | 1.0515 | 0.17522 | 0.16664 | -1.5049 | -1.0615 | -1.0113 | -2.8032 | [10 empty] | 0 | 1.6796 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014052 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.2% | -0.77211 | -0.9106 | -1.1851 | [8 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 102.4% | -0.28882 | 3.3921 | -1.683 | iRM_12 | n | 9 | 9 | 9
Lab 1 | Jun 20, 2022 | 1 | Zr_050 | iRM |  |  | 0.00029243 | 7.3103e-05 | 0.00010094 | 4.64e-05 | 0.00010953 | 7.0048e-06 | 0.00015228 | 2.3704e-05 | 3.5447e-06 | [5 empty] | 23.5369 | 5.2261 | 8.1424 | 0.00011507 | 8.5621 | 3.9001e-05 | 1.4269 | 1.4599e-05 | 3.4948e-06 |  |  |  | 1 | 0 | 0 | 23.5366 | 5.226 | 8.1423 | 6.8666e-05 | 8.562 | 3.1996e-05 | 1.4267 | -9.1053e-06 | -4.9919e-08 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6807 | 23.5366 | 5.226 | 8.1423 | 8.562 | 1.4267 | 0.22204 | 0.34594 | 0.36377 | 0.060618 | 1.558 | 1.6383 | 0.27301 | 1.0515 | 0.17523 | 0.16664 | -1.5049 | -1.0615 | -1.0112 | -2.8032 | [10 empty] | 0 | 1.6807 | 0.21795 | 0.33338 | 0.3381 | 0.054377 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014054 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 1 | 102.9% | 5.3687 | -2.5492 | -15.0364 | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 103.3% | 3.7549 | -6.4412 | -0.24384 | iRM_12 | 2SE | 2.7866 | 3.5069 | 6.7115
Lab 1 | Jun 20, 2022 | 1 | Zr_055 | iRM_12 |  |  | 0.00028448 | 8.0577e-05 | 0.0001025 | 5.2208e-05 | 0.00010469 | 8.3688e-06 | 0.00014747 | 1.3121e-05 | 2.7533e-06 | [5 empty] | 22.8414 | 5.0717 | 7.902 | 0.00010709 | 8.3094 | 4.0161e-05 | 1.3848 | 1.2748e-05 | 5.8346e-06 |  |  |  | 1 | 0 | 0 | 22.8411 | 5.0716 | 7.9019 | 5.4886e-05 | 8.3093 | 3.1792e-05 | 1.3847 | -3.7259e-07 | 3.0813e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6816 | 22.8411 | 5.0716 | 7.9019 | 8.3093 | 1.3847 | 0.22204 | 0.34595 | 0.36379 | 0.060622 | 1.5581 | 1.6384 | 0.27302 | 1.0516 | 0.17523 | 0.16664 | -1.5049 | -1.0615 | -1.0112 | -2.8031 | [10 empty] | 0 | 1.6816 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014052 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.85% | -0.47219 | 3.0238 | 8.611 | [8 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 104.2% | 6.9019 | -1.1684 | 20.1471 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_060 | iRM_1 |  |  | 0.0019267 | 0.00044217 | 0.00066344 | 4.4089e-05 | 0.00070564 | 4.5146e-06 | 0.00024586 | 3.6402e-06 | 1.8904e-06 | [5 empty] | 22.9549 | 5.0971 | 7.9417 | 0.00011007 | 8.3516 | 3.9137e-05 | 1.3919 | 1.6704e-05 | 7.0967e-06 |  |  |  | 1 | 0 | 0 | 22.9529 | 5.0966 | 7.941 | 6.5979e-05 | 8.3509 | 3.4622e-05 | 1.3917 | 1.3064e-05 | 5.2063e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6841 | 22.9529 | 5.0966 | 7.941 | 8.3509 | 1.3917 | 0.22205 | 0.34597 | 0.36383 | 0.060633 | 1.5581 | 1.6385 | 0.27306 | 1.0516 | 0.17525 | 0.16665 | -1.5049 | -1.0614 | -1.0111 | -2.8029 | [10 empty] | 0 | 1.6841 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014052 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 100.5% | 4.1243 | 2.0526 | 13.0788 | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 104.2% | -2.1442 | -6.5936 | -6.704 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_065 | iRM_12 |  |  | 0.00027985 | 7.7508e-05 | 9.2079e-05 | 3.9939e-05 | 0.00010586 | 4.5881e-06 | 0.00016005 | 5.0528e-06 | 2.6537e-06 | [5 empty] | 22.8329 | 5.07 | 7.8996 | 0.00010801 | 8.3075 | 3.9342e-05 | 1.3846 | 1.4044e-05 | 4.613e-06 |  |  |  | 1 | 0 | 0 | 22.8326 | 5.0699 | 7.8995 | 6.8071e-05 | 8.3074 | 3.4754e-05 | 1.3844 | 8.9914e-06 | 1.9593e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6849 | 22.8326 | 5.0699 | 7.8995 | 8.3074 | 1.3844 | 0.22205 | 0.34597 | 0.36384 | 0.060635 | 1.5581 | 1.6386 | 0.27307 | 1.0516 | 0.17526 | 0.16665 | -1.5049 | -1.0614 | -1.011 | -2.8029 | [10 empty] | 0 | 1.6849 | 0.21794 | 0.33338 | 0.3381 | 0.054377 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.5% | -7.6963 | -5.7943 | -18.0488 | [8 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 105.9% | 0.34907 | -1.7504 | 6.4978 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_070 | iRM_2 |  |  | 0.0027961 | 0.00062263 | 0.00095884 | 3.7223e-05 | 0.0010296 | 4.8325e-06 | 0.00031949 | 1.241e-05 | 1.2319e-06 | [5 empty] | 23.2282 | 5.1579 | 8.0365 | 0.00010663 | 8.4517 | 3.7075e-05 | 1.4087 | 2.2658e-05 | 9.4854e-07 |  |  |  | 1 | 0 | 0 | 23.2254 | 5.1572 | 8.0356 | 6.9407e-05 | 8.4506 | 3.2242e-05 | 1.4084 | 1.0248e-05 | -2.8339e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6858 | 23.2254 | 5.1572 | 8.0356 | 8.4506 | 1.4084 | 0.22205 | 0.34598 | 0.36385 | 0.060639 | 1.5581 | 1.6386 | 0.27308 | 1.0517 | 0.17527 | 0.16666 | -1.5048 | -1.0614 | -1.011 | -2.8028 | [10 empty] | 0 | 1.6858 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43909 | -1.3883 | 0.014052 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 102.2% | -0.0045312 | 2.2875 | -0.99852 | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 106.3% | 4.9269 | 8.3359 | 14.4127 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_075 | iRM_12 |  |  | 0.00030248 | 7.3684e-05 | 0.00010093 | 5.0229e-05 | 0.00011897 | 3.5695e-06 | 0.0001569 | 7.0506e-06 | 2.9015e-06 | [5 empty] | 22.614 | 5.0216 | 7.8242 | 0.00010261 | 8.2286 | 3.8465e-05 | 1.3715 | 1.4056e-05 | 5.5718e-06 |  |  |  | 1 | 0 | 0 | 22.6137 | 5.0215 | 7.8241 | 5.2383e-05 | 8.2285 | 3.4895e-05 | 1.3714 | 7.0059e-06 | 2.6702e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.687 | 22.6137 | 5.0215 | 7.8241 | 8.2285 | 1.3714 | 0.22206 | 0.34599 | 0.36387 | 0.060644 | 1.5581 | 1.6386 | 0.2731 | 1.0517 | 0.17528 | 0.16666 | -1.5048 | -1.0613 | -1.011 | -2.8027 | [10 empty] | 0 | 1.687 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43909 | -1.3883 | 0.014051 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 98.98% | 13.401 | 4.258 | 15.3195 | [8 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 107% | 1.8825 | -3.8238 | 6.3693 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_080 | iRM_10 |  |  | 0.0019662 | 0.00044899 | 0.00067892 | 3.4564e-05 | 0.00072554 | 5.8445e-06 | 0.00026554 | 1.088e-05 | 1.0849e-07 | [5 empty] | 22.9658 | 5.0997 | 7.9462 | 9.4761e-05 | 8.357 | 3.6161e-05 | 1.393 | 1.123e-05 | 1.6268e-06 |  |  |  | 1 | 0 | 0 | 22.9638 | 5.0993 | 7.9455 | 6.0197e-05 | 8.3563 | 3.0316e-05 | 1.3927 | 3.5014e-07 | 1.5183e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6881 | 22.9638 | 5.0993 | 7.9455 | 8.3563 | 1.3927 | 0.22206 | 0.346 | 0.36389 | 0.060648 | 1.5582 | 1.6387 | 0.27312 | 1.0517 | 0.17528 | 0.16667 | -1.5048 | -1.0613 | -1.0109 | -2.8027 | [10 empty] | 0 | 1.688 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014052 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 101% | -2.9173 | 0.54313 | 7.2901
Lab 1 | Jun 20, 2022 | 1 | Zr_085 | iRM_12 |  |  | 0.0002784 | 6.6936e-05 | 9.6442e-05 | 3.7923e-05 | 0.00010562 | 4.6658e-06 | 0.00015938 | 1.4044e-05 | 2.1232e-06 | [5 empty] | 22.8629 | 5.0768 | 7.9105 | 9.8143e-05 | 8.3195 | 3.7998e-05 | 1.3867 | 1.5178e-05 | 4.7655e-06 |  |  |  | 1 | 0 | 0 | 22.8626 | 5.0768 | 7.9104 | 6.022e-05 | 8.3194 | 3.3332e-05 | 1.3865 | 1.1337e-06 | 2.6424e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6879 | 22.8626 | 5.0768 | 7.9104 | 8.3194 | 1.3865 | 0.22206 | 0.346 | 0.36389 | 0.060647 | 1.5582 | 1.6387 | 0.27311 | 1.0517 | 0.17528 | 0.16666 | -1.5048 | -1.0613 | -1.0109 | -2.8027 | [10 empty] | 0 | 1.6879 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014053 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.8% | -7.1557 | -2.705 | -11.1214 | [8 empty] | iRM_12 | Lab 1 | Apr 4, 2022 | 1 | 98.91% | -9.2161 | -0.63121 | 8.0797 | iRM | Avg | -0.40694 | -0.099125 | 0.85021
Lab 1 | Jun 20, 2022 | 1 | Zr_090 | iRM_11 |  |  | 0.00027565 | 6.9166e-05 | 9.3228e-05 | 3.8202e-05 | 0.0001009 | 6.625e-06 | 0.00014706 | 1.465e-05 | 8.4566e-08 | [5 empty] | 23.2373 | 5.16 | 8.0402 | 9.0245e-05 | 8.4563 | 3.7829e-05 | 1.4095 | 1.5036e-05 | 3.1855e-06 |  |  |  | 1 | 0 | 0 | 23.237 | 5.16 | 8.0401 | 5.2043e-05 | 8.4562 | 3.1204e-05 | 1.4094 | 3.8627e-07 | 3.1009e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6893 | 23.237 | 5.16 | 8.0401 | 8.4562 | 1.4094 | 0.22206 | 0.34601 | 0.36391 | 0.060652 | 1.5582 | 1.6388 | 0.27314 | 1.0517 | 0.17529 | 0.16667 | -1.5048 | -1.0613 | -1.0108 | -2.8026 | [10 empty] | 0 | 1.6893 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.01406 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 101.9% | -7.0174 | -8.3555 | -11.2734 | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 98.55% | -7.5159 | -8.6024 | 9.85 | iRM_12 | 2SD | 14.0124 | 9.3872 | 25.3326
Lab 1 | Jun 20, 2022 | 1 | Zr_095 | iRM_12 |  |  | 0.00028608 | 7.5633e-05 | 8.4231e-05 | 3.2732e-05 | 9.9093e-05 | 2.9631e-06 | 0.00014066 | 1.6398e-05 | -7.211e-07 | [5 empty] | 22.7276 | 5.0469 | 7.864 | 9.1248e-05 | 8.271 | 3.8294e-05 | 1.3787 | 1.6975e-05 | 5.421e-06 |  |  |  | 1 | 0 | 0 | 22.7273 | 5.0469 | 7.8639 | 5.8516e-05 | 8.2709 | 3.5331e-05 | 1.3785 | 5.7697e-07 | 6.1421e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6898 | 22.7273 | 5.0469 | 7.8639 | 8.2708 | 1.3785 | 0.22206 | 0.34601 | 0.36392 | 0.060655 | 1.5582 | 1.6388 | 0.27315 | 1.0517 | 0.1753 | 0.16667 | -1.5048 | -1.0613 | -1.0108 | -2.8026 | [10 empty] | 0 | 1.6898 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0141 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.01405 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.69% | 3.8723 | 3.7446 | 3.1252 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.25% | -4.8912 | 0.45562 | 6.9415 | iRM_12 | n | 34 | 34 | 34
Lab 1 | Jun 20, 2022 | 1 | Zr_100 | iRM |  |  | 0.00032413 | 6.9353e-05 | 0.00010869 | 2.9289e-05 | 0.00012412 | 6.1534e-06 | 0.00016884 | 6.4917e-06 | 5.6054e-06 | [5 empty] | 23.6531 | 5.2525 | 8.1844 | 8.9192e-05 | 8.6081 | 4.0235e-05 | 1.4349 | 1.9408e-05 | 2.1838e-06 |  |  |  | 1 | 0 | 0 | 23.6528 | 5.2524 | 8.1843 | 5.9904e-05 | 8.608 | 3.4081e-05 | 1.4347 | 1.2916e-05 | -3.4216e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6907 | 23.6528 | 5.2524 | 8.1843 | 8.608 | 1.4347 | 0.22206 | 0.34602 | 0.36393 | 0.060659 | 1.5582 | 1.6389 | 0.27316 | 1.0518 | 0.1753 | 0.16668 | -1.5048 | -1.0613 | -1.0108 | -2.8025 | [10 empty] | 0 | 1.6907 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43909 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 1 | 104.1% | 3.4708 | -4.7714 | -8.5831 | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 100.4% | 1.6111 | 1.5565 | -17.5867 | iRM_12 | 2SE | 2.4031 | 1.6099 | 4.3445
Lab 1 | Jun 20, 2022 | 1 | Zr_105 | iRM_12 |  |  | 0.00027161 | 7.4487e-05 | 9.3546e-05 | 2.9614e-05 | 0.0001019 | 4.6542e-06 | 0.00014549 | 9.2901e-06 | 1.1417e-06 | [5 empty] | 22.7356 | 5.0488 | 7.867 | 9.1069e-05 | 8.2744 | 3.9769e-05 | 1.3793 | 2.0147e-05 | 5.6117e-06 |  |  |  | 1 | 0 | 0 | 22.7353 | 5.0487 | 7.8669 | 6.1455e-05 | 8.2743 | 3.5115e-05 | 1.3791 | 1.0856e-05 | 4.47e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6911 | 22.7353 | 5.0487 | 7.8669 | 8.2742 | 1.3791 | 0.22206 | 0.34602 | 0.36394 | 0.060661 | 1.5582 | 1.6389 | 0.27317 | 1.0518 | 0.17531 | 0.16668 | -1.5048 | -1.0613 | -1.0108 | -2.8025 | [10 empty] | 0 | 1.6911 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014055 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.4% | -3.6809 | -2.0788 | 3.8565 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.39% | -4.885 | -0.55477 | 9.0958 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_110 | iRM_1 |  |  | 0.00027674 | 4.6234e-05 | 8.4787e-05 | 2.7512e-05 | 0.00010688 | 6.3305e-06 | 0.00015263 | 1.6734e-05 | 4.014e-06 | [5 empty] | 23.0768 | 5.1246 | 7.9853 | 9.0047e-05 | 8.399 | 3.8905e-05 | 1.4001 | 1.5607e-05 | 7.2593e-06 |  |  |  | 1 | 0 | 0 | 23.0765 | 5.1246 | 7.9852 | 6.2534e-05 | 8.3989 | 3.2574e-05 | 1.3999 | -1.1278e-06 | 3.2453e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6925 | 23.0765 | 5.1246 | 7.9852 | 8.3989 | 1.3999 | 0.22207 | 0.34603 | 0.36396 | 0.060665 | 1.5582 | 1.639 | 0.27318 | 1.0518 | 0.17531 | 0.16668 | -1.5048 | -1.0612 | -1.0107 | -2.8024 | [10 empty] | 0 | 1.6925 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43909 | -1.3883 | 0.014054 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 101.9% | 4.6207 | 0.9203 | -26.4963 | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.55% | 9.831 | 3.5015 | -2.1286 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_115 | iRM_12 |  |  | 0.00027162 | 7.4835e-05 | 9.283e-05 | 2.778e-05 | 0.00010146 | 5.8574e-06 | 0.00013068 | 8.9131e-06 | 4.3918e-06 | [5 empty] | 22.57 | 5.0121 | 7.8099 | 8.2532e-05 | 8.2145 | 3.8762e-05 | 1.3693 | 1.8727e-05 | 5.027e-06 |  |  |  | 1 | 0 | 0 | 22.5697 | 5.012 | 7.8098 | 5.4752e-05 | 8.2144 | 3.2904e-05 | 1.3692 | 9.8143e-06 | 6.3518e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6923 | 22.5697 | 5.012 | 7.8098 | 8.2144 | 1.3692 | 0.22207 | 0.34603 | 0.36396 | 0.060665 | 1.5582 | 1.639 | 0.27318 | 1.0518 | 0.17532 | 0.16668 | -1.5048 | -1.0612 | -1.0107 | -2.8024 | [10 empty] | 0 | 1.6923 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014055 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.4% | 1.7402 | -0.38001 | 15.5945 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 100.9% | -8.6841 | -5.4091 | 0.09599 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_120 | iRM_2 |  |  | 0.00027126 | 7.3243e-05 | 9.4648e-05 | 2.6273e-05 | 9.5995e-05 | 5.2004e-06 | 0.00014625 | 1.2848e-05 | 6.3845e-06 | [5 empty] | 23.3185 | 5.1784 | 8.0692 | 8.5276e-05 | 8.4876 | 3.7464e-05 | 1.4149 | 1.1284e-05 | 4.2669e-06 |  |  |  | 1 | 0 | 0 | 23.3183 | 5.1783 | 8.0692 | 5.9003e-05 | 8.4875 | 3.2263e-05 | 1.4148 | -1.5638e-06 | -2.1176e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.694 | 23.3183 | 5.1783 | 8.0692 | 8.4875 | 1.4148 | 0.22207 | 0.34604 | 0.36398 | 0.060672 | 1.5583 | 1.639 | 0.27321 | 1.0518 | 0.17533 | 0.16669 | -1.5048 | -1.0612 | -1.0106 | -2.8023 | [10 empty] | 0 | 1.694 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014053 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 103.1% | 1.7339 | 2.0491 | 10.329 | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.3% | 5.2606 | 2.3574 | 0.83388 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_125 | iRM_12 |  |  | 0.0002724 | 7.1529e-05 | 9.8251e-05 | 2.5972e-05 | 9.2092e-05 | 3.2874e-06 | 0.00016081 | 6.1777e-06 | 5.5858e-06 | [5 empty] | 22.6773 | 5.036 | 7.8473 | 8.5409e-05 | 8.2541 | 3.6932e-05 | 1.376 | 1.4656e-05 | 5.5815e-06 |  |  |  | 1 | 0 | 0 | 22.6771 | 5.0359 | 7.8472 | 5.9436e-05 | 8.254 | 3.3644e-05 | 1.3758 | 8.4782e-06 | -4.2725e-09 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6937 | 22.6771 | 5.0359 | 7.8472 | 8.254 | 1.3758 | 0.22207 | 0.34604 | 0.36398 | 0.060669 | 1.5583 | 1.639 | 0.2732 | 1.0518 | 0.17532 | 0.16668 | -1.5048 | -1.0612 | -1.0107 | -2.8023 | [10 empty] | 0 | 1.6937 | 0.21794 | 0.33338 | 0.3381 | 0.054377 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014055 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.2% | -1.7481 | 2.5403 | -26.9663 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 99.96% | -3.1594 | -1.1881 | 6.2284 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_130 | iRM_10 |  |  | 0.00027322 | 7.4827e-05 | 9.9674e-05 | 2.3274e-05 | 0.0001076 | 6.3015e-06 | 0.00013906 | 2.5878e-06 | 5.6631e-06 | [5 empty] | 23.0356 | 5.1157 | 7.9716 | 8.1527e-05 | 8.3853 | 3.8456e-05 | 1.3979 | 1.6478e-05 | 3.4544e-06 |  |  |  | 1 | 0 | 0 | 23.0354 | 5.1156 | 7.9716 | 5.8253e-05 | 8.3852 | 3.2154e-05 | 1.3978 | 1.3891e-05 | -2.2087e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6958 | 23.0354 | 5.1156 | 7.9716 | 8.3852 | 1.3978 | 0.22208 | 0.34606 | 0.36401 | 0.06068 | 1.5583 | 1.6391 | 0.27324 | 1.0519 | 0.17535 | 0.1667 | -1.5047 | -1.0612 | -1.0106 | -2.8021 | [10 empty] | 0 | 1.6958 | 0.21795 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3882 | 0.014055 | -1.8133 | -1.8273 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 101.5% | 1.5505 | 1.5702 | 38.3952 | iRM_12 | Lab 1 | Apr 5, 2022 | 2 | 100.3% | 8.242 | 2.0289 | -17.0452 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_135 | iRM_12 |  |  | 0.00026722 | 8.4288e-05 | 9.6127e-05 | 2.6363e-05 | 9.9393e-05 | 3.2375e-06 | 0.0001582 | 3.3045e-06 | 1.7039e-06 | [5 empty] | 22.7008 | 5.0413 | 7.8556 | 7.7515e-05 | 8.2632 | 3.8913e-05 | 1.3775 | 1.5912e-05 | 5.7188e-06 |  |  |  | 1 | 0 | 0 | 22.7005 | 5.0412 | 7.8555 | 5.1152e-05 | 8.2631 | 3.5676e-05 | 1.3774 | 1.2607e-05 | 4.0149e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6952 | 22.7005 | 5.0412 | 7.8555 | 8.2631 | 1.3774 | 0.22207 | 0.34605 | 0.364 | 0.060676 | 1.5583 | 1.6391 | 0.27322 | 1.0519 | 0.17534 | 0.16669 | -1.5047 | -1.0612 | -1.0106 | -2.8022 | [10 empty] | 0 | 1.6952 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014059 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.4% | 6.0205 | -2.7476 | 4.0694 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.99% | 11.6984 | 10.4048 | 22.7649 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_140 | iRM_11 |  |  | 0.00027139 | 7.8898e-05 | 9.296e-05 | 2.9737e-05 | 9.2227e-05 | 3.579e-06 | 0.00015853 | 6.2465e-06 | 5.4056e-06 | [5 empty] | 23.1689 | 5.1453 | 8.0178 | 7.5316e-05 | 8.4339 | 3.6811e-05 | 1.406 | 1.4323e-05 | 3.0717e-06 |  |  |  | 1 | 0 | 0 | 23.1687 | 5.1452 | 8.0177 | 4.5579e-05 | 8.4338 | 3.3232e-05 | 1.4059 | 8.077e-06 | -2.3339e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6961 | 23.1687 | 5.1452 | 8.0177 | 8.4338 | 1.4059 | 0.22208 | 0.34606 | 0.36402 | 0.06068 | 1.5583 | 1.6392 | 0.27324 | 1.0519 | 0.17535 | 0.1667 | -1.5047 | -1.0611 | -1.0106 | -2.8021 | [10 empty] | 0 | 1.6961 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014055 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 102.4% | -0.28882 | 3.3921 | -1.683 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.3% | -14.6271 | -5.09 | -20.0219 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_145 | iRM_12 |  |  | 0.00026946 | 7.485e-05 | 8.5159e-05 | 2.0855e-05 | 9.9865e-05 | 8.0849e-06 | 0.00015293 | -3.6533e-06 | 4.4048e-06 | [5 empty] | 22.5546 | 5.0088 | 7.8052 | 7.9116e-05 | 8.2103 | 3.6969e-05 | 1.3688 | 9.1779e-06 | 6.3056e-06 |  |  |  | 1 | 0 | 0 | 22.5544 | 5.0088 | 7.8051 | 5.8261e-05 | 8.2102 | 2.8884e-05 | 1.3686 | 1.2831e-05 | 1.9007e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.696 | 22.5544 | 5.0088 | 7.8051 | 8.2102 | 1.3686 | 0.22207 | 0.34606 | 0.36402 | 0.06068 | 1.5583 | 1.6392 | 0.27324 | 1.0519 | 0.17535 | 0.1667 | -1.5047 | -1.0612 | -1.0106 | -2.8021 | [10 empty] | 0 | 1.696 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014058 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.64% | -11.4069 | -2.8357 | 11.2368 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.57% | 7.5591 | 0.42347 | -3.2582 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_150 | iRM |  |  | 0.0002748 | 6.7609e-05 | 9.3643e-05 | 2.9139e-05 | 0.00010171 | 5.8467e-06 | 0.00016185 | -5.9957e-06 | 5.7021e-06 | [5 empty] | 23.4946 | 5.2176 | 8.1306 | 7.7565e-05 | 8.5526 | 3.5299e-05 | 1.4258 | 6.3882e-06 | 1.667e-06 |  |  |  | 1 | 0 | 0 | 23.4943 | 5.2176 | 8.1305 | 4.8425e-05 | 8.5525 | 2.9453e-05 | 1.4257 | 1.2384e-05 | -4.035e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6966 | 23.4943 | 5.2176 | 8.1305 | 8.5525 | 1.4257 | 0.22208 | 0.34606 | 0.36403 | 0.060682 | 1.5583 | 1.6392 | 0.27325 | 1.0519 | 0.17535 | 0.1667 | -1.5047 | -1.0611 | -1.0105 | -2.8021 | [10 empty] | 0 | 1.6966 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014061 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 1 | 104.1% | 0.45672 | -5.3748 | 1.1419 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.3% | -2.1986 | 0.6795 | 31.3167 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_155 | iRM_12 |  |  | 0.0026487 | 0.00058861 | 0.00091874 | 1.8804e-05 | 0.00095964 | -8.6895e-07 | 0.0002967 | -5.2836e-07 | 2.9922e-06 | [5 empty] | 22.5745 | 5.0133 | 7.8123 | 7.3857e-05 | 8.2179 | 3.6931e-05 | 1.37 | 1.3756e-05 | 5.728e-06 |  |  |  | 1 | 0 | 0 | 22.5718 | 5.0128 | 7.8114 | 5.5052e-05 | 8.2169 | 3.78e-05 | 1.3697 | 1.4285e-05 | 2.7358e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6971 | 22.5718 | 5.0128 | 7.8114 | 8.2169 | 1.3697 | 0.22208 | 0.34607 | 0.36403 | 0.060684 | 1.5583 | 1.6392 | 0.27325 | 1.0519 | 0.17535 | 0.1667 | -1.5047 | -1.0611 | -1.0105 | -2.8021 | [10 empty] | 0 | 1.6971 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43909 | -1.3883 | 0.014052 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.1% | 10.5155 | 6.355 | 3.4458 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.84% | -1.3618 | -4.2492 | -28.2706 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_160 | iRM_1 |  |  | 0.00027685 | 5.4892e-05 | 9.8189e-05 | 2.253e-05 | 0.00010297 | 7.6437e-06 | 0.00015498 | 9.6144e-06 | 1.2928e-06 | [5 empty] | 23.1013 | 5.1304 | 7.9948 | 7.3457e-05 | 8.4101 | 3.7138e-05 | 1.4021 | 1.2137e-05 | 4.8529e-06 |  |  |  | 1 | 0 | 0 | 23.101 | 5.1303 | 7.9947 | 5.0927e-05 | 8.41 | 2.9495e-05 | 1.402 | 2.5226e-06 | 3.5601e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6983 | 23.101 | 5.1303 | 7.9947 | 8.41 | 1.402 | 0.22208 | 0.34608 | 0.36405 | 0.060689 | 1.5583 | 1.6393 | 0.27327 | 1.0519 | 0.17536 | 0.1667 | -1.5047 | -1.0611 | -1.0105 | -2.802 | [10 empty] | 0 | 1.6983 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014059 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 102.4% | 0.42284 | -4.3221 | 9.4976 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.4% | 1.7773 | 5.6029 | 0.85659 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_165 | iRM_12 |  |  | 0.00028413 | 6.5091e-05 | 9.3296e-05 | 1.9468e-05 | 9.7615e-05 | 3.6777e-06 | 0.00015016 | 1.4693e-06 | 3.9906e-06 | [5 empty] | 22.5382 | 5.0053 | 7.7998 | 7.1674e-05 | 8.2049 | 3.7301e-05 | 1.3679 | 1.6143e-05 | 5.9258e-06 |  |  |  | 1 | 0 | 0 | 22.5379 | 5.0052 | 7.7997 | 5.2206e-05 | 8.2048 | 3.3623e-05 | 1.3677 | 1.4674e-05 | 1.9352e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6977 | 22.5379 | 5.0052 | 7.7997 | 8.2048 | 1.3677 | 0.22208 | 0.34607 | 0.36404 | 0.060685 | 1.5583 | 1.6392 | 0.27326 | 1.0519 | 0.17536 | 0.1667 | -1.5047 | -1.0611 | -1.0105 | -2.8021 | [10 empty] | 0 | 1.6977 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014058 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.91% | 0.6426 | -2.6768 | -17.0925 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.21% | 4.1596 | -3.8485 | 8.886 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_170 | iRM_2 |  |  | 0.00027142 | 6.0734e-05 | 9.6242e-05 | 1.553e-05 | 0.00010061 | 5.9345e-06 | 0.00014849 | 5.9407e-06 | 3.9698e-06 | [5 empty] | 23.295 | 5.1734 | 8.0619 | 7.4357e-05 | 8.4809 | 3.7222e-05 | 1.4139 | 1.2401e-05 | 2.2587e-06 |  |  |  | 1 | 0 | 0 | 23.2947 | 5.1733 | 8.0618 | 5.8827e-05 | 8.4808 | 3.1287e-05 | 1.4138 | 6.4606e-06 | -1.711e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6991 | 23.2947 | 5.1733 | 8.0618 | 8.4808 | 1.4138 | 0.22208 | 0.34608 | 0.36407 | 0.060692 | 1.5583 | 1.6393 | 0.27329 | 1.052 | 0.17537 | 0.16671 | -1.5047 | -1.0611 | -1.0104 | -2.8019 | [10 empty] | 0 | 1.6991 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014059 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 1 | 103.3% | 0.30847 | -0.8 | 8.1277 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.5% | -3.7754 | 1.784 | 0.87042 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_175 | iRM_12 |  |  | 0.00029201 | 6.9627e-05 | 0.00010391 | 1.5851e-05 | 0.00010244 | 6.4889e-06 | 0.00015716 | 1.3666e-05 | 1.614e-06 | [5 empty] | 22.5437 | 5.0065 | 7.8019 | 7.0359e-05 | 8.2072 | 3.4194e-05 | 1.3683 | 1.0919e-05 | 5.8564e-06 |  |  |  | 1 | 0 | 0 | 22.5434 | 5.0064 | 7.8018 | 5.4509e-05 | 8.2071 | 2.7705e-05 | 1.3682 | -2.7469e-06 | 4.2424e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6987 | 22.5434 | 5.0064 | 7.8018 | 8.2071 | 1.3682 | 0.22208 | 0.34608 | 0.36406 | 0.06069 | 1.5583 | 1.6393 | 0.27328 | 1.052 | 0.17537 | 0.1667 | -1.5047 | -1.0611 | -1.0104 | -2.802 | [10 empty] | 0 | 1.6987 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014059 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.1% | -8.4934 | -1.5922 | 7.9727 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.4% | -3.1975 | 0.91127 | -11.3724 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_180 | iRM_10 |  |  | 0.00027911 | 6.3261e-05 | 8.6865e-05 | 1.4044e-05 | 9.5343e-05 | 2.9293e-06 | 0.00015091 | 1.1228e-05 | 3.2075e-06 | [5 empty] | 23.0908 | 5.1281 | 7.9913 | 7.3228e-05 | 8.4066 | 3.3488e-05 | 1.4015 | 1.4948e-05 | 4.1409e-06 |  |  |  | 1 | 0 | 0 | 23.0905 | 5.128 | 7.9912 | 5.9185e-05 | 8.4065 | 3.0559e-05 | 1.4014 | 3.7196e-06 | 9.3336e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6994 | 23.0905 | 5.128 | 7.9912 | 8.4065 | 1.4014 | 0.22208 | 0.34608 | 0.36407 | 0.060691 | 1.5583 | 1.6393 | 0.27328 | 1.052 | 0.17537 | 0.1667 | -1.5047 | -1.0611 | -1.0104 | -2.802 | [10 empty] | 0 | 1.6994 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014056 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 1 | 102.5% | 3.9712 | 1.9916 | -26.0589 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.46% | 7.0653 | 0.56432 | 3.4118 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_185 | iRM_12 |  |  | 0.00027144 | 5.499e-05 | 9.3027e-05 | 1.6697e-05 | 9.9599e-05 | 3.3096e-06 | 0.00014124 | 1.2881e-05 | 2.2545e-06 | [5 empty] | 22.5266 | 5.0027 | 7.796 | 7.028e-05 | 8.201 | 3.6008e-05 | 1.3673 | 1.4094e-05 | 5.4001e-06 |  |  |  | 1 | 0 | 0 | 22.5263 | 5.0027 | 7.7959 | 5.3583e-05 | 8.2009 | 3.2698e-05 | 1.3671 | 1.2132e-06 | 3.1457e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6987 | 22.5263 | 5.0027 | 7.7959 | 8.2009 | 1.3671 | 0.22208 | 0.34608 | 0.36406 | 0.06069 | 1.5583 | 1.6393 | 0.27328 | 1.052 | 0.17537 | 0.16671 | -1.5047 | -1.0611 | -1.0104 | -2.802 | [10 empty] | 0 | 1.6987 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.2% | 2.7902 | -1.8929 | 1.8057 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100% | -6.2409 | -6.1191 | -1.2318 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_190 | iRM_11 |  |  | 0.00027535 | 6.7338e-05 | 0.0001033 | 2.0629e-05 | 0.00010529 | 2.7808e-06 | 0.00013544 | 9.099e-06 | 3.0376e-06 | [5 empty] | 23.2059 | 5.1537 | 8.0313 | 6.6973e-05 | 8.4489 | 3.5176e-05 | 1.4086 | 8.3364e-06 | 1.6521e-06 |  |  |  | 1 | 0 | 0 | 23.2057 | 5.1537 | 8.0312 | 4.6344e-05 | 8.4488 | 3.2395e-05 | 1.4085 | -7.6259e-07 | -1.3854e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7002 | 23.2057 | 5.1537 | 8.0312 | 8.4488 | 1.4085 | 0.22209 | 0.34609 | 0.36408 | 0.060696 | 1.5583 | 1.6394 | 0.2733 | 1.052 | 0.17538 | 0.16671 | -1.5047 | -1.0611 | -1.0104 | -2.8019 | [10 empty] | 0 | 1.7002 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.01406 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 1 | 103.3% | 3.7549 | -6.4412 | -0.24384 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.56% | -2.8011 | 7.3851 | 14.011 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_195 | iRM_12 |  |  | 0.00028455 | 5.7303e-05 | 8.7589e-05 | 9.9282e-06 | 0.0001052 | 7.2578e-06 | 0.00013983 | 1.884e-05 | 5.559e-06 | [5 empty] | 22.4222 | 4.9796 | 7.76 | 6.7419e-05 | 8.1633 | 3.5931e-05 | 1.361 | 1.2378e-05 | 4.5925e-06 |  |  |  | 1 | 0 | 0 | 22.4219 | 4.9796 | 7.7599 | 5.7491e-05 | 8.1632 | 2.8673e-05 | 1.3609 | -6.4621e-06 | -9.665e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6995 | 22.4219 | 4.9796 | 7.7599 | 8.1632 | 1.3609 | 0.22208 | 0.34609 | 0.36407 | 0.060693 | 1.5584 | 1.6393 | 0.27329 | 1.052 | 0.17537 | 0.16671 | -1.5047 | -1.0611 | -1.0104 | -2.8019 | [10 empty] | 0 | 1.6995 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0141 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.01405 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.47% | 6.5834 | 8.1805 | 3.3359 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 101.1% | 11.1635 | -3.3328 | -12.2254 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_200 | iRM |  |  | 0.0002765 | 6.5106e-05 | 9.0246e-05 | 1.1558e-05 | 9.5884e-05 | 4.3742e-06 | 0.00013086 | 1.2128e-05 | 5.4442e-07 | [5 empty] | 23.6116 | 5.2437 | 8.1715 | 6.5168e-05 | 8.5963 | 3.3067e-05 | 1.4332 | 7.8871e-06 | 2.4936e-06 |  |  |  | 1 | 0 | 0 | 23.6113 | 5.2437 | 8.1715 | 5.361e-05 | 8.5962 | 2.8693e-05 | 1.433 | -4.2409e-06 | 1.9492e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6994 | 23.6113 | 5.2437 | 8.1715 | 8.5962 | 1.433 | 0.22208 | 0.34608 | 0.36407 | 0.060693 | 1.5583 | 1.6393 | 0.27329 | 1.052 | 0.17537 | 0.16671 | -1.5047 | -1.0611 | -1.0104 | -2.8019 | [10 empty] | 0 | 1.6994 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014062 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 1 | 105% | 1.3611 | -6.9592 | 1.6414 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.71% | -8.2804 | 0.2123 | 2.2107 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_205 | iRM_12 |  |  | 0.00031818 | 6.5538e-05 | 0.00010866 | 1.7706e-05 | 0.00011727 | 4.9209e-06 | 0.00013588 | 2.39e-06 | 3.3479e-06 | [5 empty] | 22.5573 | 5.0096 | 7.8067 | 6.7845e-05 | 8.2124 | 3.4021e-05 | 1.3692 | 7.5882e-06 | 5.1391e-06 |  |  |  | 1 | 0 | 0 | 22.557 | 5.0095 | 7.8066 | 5.014e-05 | 8.2123 | 2.91e-05 | 1.369 | 5.1982e-06 | 1.7912e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.6993 | 22.557 | 5.0095 | 7.8066 | 8.2123 | 1.369 | 0.22208 | 0.34608 | 0.36407 | 0.060692 | 1.5584 | 1.6393 | 0.27329 | 1.052 | 0.17537 | 0.16671 | -1.5047 | -1.0611 | -1.0104 | -2.8019 | [10 empty] | 0 | 1.6993 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014061 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.4% | -10.5532 | -6.6985 | 3.7208 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.27% | -0.043363 | -1.0138 | -2.9812 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_210 | iRM_1 |  |  | 0.00030846 | 6.8663e-05 | 0.00010402 | 1.2405e-05 | 0.00010044 | 5.1099e-06 | 0.00013433 | 9.4428e-07 | -2.8565e-07 | [5 empty] | 23.2505 | 5.1637 | 8.0469 | 6.8666e-05 | 8.4654 | 3.5982e-05 | 1.4114 | 8.6772e-06 | 3.0633e-06 |  |  |  | 1 | 0 | 0 | 23.2502 | 5.1636 | 8.0468 | 5.6261e-05 | 8.4653 | 3.0872e-05 | 1.4113 | 7.7329e-06 | 3.3489e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7011 | 23.2502 | 5.1636 | 8.0468 | 8.4653 | 1.4113 | 0.22209 | 0.3461 | 0.3641 | 0.060699 | 1.5584 | 1.6394 | 0.27331 | 1.052 | 0.17538 | 0.16671 | -1.5047 | -1.061 | -1.0103 | -2.8018 | [10 empty] | 0 | 1.7011 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014058 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 1 | 103.2% | 8.3736 | 1.0795 | -2.2693 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.1% | 6.8239 | 2.8429 | 10.3542 | iRM_12
Lab 1 | Jun 20, 2022 | 1 | Zr_215 | iRM_12 |  |  | 0.00028186 | 6.0371e-05 | 9.1001e-05 | 1.3407e-05 | 9.5241e-05 | 7.9006e-06 | 0.0001442 | 4.0154e-06 | 4.0257e-06 | [5 empty] | 22.502 | 4.9974 | 7.7879 | 6.2873e-05 | 8.1929 | 3.4083e-05 | 1.366 | 1.5696e-05 | 4.9572e-06 |  |  |  | 1 | 0 | 0 | 22.5018 | 4.9974 | 7.7878 | 4.9466e-05 | 8.1928 | 2.6182e-05 | 1.3658 | 1.168e-05 | 9.3149e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.701 | 22.5018 | 4.9974 | 7.7878 | 8.1928 | 1.3658 | 0.22209 | 0.3461 | 0.36409 | 0.060698 | 1.5584 | 1.6394 | 0.27331 | 1.052 | 0.17538 | 0.16671 | -1.5047 | -1.061 | -1.0103 | -2.8018 | [10 empty] | 0 | 1.701 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [22 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 101.1% | -7.0288 | -2.8042 | -10.6751 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_005 | iRM_12 |  |  | 0.00025827 | 7.0368e-05 | 8.6334e-05 | 6.5603e-06 | 9.6127e-05 | 6.3555e-06 | 0.00012961 | 4.3002e-06 | 3.4715e-06 | [5 empty] | 22.5689 | 5.0123 | 7.811 | 6.2333e-05 | 8.2173 | 3.5967e-05 | 1.37 | 1.2557e-05 | 4.0707e-06 |  |  |  | 1 | 0 | 0 | 22.5687 | 5.0122 | 7.8109 | 5.5772e-05 | 8.2172 | 2.9611e-05 | 1.3699 | 8.2565e-06 | 5.9922e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7013 | 22.5687 | 5.0122 | 7.8109 | 8.2172 | 1.3699 | 0.22209 | 0.3461 | 0.3641 | 0.0607 | 1.5584 | 1.6394 | 0.27331 | 1.052 | 0.17538 | 0.16671 | -1.5047 | -1.061 | -1.0103 | -2.8018 | [9 empty] | L1 JunS2 | 0 | 1.7013 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014061 | -1.8133 | -1.8274 | [22 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.32% | 4.2492 | 6.2924 | 11.2456 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_010 | iRM_1 |  |  | 0.00028233 | 6.3079e-05 | 9.4392e-05 | 9.7483e-06 | 0.00010034 | 6.6917e-06 | 0.00014872 | -8.9229e-06 | 4.4083e-06 | [5 empty] | 23.3239 | 5.1801 | 8.0727 | 6.9049e-05 | 8.4929 | 3.5323e-05 | 1.4161 | 9.0487e-06 | 7.3893e-06 |  |  |  | 1 | 0 | 0 | 23.3236 | 5.18 | 8.0726 | 5.93e-05 | 8.4928 | 2.8632e-05 | 1.4159 | 1.7972e-05 | 2.981e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7032 | 23.3236 | 5.18 | 8.0726 | 8.4928 | 1.4159 | 0.22209 | 0.34611 | 0.36413 | 0.060708 | 1.5584 | 1.6395 | 0.27334 | 1.0521 | 0.1754 | 0.16672 | -1.5047 | -1.061 | -1.0102 | -2.8017 | [10 empty] | 0 | 1.7032 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014059 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 103.7% | 2.283 | -0.5952 | -0.082763 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.67% | -9.6411 | -10.4282 | -9.7697 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_015 | iRM_12 |  |  | 0.00028887 | 7.2728e-05 | 0.00010005 | 1.3623e-05 | 0.00010673 | 5.9428e-06 | 0.00013406 | -4.2082e-06 | 6.6263e-06 | [5 empty] | 22.4269 | 4.9808 | 7.7621 | 6.6e-05 | 8.1659 | 3.3895e-05 | 1.3615 | 1.2788e-05 | 9.0441e-06 |  |  |  | 1 | 0 | 0 | 22.4266 | 4.9807 | 7.762 | 5.2376e-05 | 8.1658 | 2.7952e-05 | 1.3614 | 1.6996e-05 | 2.4178e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7021 | 22.4266 | 4.9807 | 7.762 | 8.1658 | 1.3614 | 0.22209 | 0.34611 | 0.36411 | 0.060704 | 1.5584 | 1.6395 | 0.27333 | 1.052 | 0.17539 | 0.16672 | -1.5047 | -1.061 | -1.0103 | -2.8017 | [10 empty] | 0 | 1.7021 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014055 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.1% | -5.8751 | 0.2491 | 18.1341 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.94% | 11.4185 | 6.0451 | -4.5778 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_020 | iRM_2 |  |  | 0.00028928 | 7.3338e-05 | 9.2762e-05 | 2.0431e-05 | 0.00010665 | 5.1816e-06 | 0.00013554 | 8.3929e-07 | 3.8891e-06 | [5 empty] | 23.4339 | 5.2046 | 8.1109 | 5.9154e-05 | 8.5332 | 3.4532e-05 | 1.4228 | 9.1321e-06 | 5.3192e-06 |  |  |  | 1 | 0 | 0 | 23.4336 | 5.2045 | 8.1108 | 3.8722e-05 | 8.5331 | 2.935e-05 | 1.4227 | 8.2928e-06 | 1.4301e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7039 | 23.4336 | 5.2045 | 8.1108 | 8.5331 | 1.4227 | 0.2221 | 0.34612 | 0.36414 | 0.060711 | 1.5584 | 1.6396 | 0.27335 | 1.0521 | 0.1754 | 0.16672 | -1.5046 | -1.061 | -1.0102 | -2.8016 | [10 empty] | 0 | 1.7039 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014058 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 104.9% | 1.8058 | -5.8402 | 6.7952 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 101% | -4.9752 | -0.1426 | 11.9754 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_025 | iRM_12 |  |  | 0.00030243 | 8.5693e-05 | 9.912e-05 | 1.296e-05 | 0.00011211 | 5.1696e-06 | 0.00014264 | 5.8511e-06 | 8.187e-06 | [5 empty] | 22.2573 | 4.9433 | 7.7037 | 6.4804e-05 | 8.1047 | 3.3944e-05 | 1.3513 | 1.2307e-05 | 8.9206e-06 |  |  |  | 1 | 0 | 0 | 22.257 | 4.9432 | 7.7036 | 5.1844e-05 | 8.1046 | 2.8774e-05 | 1.3512 | 6.4555e-06 | 7.3366e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7036 | 22.257 | 4.9432 | 7.7036 | 8.1046 | 1.3512 | 0.2221 | 0.34612 | 0.36414 | 0.060709 | 1.5584 | 1.6396 | 0.27334 | 1.0521 | 0.1754 | 0.16672 | -1.5046 | -1.061 | -1.0102 | -2.8017 | [10 empty] | 0 | 1.7036 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.2495 | 1.0141 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014049 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.5% | 9.1072 | 6.8642 | -20.6873 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.09% | 3.6893 | -3.1906 | -1.2966 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_030 | iRM_10 |  |  | 0.00029141 | 6.8916e-05 | 9.4106e-05 | 2.0242e-05 | 9.1143e-05 | 7.255e-06 | 0.00014738 | 2.8698e-06 | 4.1992e-06 | [5 empty] | 23.285 | 5.1716 | 8.0595 | 6.3008e-05 | 8.4794 | 3.3755e-05 | 1.4139 | 1.1627e-05 | 5.7715e-06 |  |  |  | 1 | 0 | 0 | 23.2847 | 5.1715 | 8.0594 | 4.2765e-05 | 8.4794 | 2.65e-05 | 1.4137 | 8.7568e-06 | 1.5723e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.705 | 23.2847 | 5.1715 | 8.0594 | 8.4794 | 1.4137 | 0.2221 | 0.34612 | 0.36416 | 0.060715 | 1.5584 | 1.6396 | 0.27337 | 1.0521 | 0.17541 | 0.16673 | -1.5046 | -1.061 | -1.0102 | -2.8016 | [10 empty] | 0 | 1.705 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014064 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 104.5% | -0.062223 | -10.9723 | 4.0616 | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 99.77% | -0.3178 | 6.3687 | -8.0645 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_035 | iRM_12 |  |  | 0.00028023 | 6.9008e-05 | 9.4648e-05 | 1.5683e-05 | 0.00010354 | 4.7344e-06 | 0.00014377 | -2.7504e-06 | 9.3398e-06 | [5 empty] | 22.3126 | 4.9556 | 7.7231 | 6.2488e-05 | 8.1256 | 3.428e-05 | 1.3549 | 1.2491e-05 | 7.6638e-06 |  |  |  | 1 | 0 | 0 | 22.3123 | 4.9555 | 7.723 | 4.6805e-05 | 8.1255 | 2.9546e-05 | 1.3548 | 1.5242e-05 | -1.676e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7057 | 22.3123 | 4.9555 | 7.723 | 8.1255 | 1.3548 | 0.2221 | 0.34613 | 0.36417 | 0.060718 | 1.5585 | 1.6397 | 0.27338 | 1.0521 | 0.17542 | 0.16673 | -1.5046 | -1.0609 | -1.0101 | -2.8015 | [10 empty] | 0 | 1.7057 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.89% | -5.2451 | -4.8404 | 16.3557 | [8 empty] | iRM_12 | Lab 1 | Apr 5, 2022 | 3 | 100.5% | -5.5439 | -6.1824 | 20.3843 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_040 | iRM_11 |  |  | 0.00029609 | 7.2243e-05 | 8.9203e-05 | 1.7484e-05 | 9.349e-05 | 7.8528e-06 | 0.00015503 | 1.106e-05 | 5.1044e-06 | [5 empty] | 23.2952 | 5.1739 | 8.0633 | 6.119e-05 | 8.4836 | 3.3361e-05 | 1.4146 | 1.2679e-05 | 3.6982e-06 |  |  |  | 1 | 0 | 0 | 23.2949 | 5.1738 | 8.0632 | 4.3706e-05 | 8.4835 | 2.5508e-05 | 1.4145 | 1.6183e-06 | -1.4061e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7063 | 23.2949 | 5.1738 | 8.0632 | 8.4835 | 1.4145 | 0.2221 | 0.34614 | 0.36418 | 0.060721 | 1.5585 | 1.6397 | 0.27339 | 1.0521 | 0.17542 | 0.16673 | -1.5046 | -1.0609 | -1.0101 | -2.8015 | [10 empty] | 0 | 1.7063 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014058 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 104.2% | 6.9019 | -1.1684 | 20.1471
Lab 1 | Jun 20, 2022 | 2 | Zr_045 | iRM_12 |  |  | 0.00067454 | 0.00015807 | 0.00022367 | 1.1288e-05 | 0.00026064 | 5.4672e-06 | 0.00017141 | 1.9342e-06 | 5.0529e-06 | [5 empty] | 22.4171 | 4.9789 | 7.7595 | 6.178e-05 | 8.1641 | 3.2289e-05 | 1.3613 | 1.5029e-05 | 6.7703e-06 |  |  |  | 1 | 0 | 0 | 22.4164 | 4.9787 | 7.7592 | 5.0492e-05 | 8.1638 | 2.6822e-05 | 1.3612 | 1.3095e-05 | 1.7174e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7069 | 22.4164 | 4.9787 | 7.7592 | 8.1638 | 1.3612 | 0.2221 | 0.34614 | 0.36419 | 0.060722 | 1.5585 | 1.6397 | 0.2734 | 1.0521 | 0.17542 | 0.16673 | -1.5046 | -1.0609 | -1.0101 | -2.8015 | [10 empty] | 0 | 1.7069 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014056 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.3% | -4.749 | -0.58938 | -14.265 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.51% | -3.7238 | -0.81394 | -10.2838 | iRM_12 | Avg | -0.17555 | 0.050546 | 0.21684
Lab 1 | Jun 20, 2022 | 2 | Zr_050 | iRM |  |  | 0.00027614 | 5.5225e-05 | 9.2533e-05 | 5.8669e-06 | 9.9641e-05 | 7.582e-06 | 0.00013045 | 1.4935e-05 | 4.1482e-06 | [5 empty] | 23.4969 | 5.2187 | 8.1332 | 5.7868e-05 | 8.5574 | 3.3866e-05 | 1.4269 | 1.4208e-05 | 4.4219e-06 |  |  |  | 1 | 0 | 0 | 23.4966 | 5.2187 | 8.1331 | 5.2001e-05 | 8.5573 | 2.6284e-05 | 1.4268 | -7.2704e-07 | 2.737e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.707 | 23.4966 | 5.2187 | 8.1331 | 8.5573 | 1.4268 | 0.2221 | 0.34614 | 0.36419 | 0.060722 | 1.5585 | 1.6397 | 0.27339 | 1.0521 | 0.17542 | 0.16673 | -1.5046 | -1.0609 | -1.0101 | -2.8015 | [10 empty] | 0 | 1.707 | 0.21795 | 0.33338 | 0.3381 | 0.054378 | 1.5296 | 1.5513 | 0.2495 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014062 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 2 | 104.9% | 3.7475 | -7.8556 | -16.8316 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.96% | 11.5767 | -1.3464 | 8.4289 | iRM_12 | 2SD | 13.8686 | 10.1618 | 27.1314
Lab 1 | Jun 20, 2022 | 2 | Zr_055 | iRM_12 |  |  | 0.00028004 | 5.6606e-05 | 9.0859e-05 | 1.6457e-05 | 9.9521e-05 | 4.1997e-06 | 0.00013973 | 1.2305e-05 | 1.6374e-06 | [5 empty] | 22.3821 | 4.9711 | 7.7474 | 6.3526e-05 | 8.1514 | 3.4233e-05 | 1.3592 | 1.5669e-05 | 7.243e-06 |  |  |  | 1 | 0 | 0 | 22.3818 | 4.9711 | 7.7473 | 4.7069e-05 | 8.1513 | 3.0033e-05 | 1.3591 | 3.3644e-06 | 5.6055e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7072 | 22.3818 | 4.9711 | 7.7473 | 8.1513 | 1.3591 | 0.2221 | 0.34614 | 0.36419 | 0.060724 | 1.5585 | 1.6397 | 0.2734 | 1.0521 | 0.17543 | 0.16673 | -1.5046 | -1.0609 | -1.0101 | -2.8014 | [10 empty] | 0 | 1.7072 | 0.21795 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014053 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.2% | 9.529 | 6.4929 | 1.1168 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.9% | -5.7185 | 1.3635 | -3.3013 | iRM_12 | n | 151 | 151 | 151
Lab 1 | Jun 20, 2022 | 2 | Zr_060 | iRM_1 |  |  | 0.00028371 | 3.0026e-05 | 0.00010471 | 8.7156e-06 | 0.00010078 | 7.9819e-06 | 0.00014213 | 1.1864e-05 | 7.2893e-06 | [5 empty] | 23.2763 | 5.1698 | 8.0573 | 6.1154e-05 | 8.478 | 3.4604e-05 | 1.4138 | 1.5552e-05 | 7.6823e-06 |  |  |  | 1 | 0 | 0 | 23.276 | 5.1698 | 8.0572 | 5.2439e-05 | 8.4779 | 2.6623e-05 | 1.4136 | 3.6879e-06 | 3.9292e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7096 | 23.276 | 5.1698 | 8.0572 | 8.4779 | 1.4136 | 0.22211 | 0.34616 | 0.36423 | 0.060733 | 1.5585 | 1.6399 | 0.27344 | 1.0522 | 0.17545 | 0.16674 | -1.5046 | -1.0609 | -1.01 | -2.8013 | [10 empty] | 0 | 1.7096 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014064 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 104.3% | -0.38187 | -6.5961 | -12.7188 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.2% | -0.77211 | -0.9106 | -1.1851 | iRM_12 | 2SE | 1.1286 | 0.82695 | 2.2079
Lab 1 | Jun 20, 2022 | 2 | Zr_065 | iRM_12 |  |  | 0.0002745 | 6.468e-05 | 9.4232e-05 | 2.3136e-05 | 0.00010606 | 9.6744e-06 | 0.00013024 | 1.4626e-05 | 7.9413e-06 | [5 empty] | 22.2533 | 4.9426 | 7.7029 | 6.0079e-05 | 8.1049 | 3.4078e-05 | 1.3515 | 2.0455e-05 | 6.6091e-06 |  |  |  | 1 | 0 | 0 | 22.253 | 4.9425 | 7.7029 | 3.6943e-05 | 8.1048 | 2.4403e-05 | 1.3514 | 5.8294e-06 | -1.3323e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7083 | 22.253 | 4.9425 | 7.7029 | 8.1048 | 1.3514 | 0.2221 | 0.34615 | 0.36421 | 0.060728 | 1.5585 | 1.6398 | 0.27342 | 1.0522 | 0.17544 | 0.16674 | -1.5046 | -1.0609 | -1.01 | -2.8013 | [10 empty] | 0 | 1.7083 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014063 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.92% | -3.9396 | -7.7846 | 9.2245 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.85% | -0.47219 | 3.0238 | 8.611 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_070 | iRM_2 |  |  | 0.00029134 | 6.2811e-05 | 8.9053e-05 | 2.1629e-05 | 0.0001048 | 5.8873e-06 | 0.00014316 | 1.2965e-05 | 7.9797e-06 | [5 empty] | 23.6331 | 5.2492 | 8.1811 | 5.9361e-05 | 8.6084 | 3.4817e-05 | 1.4356 | 1.4894e-05 | 6.1958e-06 |  |  |  | 1 | 0 | 0 | 23.6328 | 5.2491 | 8.181 | 3.7732e-05 | 8.6083 | 2.893e-05 | 1.4354 | 1.9289e-06 | -1.7839e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.711 | 23.6328 | 5.2491 | 8.181 | 8.6083 | 1.4354 | 0.22211 | 0.34617 | 0.36425 | 0.060738 | 1.5585 | 1.64 | 0.27346 | 1.0522 | 0.17546 | 0.16675 | -1.5046 | -1.0608 | -1.0099 | -2.8012 | [10 empty] | 0 | 1.711 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014058 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 106.4% | 3.0505 | 1.7192 | -14.1487 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.5% | -7.6963 | -5.7943 | -18.0488 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_075 | iRM_12 |  |  | 0.00028628 | 5.5737e-05 | 9.2214e-05 | 1.1042e-05 | 9.9964e-05 | 6.5544e-06 | 0.0001273 | 9.9806e-06 | 3.1529e-06 | [5 empty] | 22.1584 | 4.9216 | 7.6706 | 5.7985e-05 | 8.0712 | 3.4707e-05 | 1.346 | 1.3889e-05 | 7.1975e-06 |  |  |  | 1 | 0 | 0 | 22.1581 | 4.9215 | 7.6705 | 4.6943e-05 | 8.0711 | 2.8152e-05 | 1.3458 | 3.9083e-06 | 4.0446e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7107 | 22.1581 | 4.9215 | 7.6705 | 8.0711 | 1.3458 | 0.22211 | 0.34617 | 0.36425 | 0.060737 | 1.5585 | 1.6399 | 0.27346 | 1.0522 | 0.17546 | 0.16675 | -1.5046 | -1.0608 | -1.0099 | -2.8012 | [10 empty] | 0 | 1.7107 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.31% | 2.0424 | 4.3806 | -1.334 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 98.98% | 13.401 | 4.258 | 15.3195 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_080 | iRM_10 |  |  | 0.00028758 | 5.4725e-05 | 9.3213e-05 | 1.4341e-05 | 0.00010366 | 3.0385e-06 | 0.00012638 | 5.334e-07 | 9.8325e-06 | [5 empty] | 23.1338 | 5.1383 | 8.0082 | 5.98e-05 | 8.4265 | 3.3441e-05 | 1.4052 | 9.7105e-06 | 6.5205e-06 |  |  |  | 1 | 0 | 0 | 23.1335 | 5.1382 | 8.0081 | 4.5459e-05 | 8.4264 | 3.0403e-05 | 1.4051 | 9.1771e-06 | -3.312e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7108 | 23.1335 | 5.1382 | 8.0081 | 8.4264 | 1.4051 | 0.22211 | 0.34617 | 0.36425 | 0.060739 | 1.5585 | 1.6399 | 0.27346 | 1.0522 | 0.17546 | 0.16675 | -1.5046 | -1.0608 | -1.0099 | -2.8012 | [10 empty] | 0 | 1.7108 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3882 | 0.014061 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 103.9% | 7.3238 | -1.8952 | 10.3072 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.8% | -7.1557 | -2.705 | -11.1214 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_085 | iRM_12 |  |  | 0.00028507 | 6.5692e-05 | 9.4786e-05 | 7.2027e-06 | 0.00010809 | 5.61e-06 | 0.00013077 | 1.2959e-05 | 4.869e-06 | [5 empty] | 22.3713 | 4.9689 | 7.7444 | 6.3622e-05 | 8.1491 | 3.3568e-05 | 1.359 | 1.0633e-05 | 6.9668e-06 |  |  |  | 1 | 0 | 0 | 22.371 | 4.9689 | 7.7443 | 5.6419e-05 | 8.149 | 2.7958e-05 | 1.3588 | -2.3256e-06 | 2.0977e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7117 | 22.371 | 4.9689 | 7.7443 | 8.149 | 1.3588 | 0.22211 | 0.34618 | 0.36426 | 0.060741 | 1.5586 | 1.64 | 0.27347 | 1.0523 | 0.17546 | 0.16675 | -1.5046 | -1.0608 | -1.0099 | -2.8011 | [10 empty] | 0 | 1.7117 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.01406 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.3% | -5.1855 | -0.37098 | 3.2148 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.69% | 3.8723 | 3.7446 | 3.1252 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_090 | iRM_11 |  |  | 0.00028722 | 7.5868e-05 | 0.00010507 | 1.0376e-05 | 9.1859e-05 | 8.4615e-06 | 0.00014114 | -1.7939e-06 | 4.5069e-06 | [5 empty] | 23.3518 | 5.1868 | 8.084 | 6.0344e-05 | 8.5067 | 3.2434e-05 | 1.4186 | 8.2172e-06 | 2.896e-06 |  |  |  | 1 | 0 | 0 | 23.3515 | 5.1867 | 8.0839 | 4.9968e-05 | 8.5066 | 2.3973e-05 | 1.4185 | 1.0011e-05 | -1.6109e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7129 | 23.3515 | 5.1867 | 8.0839 | 8.5066 | 1.4185 | 0.22211 | 0.34618 | 0.36428 | 0.060745 | 1.5586 | 1.6401 | 0.27349 | 1.0523 | 0.17547 | 0.16675 | -1.5046 | -1.0608 | -1.0098 | -2.8011 | [10 empty] | 0 | 1.7129 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014068 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 104.2% | -2.1442 | -6.5936 | -6.704 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.4% | -3.6809 | -2.0788 | 3.8565 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_095 | iRM_12 |  |  | 0.00028641 | 7.21e-05 | 0.00010542 | 8.1873e-06 | 0.00011225 | 5.7465e-06 | 0.00014192 | 3.982e-06 | 2.3735e-06 | [5 empty] | 22.4641 | 4.9896 | 7.7767 | 6.2083e-05 | 8.1832 | 3.4286e-05 | 1.3647 | 1.3168e-05 | 7.0721e-06 |  |  |  | 1 | 0 | 0 | 22.4638 | 4.9895 | 7.7766 | 5.3896e-05 | 8.1831 | 2.8539e-05 | 1.3645 | 9.1864e-06 | 4.6986e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7125 | 22.4638 | 4.9895 | 7.7766 | 8.1831 | 1.3645 | 0.22211 | 0.34618 | 0.36428 | 0.060744 | 1.5586 | 1.64 | 0.27348 | 1.0523 | 0.17547 | 0.16675 | -1.5046 | -1.0608 | -1.0098 | -2.8011 | [10 empty] | 0 | 1.7125 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014063 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 101% | -2.1665 | -3.7596 | -22.1405 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.4% | 1.7402 | -0.38001 | 15.5945 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_100 | iRM |  |  | 0.00092079 | 0.00021218 | 0.00031499 | 1.4767e-05 | 0.00033764 | 3.4199e-06 | 0.00017495 | 6.0818e-06 | 5.5249e-06 | [5 empty] | 23.6242 | 5.2474 | 8.1786 | 5.8589e-05 | 8.6065 | 3.334e-05 | 1.4353 | 1.4815e-05 | 4.4873e-06 |  |  |  | 1 | 0 | 0 | 23.6233 | 5.2472 | 8.1783 | 4.3822e-05 | 8.6062 | 2.992e-05 | 1.4352 | 8.7331e-06 | -1.0376e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7145 | 23.6233 | 5.2472 | 8.1783 | 8.6062 | 1.4352 | 0.22212 | 0.3462 | 0.36431 | 0.060752 | 1.5586 | 1.6401 | 0.27351 | 1.0523 | 0.17548 | 0.16676 | -1.5045 | -1.0607 | -1.0098 | -2.801 | [10 empty] | 0 | 1.7145 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3883 | 0.014064 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 2 | 106% | -1.4952 | -3.3985 | -13.2823 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.2% | -1.7481 | 2.5403 | -26.9663 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_105 | iRM_12 |  |  | 0.00027823 | 5.2035e-05 | 0.00010411 | 1.4869e-05 | 0.00012422 | 7.2767e-06 | 0.00012612 | 2.6296e-06 | 8.6088e-06 | [5 empty] | 22.0921 | 4.907 | 7.648 | 5.5585e-05 | 8.0479 | 3.4268e-05 | 1.3422 | 8.2253e-06 | 6.7376e-06 |  |  |  | 1 | 0 | 0 | 22.0918 | 4.907 | 7.6479 | 4.0715e-05 | 8.0478 | 2.6991e-05 | 1.342 | 5.5957e-06 | -1.8711e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7132 | 22.0918 | 4.907 | 7.6479 | 8.0478 | 1.342 | 0.22212 | 0.34619 | 0.36429 | 0.060748 | 1.5586 | 1.6401 | 0.2735 | 1.0523 | 0.17548 | 0.16676 | -1.5045 | -1.0608 | -1.0098 | -2.801 | [10 empty] | 0 | 1.7132 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3882 | 0.014058 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99% | 6.9038 | 2.4545 | 23.7328 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.4% | 6.0205 | -2.7476 | 4.0694 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_110 | iRM_1 |  |  | 0.00034905 | 8.8228e-05 | 0.00012224 | 1.8408e-05 | 0.00014085 | 7.1314e-06 | 0.00013025 | 9.8336e-06 | 4.498e-06 | [5 empty] | 23.5642 | 5.2342 | 8.1581 | 6.2772e-05 | 8.5852 | 3.5002e-05 | 1.4318 | 1.2152e-05 | 7.8775e-06 |  |  |  | 1 | 0 | 0 | 23.5639 | 5.2341 | 8.158 | 4.4364e-05 | 8.585 | 2.7871e-05 | 1.4317 | 2.3182e-06 | 3.3796e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7158 | 23.5639 | 5.2341 | 8.158 | 8.585 | 1.4317 | 0.22212 | 0.34621 | 0.36433 | 0.060758 | 1.5586 | 1.6402 | 0.27353 | 1.0523 | 0.1755 | 0.16677 | -1.5045 | -1.0607 | -1.0097 | -2.8009 | [10 empty] | 0 | 1.7158 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014061 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 106.5% | -3.5822 | -2.7739 | -5.2892 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.64% | -11.4069 | -2.8357 | 11.2368 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_115 | iRM_12 |  |  | 0.00027581 | 7.0544e-05 | 9.7679e-05 | 8.6827e-06 | 0.00010437 | 2.3819e-06 | 0.00012678 | -1.0986e-06 | 3.674e-06 | [5 empty] | 22.1648 | 4.9233 | 7.6735 | 6.3398e-05 | 8.075 | 3.3779e-05 | 1.3467 | 1.0699e-05 | 8.2232e-06 |  |  |  | 1 | 0 | 0 | 22.1646 | 4.9232 | 7.6734 | 5.4716e-05 | 8.0749 | 3.1397e-05 | 1.3466 | 1.1797e-05 | 4.5492e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.715 | 22.1646 | 4.9232 | 7.6734 | 8.0749 | 1.3466 | 0.22212 | 0.3462 | 0.36432 | 0.060755 | 1.5586 | 1.6402 | 0.27352 | 1.0523 | 0.17549 | 0.16676 | -1.5045 | -1.0607 | -1.0097 | -2.8009 | [10 empty] | 0 | 1.715 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014058 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.98% | -1.4264 | -1.1333 | -5.6603 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.1% | 10.5155 | 6.355 | 3.4458 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_120 | iRM_2 |  |  | 0.00028243 | 7.2148e-05 | 9.3752e-05 | 6.7014e-06 | 0.00011008 | 5.6719e-06 | 0.00014295 | 7.0986e-06 | 4.32e-06 | [5 empty] | 23.7313 | 5.2713 | 8.2159 | 6.0406e-05 | 8.6459 | 3.4949e-05 | 1.4419 | 1.6552e-05 | 3.6055e-06 |  |  |  | 1 | 0 | 0 | 23.731 | 5.2712 | 8.2158 | 5.3705e-05 | 8.6458 | 2.9277e-05 | 1.4418 | 9.4537e-06 | -7.1448e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7154 | 23.731 | 5.2712 | 8.2158 | 8.6458 | 1.4418 | 0.22212 | 0.34621 | 0.36432 | 0.060756 | 1.5586 | 1.6402 | 0.27352 | 1.0523 | 0.17549 | 0.16676 | -1.5045 | -1.0607 | -1.0097 | -2.8009 | [10 empty] | 0 | 1.7154 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 106.9% | 0.43074 | -0.8313 | -7.8614 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.91% | 0.6426 | -2.6768 | -17.0925 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_125 | iRM_12 |  |  | 0.00034764 | 8.7709e-05 | 0.00011301 | 1.0133e-05 | 0.00012452 | 3.3205e-06 | 0.00013879 | 5.9139e-06 | 6.1233e-06 | [5 empty] | 22.2457 | 4.9414 | 7.7019 | 5.981e-05 | 8.1052 | 3.4521e-05 | 1.3518 | 1.1745e-05 | 7.4537e-06 |  |  |  | 1 | 0 | 0 | 22.2453 | 4.9413 | 7.7017 | 4.9677e-05 | 8.1051 | 3.12e-05 | 1.3517 | 5.8306e-06 | 1.3303e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.717 | 22.2453 | 4.9413 | 7.7017 | 8.1051 | 1.3517 | 0.22213 | 0.34622 | 0.36435 | 0.060762 | 1.5587 | 1.6403 | 0.27355 | 1.0524 | 0.1755 | 0.16677 | -1.5045 | -1.0607 | -1.0096 | -2.8008 | [10 empty] | 0 | 1.717 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014055 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.2% | 3.6163 | 5.5873 | 3.2154 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.1% | -8.4934 | -1.5922 | 7.9727 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_130 | iRM_10 |  |  | 0.00028477 | 7.908e-05 | 9.8301e-05 | 1.8348e-05 | 0.000105 | 1.4031e-06 | 0.00013606 | 1.014e-05 | 5.6415e-06 | [5 empty] | 23.2681 | 5.1685 | 8.0558 | 5.355e-05 | 8.4778 | 3.5135e-05 | 1.414 | 1.1122e-05 | 5.5226e-06 |  |  |  | 1 | 0 | 0 | 23.2679 | 5.1684 | 8.0557 | 3.5202e-05 | 8.4777 | 3.3732e-05 | 1.4138 | 9.8255e-07 | -1.1887e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7173 | 23.2679 | 5.1684 | 8.0557 | 8.4777 | 1.4138 | 0.22213 | 0.34622 | 0.36435 | 0.060763 | 1.5587 | 1.6403 | 0.27355 | 1.0524 | 0.1755 | 0.16677 | -1.5045 | -1.0607 | -1.0096 | -2.8008 | [10 empty] | 0 | 1.7173 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3883 | 0.014067 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 104.6% | -1.0247 | -7.6783 | -11.5869 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.2% | 2.7902 | -1.8929 | 1.8057 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_135 | iRM_12 |  |  | 0.00028881 | 5.7323e-05 | 9.9111e-05 | 6.086e-06 | 0.00010723 | 1.8749e-06 | 0.00013782 | 5.1343e-06 | 5.4883e-06 | [5 empty] | 22.2586 | 4.9442 | 7.7062 | 6.0962e-05 | 8.1098 | 3.2597e-05 | 1.3526 | 8.939e-06 | 8.0819e-06 |  |  |  | 1 | 0 | 0 | 22.2583 | 4.9441 | 7.7061 | 5.4876e-05 | 8.1097 | 3.0723e-05 | 1.3524 | 3.8047e-06 | 2.5936e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7166 | 22.2583 | 4.9441 | 7.7061 | 8.1097 | 1.3524 | 0.22212 | 0.34621 | 0.36434 | 0.06076 | 1.5587 | 1.6403 | 0.27354 | 1.0524 | 0.1755 | 0.16677 | -1.5045 | -1.0607 | -1.0097 | -2.8008 | [10 empty] | 0 | 1.7166 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014064 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.6% | -6.6397 | -5.3088 | -9.0591 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 99.47% | 6.5834 | 8.1805 | 3.3359 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_140 | iRM_11 |  |  | 0.00028547 | 7.9056e-05 | 9.8823e-05 | 1.4431e-05 | 0.00011234 | 8.2111e-06 | 0.0001145 | 9.5186e-06 | 8.4506e-06 | [5 empty] | 23.4532 | 5.2097 | 8.1202 | 6.0019e-05 | 8.5459 | 3.3355e-05 | 1.4254 | 1.2079e-05 | 3.1104e-06 |  |  |  | 1 | 0 | 0 | 23.4529 | 5.2096 | 8.1201 | 4.5588e-05 | 8.5458 | 2.5144e-05 | 1.4252 | 2.5604e-06 | -5.3402e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7189 | 23.4529 | 5.2096 | 8.1201 | 8.5458 | 1.4252 | 0.22213 | 0.34623 | 0.36438 | 0.06077 | 1.5587 | 1.6404 | 0.27358 | 1.0524 | 0.17552 | 0.16678 | -1.5045 | -1.0607 | -1.0096 | -2.8007 | [10 empty] | 0 | 1.719 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014065 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 105.9% | 0.34907 | -1.7504 | 6.4978 | iRM_12 | Lab 1 | Jun 20, 2022 | 1 | 100.4% | -10.5532 | -6.6985 | 3.7208 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_145 | iRM_12 |  |  | 0.00028422 | 7.327e-05 | 9.6486e-05 | 1.4482e-05 | 0.00010575 | 7.6311e-06 | 9.7659e-05 | 8.6008e-06 | 8.811e-06 | [5 empty] | 22.0174 | 4.8907 | 7.623 | 5.885e-05 | 8.0224 | 3.1115e-05 | 1.338 | 1.313e-05 | 9.2925e-06 |  |  |  | 1 | 0 | 0 | 22.0172 | 4.8906 | 7.6229 | 4.4368e-05 | 8.0223 | 2.3484e-05 | 1.3379 | 4.5297e-06 | 4.8154e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7181 | 22.0172 | 4.8906 | 7.6229 | 8.0223 | 1.3379 | 0.22213 | 0.34622 | 0.36437 | 0.060767 | 1.5587 | 1.6403 | 0.27357 | 1.0524 | 0.17551 | 0.16677 | -1.5045 | -1.0607 | -1.0096 | -2.8007 | [10 empty] | 0 | 1.7181 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014062 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.23% | -2.1368 | -1.8507 | 7.1034 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.1% | -5.8751 | 0.2491 | 18.1341 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_150 | iRM |  |  | 0.00028563 | 7.3913e-05 | 9.7267e-05 | 8.7215e-06 | 0.00010623 | 7.2945e-06 | 9.9838e-05 | 3.3616e-06 | 6.7937e-06 | [5 empty] | 23.7074 | 5.2661 | 8.2082 | 5.8096e-05 | 8.6385 | 3.2803e-05 | 1.4408 | 8.7282e-06 | 3.809e-06 |  |  |  | 1 | 0 | 0 | 23.7071 | 5.266 | 8.2081 | 4.9374e-05 | 8.6384 | 2.5508e-05 | 1.4407 | 5.3666e-06 | -2.9847e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.719 | 23.7071 | 5.266 | 8.2081 | 8.6384 | 1.4407 | 0.22213 | 0.34623 | 0.36438 | 0.06077 | 1.5587 | 1.6404 | 0.27358 | 1.0524 | 0.17552 | 0.16678 | -1.5045 | -1.0607 | -1.0096 | -2.8007 | [10 empty] | 0 | 1.719 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 2 | 107.4% | -8.6617 | -11.1932 | 2.1736 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.5% | 9.1072 | 6.8642 | -20.6873 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_155 | iRM_12 |  |  | 0.00029044 | 8.0344e-05 | 9.2156e-05 | 1.1291e-05 | 0.0001116 | 5.157e-06 | 0.00010343 | 1.3231e-05 | 5.996e-06 | [5 empty] | 22.1198 | 4.9136 | 7.6587 | 6.0682e-05 | 8.0602 | 3.224e-05 | 1.3444 | 7.6173e-06 | 7.463e-06 |  |  |  | 1 | 0 | 0 | 22.1195 | 4.9135 | 7.6586 | 4.9391e-05 | 8.0601 | 2.7083e-05 | 1.3442 | -5.6141e-06 | 1.467e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7195 | 22.1195 | 4.9135 | 7.6586 | 8.0601 | 1.3442 | 0.22213 | 0.34624 | 0.36439 | 0.060772 | 1.5587 | 1.6404 | 0.27358 | 1.0524 | 0.17552 | 0.16678 | -1.5045 | -1.0606 | -1.0095 | -2.8006 | [10 empty] | 0 | 1.7195 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.75% | 6.7979 | 6.4877 | -0.41843 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.89% | -5.2451 | -4.8404 | 16.3557 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_160 | iRM_1 |  |  | 0.00028072 | 6.8755e-05 | 0.00010409 | 1.0892e-05 | 0.0001042 | 3.2401e-06 | 9.2003e-05 | -3.0958e-06 | 6.1062e-06 | [5 empty] | 23.4815 | 5.216 | 8.1302 | 6.4589e-05 | 8.5565 | 3.5517e-05 | 1.4271 | 9.3394e-06 | 7.0649e-06 |  |  |  | 1 | 0 | 0 | 23.4812 | 5.2159 | 8.1301 | 5.3697e-05 | 8.5564 | 3.2277e-05 | 1.4271 | 1.2435e-05 | 9.5869e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7199 | 23.4812 | 5.2159 | 8.1301 | 8.5564 | 1.4271 | 0.22213 | 0.34624 | 0.36439 | 0.060774 | 1.5587 | 1.6404 | 0.2736 | 1.0524 | 0.17553 | 0.16678 | -1.5045 | -1.0606 | -1.0095 | -2.8006 | [10 empty] | 0 | 1.7199 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014067 | -1.8133 | -1.8273 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 105.6% | -8.2727 | -6.3618 | 16.8375 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.3% | -4.749 | -0.58938 | -14.265 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_165 | iRM_12 |  |  | 0.00036833 | 8.751e-05 | 0.00012707 | 1.9489e-05 | 0.00013993 | 9.578e-06 | 9.4299e-05 | 4.1057e-06 | 6.2451e-06 | [5 empty] | 22.3343 | 4.9612 | 7.7331 | 6.5428e-05 | 8.1387 | 3.3059e-05 | 1.3575 | 1.0902e-05 | 7.8132e-06 |  |  |  | 1 | 0 | 0 | 22.3339 | 4.9611 | 7.7329 | 4.594e-05 | 8.1386 | 2.3481e-05 | 1.3574 | 6.7963e-06 | 1.5681e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7205 | 22.3339 | 4.9611 | 7.7329 | 8.1386 | 1.3574 | 0.22213 | 0.34624 | 0.3644 | 0.060776 | 1.5587 | 1.6405 | 0.2736 | 1.0525 | 0.17553 | 0.16678 | -1.5045 | -1.0606 | -1.0095 | -2.8006 | [10 empty] | 0 | 1.7205 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3883 | 0.014064 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.8% | -3.1634 | -6.4499 | -10.2962 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.2% | 9.529 | 6.4929 | 1.1168 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_170 | iRM_2 |  |  | 0.0002868 | 5.052e-05 | 0.00010449 | 2.2274e-05 | 0.00010801 | 4.8461e-06 | 9.3223e-05 | -4.922e-07 | 4.5473e-06 | [5 empty] | 23.5773 | 5.2374 | 8.1636 | 6.6498e-05 | 8.592 | 3.2816e-05 | 1.4331 | 6.8822e-06 | 5.1725e-06 |  |  |  | 1 | 0 | 0 | 23.577 | 5.2373 | 8.1635 | 4.4224e-05 | 8.5919 | 2.797e-05 | 1.433 | 7.3744e-06 | 6.2516e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7214 | 23.577 | 5.2373 | 8.1635 | 8.5919 | 1.433 | 0.22214 | 0.34625 | 0.36442 | 0.06078 | 1.5587 | 1.6405 | 0.27362 | 1.0525 | 0.17554 | 0.16679 | -1.5045 | -1.0606 | -1.0095 | -2.8005 | [10 empty] | 0 | 1.7214 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014061 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 105.9% | -1.1484 | 0.9976 | 5.8119 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.92% | -3.9396 | -7.7846 | 9.2245 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_175 | iRM_12 |  |  | 0.0010186 | 0.00023554 | 0.00035049 | 8.1092e-06 | 0.00036991 | 4.5023e-06 | 0.00011795 | 1.1114e-05 | 6.6879e-06 | [5 empty] | 22.2058 | 4.9327 | 7.6886 | 5.9361e-05 | 8.0919 | 3.192e-05 | 1.3497 | 1.0727e-05 | 6.6915e-06 |  |  |  | 1 | 0 | 0 | 22.2048 | 4.9325 | 7.6883 | 5.1251e-05 | 8.0915 | 2.7418e-05 | 1.3495 | -3.87e-07 | 3.6327e-09 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7205 | 22.2048 | 4.9325 | 7.6883 | 8.0915 | 1.3495 | 0.22213 | 0.34624 | 0.3644 | 0.060777 | 1.5587 | 1.6405 | 0.2736 | 1.0525 | 0.17553 | 0.16678 | -1.5045 | -1.0606 | -1.0095 | -2.8005 | [10 empty] | 0 | 1.7205 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014059 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.1% | 4.8643 | 11.158 | 18.878 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.31% | 2.0424 | 4.3806 | -1.334 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_180 | iRM_10 |  |  | 0.00028767 | 6.2049e-05 | 9.5635e-05 | 2.0242e-05 | 0.00010028 | 7.2306e-06 | 0.00011172 | 1.0928e-05 | 3.4521e-06 | [5 empty] | 23.4997 | 5.2202 | 8.1371 | 6.2033e-05 | 8.5644 | 3.2114e-05 | 1.4286 | 8.2919e-06 | 4.5755e-06 |  |  |  | 1 | 0 | 0 | 23.4994 | 5.2202 | 8.137 | 4.1791e-05 | 8.5643 | 2.4884e-05 | 1.4284 | -2.6364e-06 | 1.1234e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7231 | 23.4994 | 5.2202 | 8.137 | 8.5643 | 1.4284 | 0.22214 | 0.34626 | 0.36445 | 0.060786 | 1.5588 | 1.6406 | 0.27364 | 1.0525 | 0.17555 | 0.16679 | -1.5044 | -1.0606 | -1.0094 | -2.8004 | [10 empty] | 0 | 1.7231 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014057 | -1.8133 | -1.8274 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 106.2% | -0.36255 | 10.1373 | -5.6021 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.3% | -5.1855 | -0.37098 | 3.2148 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_185 | iRM_12 |  |  | 0.00028259 | 6.4523e-05 | 9.9756e-05 | 1.3876e-05 | 9.7932e-05 | 4.8062e-06 | 9.2081e-05 | 7.7014e-06 | 4.4379e-06 | [5 empty] | 22.0366 | 4.8951 | 7.6302 | 5.7602e-05 | 8.0309 | 3.0705e-05 | 1.3395 | 1.1625e-05 | 6.7896e-06 |  |  |  | 1 | 0 | 0 | 22.0363 | 4.8951 | 7.6301 | 4.3726e-05 | 8.0308 | 2.5898e-05 | 1.3394 | 3.9234e-06 | 2.3517e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7223 | 22.0363 | 4.8951 | 7.6301 | 8.0308 | 1.3394 | 0.22214 | 0.34625 | 0.36443 | 0.060783 | 1.5587 | 1.6406 | 0.27363 | 1.0525 | 0.17554 | 0.16679 | -1.5045 | -1.0606 | -1.0094 | -2.8005 | [10 empty] | 0 | 1.7223 | 0.21794 | 0.33337 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3883 | 0.014076 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.06% | -2.8093 | -15.4027 | -16.1787 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 101% | -2.1665 | -3.7596 | -22.1405 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_190 | iRM_11 |  |  | 0.00028721 | 6.9005e-05 | 0.00010066 | 1.3857e-05 | 0.00011221 | 8.4132e-06 | 8.3144e-05 | 2.0631e-06 | 5.4764e-06 | [5 empty] | 23.5681 | 5.2354 | 8.1607 | 5.6131e-05 | 8.5893 | 3.2397e-05 | 1.4327 | 7.0944e-06 | 4.4142e-06 |  |  |  | 1 | 0 | 0 | 23.5678 | 5.2353 | 8.1606 | 4.2274e-05 | 8.5892 | 2.3983e-05 | 1.4326 | 5.0313e-06 | -1.0621e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.723 | 23.5678 | 5.2353 | 8.1606 | 8.5892 | 1.4326 | 0.22214 | 0.34626 | 0.36444 | 0.060787 | 1.5588 | 1.6406 | 0.27364 | 1.0525 | 0.17555 | 0.16679 | -1.5044 | -1.0606 | -1.0094 | -2.8004 | [10 empty] | 0 | 1.723 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.01406 | -1.8133 | -1.8274 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 106.3% | 4.9269 | 8.3359 | 14.4127 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99% | 6.9038 | 2.4545 | 23.7328 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_195 | iRM_12 |  |  | 0.0002895 | 7.5315e-05 | 9.8928e-05 | 7.9382e-06 | 0.0001049 | 1.0272e-05 | 9.7745e-05 | 3.9616e-06 | 6.5555e-06 | [5 empty] | 22.2865 | 4.9508 | 7.7172 | 5.3992e-05 | 8.1226 | 3.3826e-05 | 1.3549 | 5.8941e-06 | 7.0851e-06 |  |  |  | 1 | 0 | 0 | 22.2862 | 4.9507 | 7.7171 | 4.6054e-05 | 8.1225 | 2.3555e-05 | 1.3548 | 1.9325e-06 | 5.2956e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7242 | 22.2862 | 4.9507 | 7.7171 | 8.1225 | 1.3548 | 0.22214 | 0.34627 | 0.36446 | 0.06079 | 1.5588 | 1.6407 | 0.27366 | 1.0525 | 0.17556 | 0.16679 | -1.5044 | -1.0605 | -1.0093 | -2.8003 | [10 empty] | 0 | 1.7242 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014062 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.5% | -5.3999 | 8.0389 | -1.0092 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.98% | -1.4264 | -1.1333 | -5.6603 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_200 | iRM |  |  | 0.00088941 | 0.00020074 | 0.00030965 | 1.0925e-05 | 0.00033192 | 4.2111e-06 | 0.00016044 | 7.1616e-06 | 2.4283e-06 | [5 empty] | 23.8353 | 5.2949 | 8.2536 | 5.3963e-05 | 8.6872 | 4.0475e-05 | 1.4491 | 1.8716e-06 | 5.526e-06 |  |  |  | 1 | 0 | 0 | 23.8344 | 5.2947 | 8.2532 | 4.3038e-05 | 8.6869 | 3.6264e-05 | 1.4489 | -5.29e-06 | 3.0978e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7246 | 23.8344 | 5.2947 | 8.2532 | 8.6869 | 1.4489 | 0.22214 | 0.34627 | 0.36447 | 0.060792 | 1.5588 | 1.6407 | 0.27366 | 1.0525 | 0.17556 | 0.1668 | -1.5044 | -1.0605 | -1.0093 | -2.8003 | [10 empty] | 0 | 1.7246 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014061 | -1.8133 | -1.8274 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 2 | 106.9% | 0.63437 | 1.6652 | -9.1426 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.2% | 3.6163 | 5.5873 | 3.2154 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_205 | iRM_12 |  |  | 0.00028513 | 6.2636e-05 | 0.00010428 | 6.5153e-06 | 0.00010867 | 2.1311e-05 | 0.00010422 | 7.3446e-06 | 6.6771e-06 | [5 empty] | 22.3214 | 4.9585 | 7.7292 | 5.7228e-05 | 8.1353 | 3.4809e-05 | 1.357 | 8.0937e-06 | 6.905e-06 |  |  |  | 1 | 0 | 0 | 22.3211 | 4.9585 | 7.7291 | 5.0713e-05 | 8.1352 | 1.3498e-05 | 1.3569 | 7.4908e-07 | 2.2785e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7241 | 22.3211 | 4.9585 | 7.7291 | 8.1352 | 1.3569 | 0.22214 | 0.34627 | 0.36446 | 0.060791 | 1.5588 | 1.6407 | 0.27366 | 1.0525 | 0.17556 | 0.1668 | -1.5044 | -1.0605 | -1.0093 | -2.8003 | [10 empty] | 0 | 1.7241 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014064 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.4% | 3.7808 | -1.1817 | 10.7184 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.6% | -6.6397 | -5.3088 | -9.0591 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_210 | iRM_1 |  |  | 0.00029299 | 7.1761e-05 | 9.5629e-05 | 1.1711e-05 | 0.0001019 | 2.9429e-06 | 8.7965e-05 | 3.0524e-06 | 8.598e-06 | [5 empty] | 23.7682 | 5.28 | 8.2306 | 5.983e-05 | 8.6635 | 3.4079e-05 | 1.4452 | 9.4518e-06 | 6.6726e-06 |  |  |  | 1 | 0 | 0 | 23.7679 | 5.28 | 8.2305 | 4.8119e-05 | 8.6634 | 3.1136e-05 | 1.4451 | 6.3994e-06 | -1.9254e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7265 | 23.7679 | 5.28 | 8.2305 | 8.6634 | 1.4451 | 0.22215 | 0.34629 | 0.3645 | 0.0608 | 1.5588 | 1.6408 | 0.27369 | 1.0526 | 0.17558 | 0.1668 | -1.5044 | -1.0605 | -1.0092 | -2.8002 | [10 empty] | 0 | 1.7265 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014062 | -1.8133 | -1.8274 | [14 empty] | iRM_1 | Lab 1 | Jun 20, 2022 | 2 | 106.8% | -8.5396 | 1.4557 | 2.9679 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.23% | -2.1368 | -1.8507 | 7.1034 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_215 | iRM_12 |  |  | 0.00027862 | 7.135e-05 | 9.4707e-05 | 1.4254e-05 | 0.00011229 | 5.246e-06 | 9.021e-05 | 8.145e-06 | 7.8627e-06 | [5 empty] | 22.1812 | 4.9275 | 7.6809 | 5.9699e-05 | 8.0847 | 3.1154e-05 | 1.3486 | 1.5963e-05 | 7.9189e-06 |  |  |  | 1 | 0 | 0 | 22.1809 | 4.9274 | 7.6808 | 4.5445e-05 | 8.0846 | 2.5908e-05 | 1.3485 | 7.818e-06 | 5.6217e-08 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7254 | 22.1809 | 4.9274 | 7.6808 | 8.0846 | 1.3485 | 0.22215 | 0.34628 | 0.36448 | 0.060796 | 1.5588 | 1.6407 | 0.27367 | 1.0526 | 0.17557 | 0.1668 | -1.5044 | -1.0605 | -1.0093 | -2.8002 | [10 empty] | 0 | 1.7254 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014064 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.97% | -0.24638 | 2.1959 | -10.3957 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.75% | 6.7979 | 6.4877 | -0.41843 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_220 | iRM_2 |  |  | 0.00028426 | 6.0124e-05 | 0.00010907 | 9.457e-06 | 9.7312e-05 | 6.6414e-06 | 9.5001e-05 | 4.3366e-06 | 5.7068e-06 | [5 empty] | 23.9184 | 5.3134 | 8.2827 | 5.5269e-05 | 8.7183 | 3.2627e-05 | 1.4543 | 1.0358e-05 | 3.5245e-06 |  |  |  | 1 | 0 | 0 | 23.9181 | 5.3134 | 8.2825 | 4.5812e-05 | 8.7182 | 2.5985e-05 | 1.4542 | 6.0211e-06 | -2.1824e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7267 | 23.9181 | 5.3134 | 8.2825 | 8.7182 | 1.4542 | 0.22215 | 0.34629 | 0.3645 | 0.060801 | 1.5588 | 1.6408 | 0.2737 | 1.0526 | 0.17558 | 0.1668 | -1.5044 | -1.0605 | -1.0092 | -2.8002 | [10 empty] | 0 | 1.7267 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8274 | [14 empty] | iRM_2 | Lab 1 | Jun 20, 2022 | 2 | 108.1% | -8.1778 | -3.37 | -3.2575 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.8% | -3.1634 | -6.4499 | -10.2962 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_225 | iRM_12 |  |  | 0.00028211 | 6.8282e-05 | 9.2848e-05 | 1.4641e-05 | 0.00010018 | 2.8151e-06 | 8.72e-05 | 2.47e-06 | 3.1927e-06 | [5 empty] | 22.0558 | 4.8997 | 7.6377 | 5.8523e-05 | 8.0395 | 3.2096e-05 | 1.3411 | 1.2109e-05 | 8.0913e-06 |  |  |  | 1 | 0 | 0 | 22.0555 | 4.8996 | 7.6376 | 4.3882e-05 | 8.0394 | 2.9281e-05 | 1.341 | 9.6386e-06 | 4.8986e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7268 | 22.0555 | 4.8996 | 7.6376 | 8.0394 | 1.341 | 0.22215 | 0.34629 | 0.36451 | 0.060802 | 1.5588 | 1.6408 | 0.2737 | 1.0526 | 0.17558 | 0.16681 | -1.5044 | -1.0605 | -1.0092 | -2.8001 | [10 empty] | 0 | 1.7268 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014068 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.64% | 3.3589 | -1.8646 | 9.4448 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.1% | 4.8643 | 11.158 | 18.878 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_230 | iRM_10 |  |  | 0.00028293 | 7.3008e-05 | 0.00010469 | 1.465e-05 | 9.4445e-05 | -3.562e-07 | 9.1301e-05 | 4.7649e-07 | 3.9871e-06 | [5 empty] | 23.4903 | 5.2183 | 8.1345 | 5.9706e-05 | 8.5625 | 3.2805e-05 | 1.4284 | 3.5828e-06 | 2.7707e-06 |  |  |  | 1 | 0 | 0 | 23.49 | 5.2183 | 8.1344 | 4.5056e-05 | 8.5624 | 3.3161e-05 | 1.4283 | 3.1063e-06 | -1.2163e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7272 | 23.49 | 5.2183 | 8.1344 | 8.5624 | 1.4283 | 0.22215 | 0.34629 | 0.36451 | 0.060803 | 1.5588 | 1.6408 | 0.27371 | 1.0526 | 0.17558 | 0.16681 | -1.5044 | -1.0605 | -1.0092 | -2.8001 | [10 empty] | 0 | 1.7272 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8273 | [14 empty] | iRM_10 | Lab 1 | Jun 20, 2022 | 2 | 106.4% | -8.6197 | 1.4718 | -0.35788 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.06% | -2.8093 | -15.4027 | -16.1787 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_235 | iRM_12 |  |  | 0.00028928 | 7.5514e-05 | 0.00010262 | 8.9912e-07 | 0.00011088 | 1.7576e-06 | 9.9174e-05 | 2.2469e-07 | 3.9295e-06 | [5 empty] | 22.0908 | 4.9075 | 7.6498 | 5.405e-05 | 8.0523 | 3.2396e-05 | 1.3433 | 1.1254e-05 | 7.917e-06 |  |  |  | 1 | 0 | 0 | 22.0905 | 4.9074 | 7.6497 | 5.3151e-05 | 8.0522 | 3.0639e-05 | 1.3432 | 1.1029e-05 | 3.9875e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.727 | 22.0905 | 4.9074 | 7.6497 | 8.0522 | 1.3432 | 0.22215 | 0.34629 | 0.36451 | 0.060802 | 1.5588 | 1.6408 | 0.2737 | 1.0526 | 0.17558 | 0.16681 | -1.5044 | -1.0605 | -1.0092 | -2.8001 | [10 empty] | 0 | 1.727 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.2% | 0.34278 | -3.204 | -0.81675 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.5% | -5.3999 | 8.0389 | -1.0092 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_240 | iRM_11 |  |  | 0.00028808 | 7.5252e-05 | 0.00010396 | 1.163e-05 | 9.9407e-05 | 9.3709e-06 | 9.0494e-05 | 2.7919e-06 | 5.9523e-06 | [5 empty] | 23.6024 | 5.2434 | 8.1736 | 5.2522e-05 | 8.604 | 3.3451e-05 | 1.4353 | 5.4206e-06 | 4.0285e-06 |  |  |  | 1 | 0 | 0 | 23.6021 | 5.2433 | 8.1735 | 4.0891e-05 | 8.6039 | 2.408e-05 | 1.4352 | 2.6287e-06 | -1.9238e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7289 | 23.6021 | 5.2433 | 8.1735 | 8.6039 | 1.4352 | 0.22215 | 0.3463 | 0.36454 | 0.06081 | 1.5589 | 1.6409 | 0.27373 | 1.0527 | 0.1756 | 0.16681 | -1.5044 | -1.0604 | -1.0091 | -2.8 | [10 empty] | 0 | 1.7289 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8273 | [14 empty] | iRM_11 | Lab 1 | Jun 20, 2022 | 2 | 107% | 1.8825 | -3.8238 | 6.3693 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.4% | 3.7808 | -1.1817 | 10.7184 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_245 | iRM_12 |  |  | 0.00029178 | 6.3704e-05 | 9.7526e-05 | 4.3993e-06 | 0.00010158 | 9.1963e-06 | 9.6227e-05 | 6.9302e-06 | 4.8822e-07 | [5 empty] | 22.0376 | 4.8957 | 7.6317 | 5.1403e-05 | 8.0333 | 3.2658e-05 | 1.3401 | 1.3115e-05 | 6.7679e-06 |  |  |  | 1 | 0 | 0 | 22.0373 | 4.8956 | 7.6316 | 4.7004e-05 | 8.0332 | 2.3462e-05 | 1.34 | 6.1844e-06 | 6.2796e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7282 | 22.0373 | 4.8956 | 7.6316 | 8.0332 | 1.34 | 0.22215 | 0.3463 | 0.36453 | 0.060807 | 1.5589 | 1.6409 | 0.27372 | 1.0526 | 0.17559 | 0.16681 | -1.5044 | -1.0604 | -1.0092 | -2.8001 | [10 empty] | 0 | 1.7282 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014062 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100% | -5.6523 | 7.72 | -6.694 | [8 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.97% | -0.24638 | 2.1959 | -10.3957 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_250 | iRM |  |  | 0.0002847 | 6.7395e-05 | 0.00010052 | 3.6099e-06 | 8.7096e-05 | 6.8854e-06 | 8.8342e-05 | 4.7442e-06 | 3.8327e-06 | [5 empty] | 23.9207 | 5.3141 | 8.2839 | 4.9845e-05 | 8.7201 | 3.2601e-05 | 1.4547 | 1.067e-05 | 4.8085e-06 |  |  |  | 1 | 0 | 0 | 23.9204 | 5.314 | 8.2838 | 4.6236e-05 | 8.72 | 2.5715e-05 | 1.4546 | 5.926e-06 | 9.7582e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7291 | 23.9204 | 5.314 | 8.2838 | 8.72 | 1.4546 | 0.22215 | 0.34631 | 0.36454 | 0.060811 | 1.5589 | 1.6409 | 0.27373 | 1.0527 | 0.1756 | 0.16681 | -1.5044 | -1.0604 | -1.0091 | -2.8 | [10 empty] | 0 | 1.7291 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014066 | -1.8133 | -1.8273 | [14 empty] | iRM | Lab 1 | Jun 20, 2022 | 2 | 108.7% | 7.2755 | 0.88182 | 6.9992 | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 99.64% | 3.3589 | -1.8646 | 9.4448 | iRM_12
Lab 1 | Jun 20, 2022 | 2 | Zr_255 | iRM_12 |  |  | 0.00033158 | 7.9226e-05 | 0.00011301 | 6.7525e-06 | 0.00012395 | 7.3348e-06 | 9.1926e-05 | 1.6039e-06 | 5.5097e-06 | [5 empty] | 21.9863 | 4.8844 | 7.6141 | 5.1844e-05 | 8.0152 | 3.0216e-05 | 1.3371 | 1.0535e-05 | 9.0029e-06 |  |  |  | 1 | 0 | 0 | 21.986 | 4.8843 | 7.614 | 4.5091e-05 | 8.015 | 2.2882e-05 | 1.337 | 8.9313e-06 | 3.4932e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7298 | 21.986 | 4.8843 | 7.614 | 8.015 | 1.337 | 0.22216 | 0.34631 | 0.36455 | 0.060814 | 1.5589 | 1.641 | 0.27374 | 1.0527 | 0.1756 | 0.16682 | -1.5044 | -1.0604 | -1.0091 | -2.7999 | [10 empty] | 0 | 1.7298 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014072 | -1.8133 | -1.8273 | [22 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100.2% | 0.34278 | -3.204 | -0.81675 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_005 | iRM_12 |  |  | 0.0002015 | 4.4719e-05 | 6.9993e-05 | 1.326e-05 | 6.5544e-05 | 1.8025e-06 | 0.00011417 | -1.3958e-06 | -4.9971e-06 | [5 empty] | 20.8665 | 4.6356 | 7.2261 | 5.4659e-05 | 7.6064 | 2.4791e-05 | 1.2689 | 1.2068e-05 | 1.0138e-06 |  |  |  | 1 | 0 | 0 | 20.8663 | 4.6355 | 7.226 | 4.1399e-05 | 7.6063 | 2.2988e-05 | 1.2688 | 1.3464e-05 | 6.0108e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7281 | 20.8663 | 4.6355 | 7.226 | 7.6063 | 1.2688 | 0.22215 | 0.3463 | 0.36453 | 0.060805 | 1.5588 | 1.6409 | 0.27371 | 1.0526 | 0.17559 | 0.16681 | -1.5044 | -1.0604 | -1.0092 | -2.8001 | [9 empty] | L1 JunS3 | 0 | 1.7281 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014064 | -1.8133 | -1.8274 | [22 empty] | iRM_12 | Lab 1 | Jun 20, 2022 | 2 | 100% | -5.6523 | 7.72 | -6.694 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_010 | iRM_UR |  |  | 0.00024438 | 5.0418e-05 | 8.7866e-05 | 2.4174e-05 | 7.6885e-05 | -1.5017e-06 | 0.00010545 | -4.1821e-07 | -4.423e-06 | [5 empty] | 22.0676 | 4.9025 | 7.6423 | 4.3331e-05 | 8.0448 | 2.3683e-05 | 1.3421 | 6.5085e-06 | -2.7251e-06 |  |  |  | 1 | 0 | 0 | 22.0673 | 4.9025 | 7.6422 | 1.9157e-05 | 8.0447 | 2.5185e-05 | 1.342 | 6.9267e-06 | 1.6979e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7298 | 22.0673 | 4.9025 | 7.6422 | 8.0447 | 1.342 | 0.22216 | 0.34631 | 0.36455 | 0.060814 | 1.5589 | 1.641 | 0.27374 | 1.0527 | 0.1756 | 0.16682 | -1.5044 | -1.0604 | -1.0091 | -2.7999 | [10 empty] | 0 | 1.7298 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3882 | 0.014063 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 21, 2022 | 3 | 107.2% | 14.5608 | 3.0016 | 13.4888 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.11% | -11.8179 | -1.5137 | 16.6074 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_015 | iRM_12 |  |  | 0.00023792 | 6.0207e-05 | 7.8018e-05 | 6.912e-06 | 7.3009e-05 | -2.3257e-06 | 0.00010127 | 1.057e-05 | 1.6645e-06 | [5 empty] | 20.3007 | 4.5101 | 7.0309 | 5.4254e-05 | 7.4018 | 2.5165e-05 | 1.2349 | 1.7734e-05 | -2.3799e-07 |  |  |  | 1 | 0 | 0 | 20.3005 | 4.51 | 7.0309 | 4.7342e-05 | 7.4018 | 2.7491e-05 | 1.2348 | 7.1642e-06 | -1.9025e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7334 | 20.3005 | 4.51 | 7.0309 | 7.4018 | 1.2348 | 0.22216 | 0.34634 | 0.36461 | 0.060828 | 1.5589 | 1.6412 | 0.2738 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.7334 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.11% | -11.8179 | -1.5137 | 16.6074 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.79% | 6.5884 | -2.6304 | 0.38213 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_020 | SRM |  |  | 0.00022809 | 4.0951e-05 | 8.8511e-05 | 9.9166e-06 | 6.2641e-05 | 2.5823e-07 | 9.5747e-05 | 9.4366e-06 | -3.9687e-06 | [5 empty] | 20.5856 | 4.5735 | 7.1297 | 5.1603e-05 | 7.5059 | 2.5313e-05 | 1.2523 | 1.1377e-05 | 1.8624e-06 |  |  |  | 1 | 0 | 0 | 20.5853 | 4.5734 | 7.1296 | 4.1687e-05 | 7.5058 | 2.5055e-05 | 1.2522 | 1.9406e-06 | 5.8311e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.734 | 20.5853 | 4.5734 | 7.1296 | 7.5058 | 1.2522 | 0.22217 | 0.34634 | 0.36462 | 0.060829 | 1.5589 | 1.6412 | 0.2738 | 1.0528 | 0.17563 | 0.16683 | -1.5043 | -1.0603 | -1.0089 | -2.7997 | [10 empty] | 0 | 1.734 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5296 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.4391 | -1.3883 | 0.014068 | -1.8133 | -1.8274 | [14 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 101.9% | 13.4969 | 0.93421 | -14.066 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.84% | -1.7356 | 2.4936 | 5.997 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_025 | iRM_12 |  |  | 0.00078 | 0.00016702 | 0.00027612 | 1.619e-05 | 0.00026269 | 1.4862e-06 | 0.00012311 | 1.4643e-05 | -3.3333e-06 | [5 empty] | 20.1015 | 4.466 | 6.9624 | 5.4287e-05 | 7.3301 | 2.3253e-05 | 1.223 | 1.1577e-05 | -7.2746e-08 |  |  |  | 1 | 0 | 0 | 20.1008 | 4.4658 | 6.9621 | 3.8097e-05 | 7.3299 | 2.1766e-05 | 1.2229 | -3.0666e-06 | 3.2605e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7364 | 20.1008 | 4.4658 | 6.9621 | 7.3299 | 1.2229 | 0.22217 | 0.34636 | 0.36466 | 0.060839 | 1.559 | 1.6413 | 0.27384 | 1.0528 | 0.17565 | 0.16684 | -1.5043 | -1.0603 | -1.0088 | -2.7995 | [10 empty] | 0 | 1.7364 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014069 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.79% | 6.5884 | -2.6304 | 0.38213 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.2% | 1.3515 | 1.3094 | -18.2097 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_030 | alfaA |  |  | 0.00032017 | 5.8673e-05 | 0.00011438 | 8.3257e-06 | 0.00010605 | 4.8157e-07 | 9.4788e-05 | 3.1479e-06 | -7.6546e-07 | [5 empty] | 20.5384 | 4.5625 | 7.1119 | 4.1234e-05 | 7.4856 | 2.4156e-05 | 1.2487 | 1.0572e-05 | -2.5311e-08 |  |  |  | 1 | 0 | 0 | 20.5381 | 4.5624 | 7.1117 | 3.2908e-05 | 7.4855 | 2.3675e-05 | 1.2486 | 7.4238e-06 | 7.4015e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7245 | 20.5381 | 4.5624 | 7.1117 | 7.4855 | 1.2486 | 0.22214 | 0.34627 | 0.36447 | 0.060793 | 1.5588 | 1.6407 | 0.27367 | 1.0526 | 0.17557 | 0.1668 | -1.5044 | -1.0605 | -1.0093 | -2.8003 | [10 empty] | 0 | 1.7245 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43911 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 21, 2022 | 3 | 102.5% | 0.48875 | -3.5645 | 19.444 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.48% | -1.3921 | -4.1706 | 17.2662 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_035 | iRM_12 |  |  | 0.00025432 | 5.8472e-05 | 8.6458e-05 | 6.3867e-06 | 7.8944e-05 | -7.6854e-07 | 0.00010557 | 2.6255e-06 | -3.2433e-06 | [5 empty] | 19.9856 | 4.4403 | 6.9226 | 5.571e-05 | 7.2885 | 2.5232e-05 | 1.2162 | 1.2195e-05 | 6.4112e-07 |  |  |  | 1 | 0 | 0 | 19.9853 | 4.4403 | 6.9225 | 4.9323e-05 | 7.2884 | 2.6e-05 | 1.216 | 9.5693e-06 | 3.8844e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7384 | 19.9853 | 4.4403 | 6.9225 | 7.2884 | 1.216 | 0.22218 | 0.34638 | 0.36469 | 0.060847 | 1.559 | 1.6414 | 0.27387 | 1.0529 | 0.17567 | 0.16685 | -1.5043 | -1.0602 | -1.0087 | -2.7994 | [10 empty] | 0 | 1.7384 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014065 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.84% | -1.7356 | 2.4936 | 5.997 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.3% | -1.4995 | 1.7153 | -19.2164 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_040 | Ames |  |  | 0.00023705 | 5.3088e-05 | 8.0964e-05 | 8.4248e-06 | 6.4245e-05 | 3.9452e-07 | 8.3217e-05 | 7.2927e-06 | -1.4782e-06 | [5 empty] | 20.9183 | 4.6479 | 7.2469 | 3.772e-05 | 7.6311 | 2.9707e-05 | 1.2735 | 1.1976e-05 | 9.3356e-06 |  |  |  | 1 | 0 | 0 | 20.9181 | 4.6479 | 7.2468 | 2.9296e-05 | 7.631 | 2.9313e-05 | 1.2734 | 4.6836e-06 | 1.0814e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7458 | 20.9181 | 4.6479 | 7.2468 | 7.631 | 1.2734 | 0.22219 | 0.34644 | 0.36481 | 0.060875 | 1.5592 | 1.6418 | 0.27397 | 1.053 | 0.17572 | 0.16687 | -1.5042 | -1.0601 | -1.0084 | -2.7989 | [10 empty] | 0 | 1.7457 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014062 | -1.8133 | -1.8274 | [14 empty] | Ames | Lab 1 | Jun 21, 2022 | 3 | 104.8% | -1.3223 | 3.1852 | 2.9575 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.99% | 1.4999 | 3.5417 | 23.8293 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_045 | iRM_12 |  |  | 0.0019442 | 0.00042442 | 0.00067178 | 1.4119e-05 | 0.00070341 | 1.7401e-06 | 0.00019431 | 3.7959e-06 | -8.2313e-06 | [5 empty] | 19.9344 | 4.429 | 6.9051 | 5.5404e-05 | 7.2704 | 2.5158e-05 | 1.2131 | 1.3737e-05 | 1.3502e-06 |  |  |  | 1 | 0 | 0 | 19.9324 | 4.4286 | 6.9044 | 4.1285e-05 | 7.2697 | 2.3417e-05 | 1.213 | 9.9415e-06 | 9.5815e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7402 | 19.9324 | 4.4286 | 6.9044 | 7.2697 | 1.2129 | 0.22218 | 0.34639 | 0.36472 | 0.060853 | 1.5591 | 1.6415 | 0.27389 | 1.0529 | 0.17568 | 0.16685 | -1.5043 | -1.0602 | -1.0086 | -2.7993 | [10 empty] | 0 | 1.7402 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014066 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.2% | 1.3515 | 1.3094 | -18.2097 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.89% | 5.039 | -3.6327 | -19.5207 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_050 | IPGP |  |  | 0.00023917 | 4.5168e-05 | 8.4508e-05 | 8.5299e-06 | 6.9064e-05 | -7.2585e-07 | 0.0001002 | 8.942e-06 | -2.4627e-06 | [5 empty] | 20.7566 | 4.6116 | 7.1897 | 3.6217e-05 | 7.5698 | 2.552e-05 | 1.2631 | 1.5913e-05 | 1.8338e-06 |  |  |  | 1 | 0 | 0 | 20.7564 | 4.6115 | 7.1896 | 2.7687e-05 | 7.5697 | 2.6246e-05 | 1.263 | 6.971e-06 | 4.2965e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7387 | 20.7564 | 4.6115 | 7.1896 | 7.5697 | 1.263 | 0.22217 | 0.34638 | 0.36469 | 0.060847 | 1.559 | 1.6415 | 0.27387 | 1.0529 | 0.17567 | 0.16684 | -1.5043 | -1.0602 | -1.0087 | -2.7994 | [10 empty] | 0 | 1.7387 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5514 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [14 empty] | IPGP | Lab 1 | Jun 21, 2022 | 3 | 104.4% | -10.6311 | -0.19379 | -2.644 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.7% | -14.2281 | -3.0436 | 1.9208 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_055 | iRM_12 |  |  | 0.00028394 | 5.6906e-05 | 9.8978e-05 | 9.2232e-06 | 7.7062e-05 | 3.9901e-06 | 9.2094e-05 | 7.8447e-06 | 1.7328e-06 | [5 empty] | 19.8171 | 4.403 | 6.8647 | 5.0093e-05 | 7.228 | 2.2987e-05 | 1.2061 | 1.8677e-05 | 1.9903e-06 |  |  |  | 1 | 0 | 0 | 19.8168 | 4.403 | 6.8646 | 4.0869e-05 | 7.228 | 1.8997e-05 | 1.206 | 1.0833e-05 | 2.5747e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7415 | 19.8168 | 4.403 | 6.8646 | 7.228 | 1.206 | 0.22218 | 0.3464 | 0.36474 | 0.060859 | 1.5591 | 1.6416 | 0.27391 | 1.0529 | 0.17569 | 0.16686 | -1.5042 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7415 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.48% | -1.3921 | -4.1706 | 17.2662 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | 16.9696 | 7.8761 | -3.6976 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_060 | iRM_UR |  |  | 0.00025365 | 4.1499e-05 | 8.9707e-05 | 9.9106e-06 | 8.151e-05 | 6.1196e-07 | 9.1921e-05 | 3.0451e-06 | -2.726e-06 | [5 empty] | 21.1415 | 4.6973 | 7.3235 | 3.8889e-05 | 7.711 | 2.2932e-05 | 1.2867 | 1.0321e-05 | -3.6322e-06 |  |  |  | 1 | 0 | 0 | 21.1413 | 4.6973 | 7.3234 | 2.8978e-05 | 7.7109 | 2.232e-05 | 1.2866 | 7.2759e-06 | -9.0621e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7412 | 21.1413 | 4.6973 | 7.3234 | 7.7109 | 1.2866 | 0.22218 | 0.3464 | 0.36473 | 0.060858 | 1.5591 | 1.6416 | 0.27391 | 1.0529 | 0.17569 | 0.16686 | -1.5042 | -1.0602 | -1.0086 | -2.7992 | [10 empty] | 0 | 1.7412 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014057 | -1.8133 | -1.8274 | [14 empty] | iRM_UR | Lab 1 | Jun 21, 2022 | 3 | 106.4% | 7.1602 | 9.5924 | 4.6643 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.81% | -8.587 | -5.6283 | 16.5909 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_065 | iRM_12 |  |  | 0.00026701 | 5.7195e-05 | 9.3183e-05 | 1.9263e-06 | 8.2258e-05 | -1.4254e-06 | 9.9749e-05 | 8.0607e-06 | -2.2232e-06 | [5 empty] | 19.9094 | 4.4236 | 6.8968 | 5.0533e-05 | 7.2621 | 2.4598e-05 | 1.2118 | 1.1991e-05 | -1.0788e-06 |  |  |  | 1 | 0 | 0 | 19.9091 | 4.4236 | 6.8968 | 4.8606e-05 | 7.262 | 2.6023e-05 | 1.2117 | 3.9307e-06 | 1.1445e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7428 | 19.9091 | 4.4236 | 6.8968 | 7.262 | 1.2117 | 0.22219 | 0.34641 | 0.36476 | 0.060863 | 1.5591 | 1.6417 | 0.27393 | 1.053 | 0.1757 | 0.16686 | -1.5042 | -1.0601 | -1.0085 | -2.7991 | [10 empty] | 0 | 1.7428 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014064 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.3% | -1.4995 | 1.7153 | -19.2164 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | -0.66443 | 2.219 | -0.45114 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_070 | SRM |  |  | 0.00024773 | 5.5748e-05 | 8.2301e-05 | 1.6127e-05 | 7.7792e-05 | -9.9929e-08 | 9.9038e-05 | 1.5375e-05 | -2.1899e-06 | [5 empty] | 20.5306 | 4.5617 | 7.1121 | 5.2881e-05 | 7.4889 | 2.6565e-05 | 1.2497 | 1.2837e-05 | -1.8999e-07 |  |  |  | 1 | 0 | 0 | 20.5303 | 4.5616 | 7.112 | 3.6754e-05 | 7.4889 | 2.6665e-05 | 1.2496 | -2.5383e-06 | 1.9999e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7436 | 20.5303 | 4.5616 | 7.112 | 7.4889 | 1.2496 | 0.22219 | 0.34642 | 0.36477 | 0.060866 | 1.5591 | 1.6417 | 0.27394 | 1.053 | 0.1757 | 0.16686 | -1.5042 | -1.0601 | -1.0085 | -2.7991 | [10 empty] | 0 | 1.7436 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.01407 | -1.8133 | -1.8274 | [14 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 103.2% | -1.625 | -6.0986 | -18.6138 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.98% | 5.8598 | 3.9872 | -6.4581 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_075 | iRM_12 |  |  | 0.00026762 | 5.1245e-05 | 9.182e-05 | 1.5379e-05 | 8.8995e-05 | 1.8362e-05 | 0.0001015 | 1.3887e-05 | -6.4923e-07 | [5 empty] | 19.8816 | 4.4175 | 6.8874 | 4.823e-05 | 7.2523 | 3.753e-05 | 1.2103 | 1.3228e-05 | 1.2997e-06 |  |  |  | 1 | 0 | 0 | 19.8813 | 4.4174 | 6.8873 | 3.2851e-05 | 7.2522 | 1.9168e-05 | 1.2102 | -6.597e-07 | 1.949e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7439 | 19.8813 | 4.4174 | 6.8873 | 7.2522 | 1.2102 | 0.22219 | 0.34642 | 0.36478 | 0.060869 | 1.5591 | 1.6417 | 0.27395 | 1.053 | 0.17571 | 0.16687 | -1.5042 | -1.0601 | -1.0085 | -2.799 | [10 empty] | 0 | 1.7439 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014063 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.99% | 1.4999 | 3.5417 | 23.8293 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.63% | -9.4498 | -9.4533 | -9.3109 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_080 | alfaA |  |  | 0.0002687 | 5.3433e-05 | 9.6012e-05 | 1.7775e-05 | 7.8645e-05 | 3.793e-06 | 9.8099e-05 | 2.4698e-06 | -2.7045e-06 | [5 empty] | 20.3547 | 4.522 | 7.0495 | 3.8612e-05 | 7.4211 | 2.3888e-05 | 1.2381 | 1.2135e-05 | -2.5247e-07 |  |  |  | 1 | 0 | 0 | 20.3544 | 4.522 | 7.0494 | 2.0837e-05 | 7.421 | 2.0095e-05 | 1.238 | 9.6653e-06 | 2.452e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7322 | 20.3544 | 4.522 | 7.0494 | 7.421 | 1.238 | 0.22216 | 0.34633 | 0.36459 | 0.060824 | 1.5589 | 1.6411 | 0.27378 | 1.0527 | 0.17562 | 0.16683 | -1.5044 | -1.0604 | -1.009 | -2.7998 | [10 empty] | 0 | 1.7322 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014065 | -1.8133 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 21, 2022 | 3 | 102.4% | -5.192 | 0.28368 | 27.6043 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.2% | 12.1495 | 6.7028 | 16.4242 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_085 | iRM_12 |  |  | 0.00026709 | 4.8683e-05 | 9.4297e-05 | 9.1151e-06 | 7.8921e-05 | 2.9316e-06 | 0.00010189 | 9.5657e-06 | -1.5936e-06 | [5 empty] | 19.8595 | 4.4127 | 6.8799 | 4.9385e-05 | 7.2447 | 2.4253e-05 | 1.209 | 1.301e-05 | 2.284e-06 |  |  |  | 1 | 0 | 0 | 19.8592 | 4.4126 | 6.8798 | 4.027e-05 | 7.2446 | 2.1321e-05 | 1.2089 | 3.4439e-06 | 3.8776e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7454 | 19.8592 | 4.4126 | 6.8798 | 7.2446 | 1.2089 | 0.22219 | 0.34643 | 0.3648 | 0.060873 | 1.5591 | 1.6418 | 0.27396 | 1.053 | 0.17572 | 0.16687 | -1.5042 | -1.0601 | -1.0084 | -2.799 | [10 empty] | 0 | 1.7454 | 0.21794 | 0.33338 | 0.3381 | 0.054378 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3883 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.89% | 5.039 | -3.6327 | -19.5207 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.5% | -9.7873 | -1.9428 | -7.5515 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_090 | Ames |  |  | 0.00026476 | 5.2323e-05 | 9.0906e-05 | 1.3894e-05 | 7.0324e-05 | -2.2498e-08 | 0.00010779 | 7.952e-06 | -6.9249e-06 | [5 empty] | 20.9632 | 4.6582 | 7.2634 | 4.559e-05 | 7.6496 | 3.2348e-05 | 1.2768 | 1.5703e-05 | 8.9507e-06 |  |  |  | 1 | 0 | 0 | 20.9629 | 4.6582 | 7.2633 | 3.1696e-05 | 7.6495 | 3.2371e-05 | 1.2767 | 7.7513e-06 | 1.5876e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7522 | 20.9629 | 4.6582 | 7.2633 | 7.6495 | 1.2766 | 0.22221 | 0.34648 | 0.36491 | 0.0609 | 1.5593 | 1.6422 | 0.27407 | 1.0532 | 0.17577 | 0.16689 | -1.5041 | -1.0599 | -1.0081 | -2.7985 | [10 empty] | 0 | 1.7522 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014063 | -1.8133 | -1.8274 | [14 empty] | Ames | Lab 1 | Jun 21, 2022 | 3 | 105.5% | 0.49654 | 5.0129 | -3.0264 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.56% | 2.1434 | -0.16094 | -3.3429 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_095 | iRM_12 |  |  | 0.00025829 | 4.8725e-05 | 8.7486e-05 | 1.3149e-05 | 8.0635e-05 | 3.4186e-06 | 9.2175e-05 | 2.7874e-06 | 6.6491e-07 | [5 empty] | 19.8806 | 4.4173 | 6.8873 | 5.3251e-05 | 7.2526 | 2.3164e-05 | 1.2103 | 1.3294e-05 | 2.7449e-07 |  |  |  | 1 | 0 | 0 | 19.8803 | 4.4173 | 6.8873 | 4.0101e-05 | 7.2525 | 1.9745e-05 | 1.2102 | 1.0507e-05 | -3.9042e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.746 | 19.8803 | 4.4173 | 6.8873 | 7.2525 | 1.2102 | 0.22219 | 0.34644 | 0.36481 | 0.060877 | 1.5592 | 1.6419 | 0.27398 | 1.053 | 0.17572 | 0.16687 | -1.5042 | -1.0601 | -1.0084 | -2.7989 | [10 empty] | 0 | 1.746 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.7% | -14.2281 | -3.0436 | 1.9208 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.83% | 4.9071 | 3.5781 | 1.1253 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_100 | IPGP |  |  | 0.00024552 | 5.2671e-05 | 8.0852e-05 | 6.5968e-06 | 7.6463e-05 | 2.2919e-06 | 8.9488e-05 | 1.0508e-05 | -9.2147e-06 | [5 empty] | 21.0683 | 4.6812 | 7.2986 | 4.1065e-05 | 7.6855 | 2.4944e-05 | 1.2825 | 1.6552e-05 | 5.2211e-07 |  |  |  | 1 | 0 | 0 | 21.068 | 4.6812 | 7.2986 | 3.4468e-05 | 7.6854 | 2.2652e-05 | 1.2824 | 6.0448e-06 | 9.7368e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7448 | 21.068 | 4.6812 | 7.2986 | 7.6854 | 1.2824 | 0.22219 | 0.34643 | 0.36479 | 0.060871 | 1.5591 | 1.6418 | 0.27396 | 1.053 | 0.17571 | 0.16687 | -1.5042 | -1.0601 | -1.0084 | -2.799 | [10 empty] | 0 | 1.7448 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3883 | 0.014063 | -1.8133 | -1.8274 | [14 empty] | IPGP | Lab 1 | Jun 21, 2022 | 3 | 105.6% | 4.0962 | 0.13366 | -12.4773 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | -10.1474 | -8.323 | 4.5738 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_105 | iRM_12 |  |  | 0.0002677 | 4.8329e-05 | 9.81e-05 | 1.3014e-05 | 6.3673e-05 | -3.5482e-06 | 9.9931e-05 | 1.1254e-05 | -4.0168e-06 | [5 empty] | 20.0227 | 4.4491 | 6.9369 | 5.5265e-05 | 7.305 | 2.5106e-05 | 1.2191 | 1.8468e-05 | -3.6687e-07 |  |  |  | 1 | 0 | 0 | 20.0224 | 4.449 | 6.9368 | 4.225e-05 | 7.3049 | 2.8654e-05 | 1.219 | 7.2134e-06 | 3.6499e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7478 | 20.0224 | 4.449 | 6.9368 | 7.3049 | 1.219 | 0.2222 | 0.34645 | 0.36484 | 0.060884 | 1.5592 | 1.6419 | 0.274 | 1.0531 | 0.17574 | 0.16688 | -1.5042 | -1.06 | -1.0083 | -2.7988 | [10 empty] | 0 | 1.7478 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.4391 | -1.3882 | 0.014059 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | 16.9696 | 7.8761 | -3.6976 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.43% | 13.1667 | 5.4013 | -11.6894 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_110 | iRM_UR |  |  | 0.00025135 | 4.4394e-05 | 9.2223e-05 | 9.4693e-06 | 8.9456e-05 | 1.8874e-06 | 9.3511e-05 | 2.9048e-06 | -2.1694e-06 | [5 empty] | 21.3318 | 4.7399 | 7.3902 | 3.8635e-05 | 7.7822 | 2.4145e-05 | 1.2988 | 8.314e-06 | -5.7552e-06 |  |  |  | 1 | 0 | 0 | 21.3315 | 4.7398 | 7.3901 | 2.9165e-05 | 7.7821 | 2.2258e-05 | 1.2987 | 5.4092e-06 | -3.5859e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7466 | 21.3315 | 4.7398 | 7.3901 | 7.7822 | 1.2987 | 0.2222 | 0.34644 | 0.36482 | 0.060881 | 1.5592 | 1.6419 | 0.27399 | 1.053 | 0.17573 | 0.16688 | -1.5042 | -1.06 | -1.0084 | -2.7988 | [10 empty] | 0 | 1.7466 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43911 | -1.3882 | 0.01406 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 21, 2022 | 3 | 106.6% | -2.2618 | 2.5281 | 7.0059 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100% | -11.0139 | 0.42583 | 25.8053 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_115 | iRM_12 |  |  | 0.0002814 | 6.1234e-05 | 0.00010587 | 7.6295e-06 | 8.9672e-05 | 1.3064e-06 | 8.4099e-05 | 8.6961e-06 | -4.9954e-06 | [5 empty] | 20.0106 | 4.4464 | 6.9329 | 5.6398e-05 | 7.301 | 2.468e-05 | 1.2185 | 1.1975e-05 | -6.6632e-07 |  |  |  | 1 | 0 | 0 | 20.0103 | 4.4463 | 6.9327 | 4.8769e-05 | 7.3009 | 2.3373e-05 | 1.2184 | 3.2785e-06 | 4.329e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.749 | 20.0103 | 4.4463 | 6.9327 | 7.3009 | 1.2184 | 0.2222 | 0.34646 | 0.36486 | 0.06089 | 1.5592 | 1.642 | 0.27403 | 1.0531 | 0.17575 | 0.16689 | -1.5042 | -1.06 | -1.0082 | -2.7987 | [10 empty] | 0 | 1.749 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.81% | -8.587 | -5.6283 | 16.5909 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | 4.1844 | 1.1252 | -24.4888 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_120 | SRM |  |  | 0.0017976 | 0.0003957 | 0.00062088 | 1.7601e-05 | 0.00063618 | 6.1715e-06 | 0.00018704 | 1.0543e-06 | -3.766e-06 | [5 empty] | 20.5506 | 4.5664 | 7.1198 | 4.8621e-05 | 7.4978 | 2.8057e-05 | 1.2513 | 1.4885e-05 | 2.5969e-08 |  |  |  | 1 | 0 | 0 | 20.5488 | 4.566 | 7.1192 | 3.102e-05 | 7.4972 | 2.1885e-05 | 1.2511 | 1.3831e-05 | 3.792e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7484 | 20.5488 | 4.566 | 7.1192 | 7.4972 | 1.2511 | 0.2222 | 0.34645 | 0.36485 | 0.060886 | 1.5592 | 1.642 | 0.27401 | 1.0531 | 0.17574 | 0.16688 | -1.5042 | -1.06 | -1.0083 | -2.7987 | [10 empty] | 0 | 1.7484 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [14 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 102.5% | -0.2557 | -2.8863 | -18.6487 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.99% | 1.3538 | -1.198 | 9.2785 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_125 | iRM_12 |  |  | 0.00031446 | 6.4289e-05 | 0.00010267 | 1.0199e-05 | 9.2252e-05 | 5.2989e-06 | 9.7989e-05 | 5.5052e-06 | -4.7563e-06 | [5 empty] | 20.0757 | 4.4609 | 6.9555 | 5.4081e-05 | 7.325 | 2.4599e-05 | 1.2225 | 1.2841e-05 | -1.9677e-06 |  |  |  | 1 | 0 | 0 | 20.0754 | 4.4608 | 6.9554 | 4.3882e-05 | 7.3249 | 1.93e-05 | 1.2224 | 7.3362e-06 | 2.7886e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7498 | 20.0754 | 4.4608 | 6.9554 | 7.3249 | 1.2224 | 0.2222 | 0.34647 | 0.36487 | 0.060892 | 1.5592 | 1.642 | 0.27404 | 1.0531 | 0.17575 | 0.16689 | -1.5042 | -1.06 | -1.0082 | -2.7986 | [10 empty] | 0 | 1.7498 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014064 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | -0.66443 | 2.219 | -0.45114 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.72% | -2.2327 | -0.61844 | 0.11641 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_130 | alfaA |  |  | 0.00027291 | 6.5112e-05 | 8.7866e-05 | 2.4415e-05 | 9.2833e-05 | 3.0333e-06 | 0.00010478 | -6.0992e-06 | -4.9622e-06 | [5 empty] | 20.6278 | 4.583 | 7.1449 | 4.2575e-05 | 7.5226 | 2.642e-05 | 1.2552 | 1.4784e-05 | -2.4415e-06 |  |  |  | 1 | 0 | 0 | 20.6276 | 4.5829 | 7.1448 | 1.8161e-05 | 7.5225 | 2.3387e-05 | 1.2551 | 2.0883e-05 | 2.5207e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7379 | 20.6276 | 4.5829 | 7.1448 | 7.5225 | 1.2551 | 0.22217 | 0.34637 | 0.36468 | 0.060845 | 1.559 | 1.6414 | 0.27386 | 1.0529 | 0.17566 | 0.16684 | -1.5043 | -1.0602 | -1.0087 | -2.7994 | [10 empty] | 0 | 1.7379 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [14 empty] | alfaA | Lab 1 | Jun 21, 2022 | 3 | 103% | -10.9214 | -6.3785 | -0.95754 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.95% | 3.9704 | 0.8925 | -12.5314 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_135 | iRM_12 |  |  | 0.0012238 | 0.00026507 | 0.00043089 | 1.4137e-05 | 0.00043461 | 2.0905e-06 | 0.00015747 | 8.0513e-06 | -6.6152e-06 | [5 empty] | 19.9831 | 4.4404 | 6.9235 | 5.1083e-05 | 7.2913 | 2.3193e-05 | 1.2169 | 1.1982e-05 | 2.9274e-07 |  |  |  | 1 | 0 | 0 | 19.9819 | 4.4401 | 6.9231 | 3.6946e-05 | 7.2909 | 2.1102e-05 | 1.2168 | 3.9311e-06 | 6.908e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7501 | 19.9819 | 4.4401 | 6.9231 | 7.2909 | 1.2167 | 0.22221 | 0.34647 | 0.36487 | 0.060893 | 1.5592 | 1.6421 | 0.27404 | 1.0531 | 0.17575 | 0.16689 | -1.5042 | -1.06 | -1.0082 | -2.7986 | [10 empty] | 0 | 1.7501 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014065 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.98% | 5.8598 | 3.9872 | -6.4581 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.7% | -8.061 | -3.5431 | 19.8195 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_140 | Ames |  |  | 0.00026936 | 6.6014e-05 | 9.4771e-05 | 1.8075e-05 | 7.3434e-05 | 1.3271e-06 | 9.1222e-05 | -2.1138e-07 | -3.7412e-06 | [5 empty] | 21.1361 | 4.697 | 7.3243 | 4.7259e-05 | 7.7148 | 3.2363e-05 | 1.2878 | 2.0275e-05 | 7.6551e-06 |  |  |  | 1 | 0 | 0 | 21.1358 | 4.6969 | 7.3242 | 2.9184e-05 | 7.7147 | 3.1036e-05 | 1.2877 | 2.0486e-05 | 1.1396e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7585 | 21.1358 | 4.6969 | 7.3242 | 7.7147 | 1.2877 | 0.22223 | 0.34653 | 0.36501 | 0.060926 | 1.5594 | 1.6425 | 0.27416 | 1.0533 | 0.17582 | 0.16692 | -1.5041 | -1.0598 | -1.0078 | -2.7981 | [10 empty] | 0 | 1.7585 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014062 | -1.8133 | -1.8274 | [14 empty] | Ames | Lab 1 | Jun 21, 2022 | 3 | 106% | -1.9393 | 6.9335 | 10.2852 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.5% | 4.1712 | 5.9696 | -10.5606 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_145 | iRM_12 |  |  | 0.00025651 | 5.4921e-05 | 8.7583e-05 | 1.4569e-05 | 7.3317e-05 | -2.4222e-06 | 9.4907e-05 | 6.5697e-06 | -5.0024e-06 | [5 empty] | 19.8955 | 4.4209 | 6.8933 | 5.5387e-05 | 7.2599 | 2.3838e-05 | 1.2117 | 1.377e-05 | -8.261e-07 |  |  |  | 1 | 0 | 0 | 19.8952 | 4.4209 | 6.8932 | 4.0818e-05 | 7.2598 | 2.626e-05 | 1.2116 | 7.2007e-06 | 4.1763e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7518 | 19.8952 | 4.4209 | 6.8932 | 7.2598 | 1.2116 | 0.22221 | 0.34648 | 0.3649 | 0.060899 | 1.5592 | 1.6422 | 0.27406 | 1.0532 | 0.17577 | 0.16689 | -1.5041 | -1.0599 | -1.0081 | -2.7985 | [10 empty] | 0 | 1.7518 | 0.21794 | 0.33337 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014074 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.63% | -9.4498 | -9.4533 | -9.3109 | [8 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.78% | 1.6991 | -7.5514 | -1.9192 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_150 | IPGP |  |  | 0.00075368 | 0.00017053 | 0.00025824 | 2.2584e-05 | 0.00025207 | 3.0111e-06 | 0.00013209 | 1.2472e-05 | -2.0277e-06 | [5 empty] | 21.0768 | 4.6834 | 7.3025 | 3.7101e-05 | 7.6906 | 2.5299e-05 | 1.2836 | 1.8676e-05 | 4.7434e-07 |  |  |  | 1 | 0 | 0 | 21.0761 | 4.6833 | 7.3023 | 1.4517e-05 | 7.6903 | 2.2288e-05 | 1.2834 | 6.2039e-06 | 2.502e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7507 | 21.0761 | 4.6833 | 7.3023 | 7.6903 | 1.2834 | 0.22221 | 0.34647 | 0.36488 | 0.060895 | 1.5592 | 1.6421 | 0.27405 | 1.0531 | 0.17576 | 0.16689 | -1.5041 | -1.06 | -1.0082 | -2.7986 | [10 empty] | 0 | 1.7507 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014065 | -1.8133 | -1.8274 | [14 empty] | IPGP | Lab 1 | Jun 21, 2022 | 3 | 105.8% | 1.8167 | 4.4055 | -5.8709 | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | 1.0968 | 8.5033 | 3.6459 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_155 | iRM_12 |  |  | 0.00026397 | 6.1424e-05 | 9.8263e-05 | 1.3684e-05 | 7.8609e-05 | 4.0393e-06 | 9.4979e-05 | 1.967e-05 | -7.9326e-06 | [5 empty] | 19.9546 | 4.4342 | 6.9141 | 5.1481e-05 | 7.2818 | 2.4446e-05 | 1.2154 | 9.8542e-06 | 3.7708e-08 |  |  |  | 1 | 0 | 0 | 19.9544 | 4.4341 | 6.914 | 3.7797e-05 | 7.2818 | 2.0407e-05 | 1.2153 | -9.8154e-06 | 7.9703e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.753 | 19.9544 | 4.4341 | 6.914 | 7.2818 | 1.2153 | 0.22221 | 0.34649 | 0.36492 | 0.060905 | 1.5593 | 1.6422 | 0.27408 | 1.0532 | 0.17578 | 0.1669 | -1.5041 | -1.0599 | -1.0081 | -2.7984 | [10 empty] | 0 | 1.753 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014064 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.2% | 12.1495 | 6.7028 | 16.4242 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.03% | 1.7231 | -6.403 | 4.264 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_160 | iRM_UR |  |  | 0.00025255 | 5.3629e-05 | 8.1712e-05 | 1.142e-05 | 7.0171e-05 | 1.2251e-06 | 9.8272e-05 | 1.3075e-05 | -4.214e-06 | [5 empty] | 21.1819 | 4.7069 | 7.3391 | 4.1715e-05 | 7.7294 | 2.4031e-05 | 1.2901 | 1.9619e-05 | -3.3779e-06 |  |  |  | 1 | 0 | 0 | 21.1817 | 4.7068 | 7.3391 | 3.0294e-05 | 7.7293 | 2.2806e-05 | 1.29 | 6.5444e-06 | 8.3606e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.752 | 21.1817 | 4.7068 | 7.339 | 7.7293 | 1.29 | 0.22221 | 0.34648 | 0.3649 | 0.060901 | 1.5592 | 1.6422 | 0.27407 | 1.0532 | 0.17577 | 0.16689 | -1.5041 | -1.0599 | -1.0081 | -2.7985 | [10 empty] | 0 | 1.752 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43911 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [14 empty] | iRM_UR | Lab 1 | Jun 21, 2022 | 3 | 106.2% | 6.8249 | -2.3206 | 1.2686 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.1% | -12.6831 | -4.9748 | 20.7813 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_165 | iRM_12 |  |  | 0.00027378 | 7.0308e-05 | 0.00010098 | 1.8537e-05 | 8.3832e-05 | 1.8416e-06 | 0.00010216 | 4.0895e-06 | -5.9254e-06 | [5 empty] | 19.9455 | 4.4322 | 6.9111 | 5.6919e-05 | 7.2789 | 2.5533e-05 | 1.2149 | 1.1972e-05 | -2.319e-08 |  |  |  | 1 | 0 | 0 | 19.9452 | 4.4321 | 6.911 | 3.8382e-05 | 7.2788 | 2.3691e-05 | 1.2148 | 7.8824e-06 | 5.9022e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7542 | 19.9452 | 4.4321 | 6.911 | 7.2788 | 1.2148 | 0.22221 | 0.3465 | 0.36494 | 0.060909 | 1.5593 | 1.6423 | 0.2741 | 1.0532 | 0.17578 | 0.1669 | -1.5041 | -1.0599 | -1.008 | -2.7984 | [10 empty] | 0 | 1.7542 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.5% | -9.7873 | -1.9428 | -7.5515 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.96% | 9.4255 | 9.1951 | -15.1327 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_170 | SRM |  |  | 0.00045837 | 0.0001063 | 0.00015603 | 2.0786e-05 | 0.00014581 | -4.6295e-06 | 0.00011922 | 1.1908e-05 | -4.5837e-06 | [5 empty] | 20.2851 | 4.5076 | 7.0286 | 5.234e-05 | 7.4026 | 2.4723e-05 | 1.2356 | 1.5031e-05 | 8.1318e-08 |  |  |  | 1 | 0 | 0 | 20.2846 | 4.5075 | 7.0285 | 3.1554e-05 | 7.4024 | 2.9352e-05 | 1.2355 | 3.1223e-06 | 4.665e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7534 | 20.2846 | 4.5075 | 7.0285 | 7.4024 | 1.2355 | 0.22221 | 0.34649 | 0.36493 | 0.060906 | 1.5593 | 1.6422 | 0.27409 | 1.0532 | 0.17578 | 0.1669 | -1.5041 | -1.0599 | -1.0081 | -2.7984 | [10 empty] | 0 | 1.7534 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014066 | -1.8133 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 102.3% | 1.3769 | 1.5391 | 12.3557 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.65% | -7.4796 | -6.6957 | 3.5884 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_175 | iRM_12 |  |  | 0.00028171 | 6.7004e-05 | 9.0049e-05 | 1.9531e-05 | 9.0495e-05 | 4.2963e-06 | 0.00010102 | 6.9838e-06 | -3.2469e-06 | [5 empty] | 19.7219 | 4.3825 | 6.8336 | 5.0202e-05 | 7.1973 | 2.1727e-05 | 1.2013 | 1.4928e-05 | 9.8977e-08 |  |  |  | 1 | 0 | 0 | 19.7216 | 4.3824 | 6.8335 | 3.0671e-05 | 7.1973 | 1.7431e-05 | 1.2012 | 7.9438e-06 | 3.3459e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7544 | 19.7216 | 4.3824 | 6.8335 | 7.1973 | 1.2012 | 0.22222 | 0.3465 | 0.36494 | 0.060909 | 1.5593 | 1.6423 | 0.2741 | 1.0532 | 0.17579 | 0.1669 | -1.5041 | -1.0599 | -1.008 | -2.7984 | [10 empty] | 0 | 1.7544 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.56% | 2.1434 | -0.16094 | -3.3429 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.5% | 13.3519 | 0.70557 | 5.3558 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_180 | alfaA |  |  | 0.000256 | 5.0888e-05 | 9.4129e-05 | 1.1345e-05 | 6.9932e-05 | 2.205e-06 | 9.3296e-05 | 8.9298e-06 | -3.3738e-06 | [5 empty] | 20.4557 | 4.545 | 7.0861 | 4.173e-05 | 7.4614 | 2.2269e-05 | 1.2451 | 1.2274e-05 | -1.4967e-06 |  |  |  | 1 | 0 | 0 | 20.4555 | 4.5449 | 7.086 | 3.0384e-05 | 7.4613 | 2.0064e-05 | 1.245 | 3.3443e-06 | 1.8771e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7427 | 20.4555 | 4.5449 | 7.086 | 7.4613 | 1.245 | 0.22219 | 0.34641 | 0.36476 | 0.060865 | 1.5591 | 1.6417 | 0.27394 | 1.053 | 0.1757 | 0.16686 | -1.5042 | -1.0601 | -1.0085 | -2.7991 | [10 empty] | 0 | 1.7427 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.01407 | -1.8133 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 21, 2022 | 3 | 103.9% | -1.3967 | -3.1624 | 20.1242 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.62% | -14.6895 | 0.77946 | -16.2352 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_185 | iRM_12 |  |  | 0.00034794 | 7.2089e-05 | 0.0001222 | 1.2351e-05 | 9.6055e-05 | 2.4244e-07 | 0.0001025 | 2.4103e-06 | -1.8555e-06 | [5 empty] | 19.673 | 4.3717 | 6.8168 | 5.4071e-05 | 7.1798 | 2.5084e-05 | 1.1984 | 2.1464e-05 | 1.8038e-06 |  |  |  | 1 | 0 | 0 | 19.6726 | 4.3716 | 6.8167 | 4.172e-05 | 7.1797 | 2.4841e-05 | 1.1983 | 1.9053e-05 | 3.6593e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7554 | 19.6726 | 4.3716 | 6.8167 | 7.1797 | 1.1983 | 0.22222 | 0.34651 | 0.36496 | 0.060914 | 1.5593 | 1.6423 | 0.27412 | 1.0533 | 0.17579 | 0.16691 | -1.5041 | -1.0599 | -1.008 | -2.7983 | [10 empty] | 0 | 1.7554 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.83% | 4.9071 | 3.5781 | 1.1253 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.95% | 12.0145 | 6.188 | 11.6854 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_190 | Ames |  |  | 0.00026663 | 1.544e-05 | 9.9391e-05 | 8.5779e-06 | 7.8579e-05 | -6.6571e-07 | 7.3886e-05 | 9.2778e-06 | -4.2141e-06 | [5 empty] | 21.1027 | 4.6898 | 7.3135 | 3.8649e-05 | 7.7042 | 3.1937e-05 | 1.2862 | 1.8578e-05 | 8.1845e-06 |  |  |  | 1 | 0 | 0 | 21.1024 | 4.6898 | 7.3134 | 3.0071e-05 | 7.7041 | 3.2603e-05 | 1.2861 | 9.3003e-06 | 1.2399e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7632 | 21.1024 | 4.6898 | 7.3134 | 7.7041 | 1.2861 | 0.22224 | 0.34657 | 0.36508 | 0.060946 | 1.5594 | 1.6427 | 0.27423 | 1.0534 | 0.17586 | 0.16694 | -1.504 | -1.0597 | -1.0076 | -2.7978 | [10 empty] | 0 | 1.7632 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43911 | -1.3882 | 0.014065 | -1.8133 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 21, 2022 | 3 | 107.2% | 13.566 | 5.3926 | 20.0283 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.3% | -6.9055 | -5.9719 | 0.32936 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_195 | iRM_12 |  |  | 0.0003861 | 8.7318e-05 | 0.00013541 | 1.7409e-05 | 0.00011564 | -1.5199e-06 | 9.5006e-05 | 1.6578e-05 | -1.7602e-06 | [5 empty] | 19.6919 | 4.3759 | 6.8235 | 5.1991e-05 | 7.1871 | 2.2578e-05 | 1.1997 | 1.8662e-05 | -1.3657e-06 |  |  |  | 1 | 0 | 0 | 19.6915 | 4.3758 | 6.8234 | 3.4583e-05 | 7.187 | 2.4098e-05 | 1.1996 | 2.0841e-06 | 3.9449e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7566 | 19.6915 | 4.3758 | 6.8234 | 7.187 | 1.1996 | 0.22222 | 0.34651 | 0.36498 | 0.060919 | 1.5593 | 1.6424 | 0.27414 | 1.0533 | 0.17581 | 0.16691 | -1.5041 | -1.0598 | -1.0079 | -2.7982 | [10 empty] | 0 | 1.7566 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014073 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | -10.1474 | -8.323 | 4.5738 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.17% | -4.1228 | 0.45344 | 2.9983 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_200 | IPGP |  |  | 0.0002536 | 5.8731e-05 | 9.3242e-05 | 9.8416e-06 | 7.5523e-05 | 2.3219e-06 | 9.8858e-05 | 8.0749e-06 | -5.2264e-06 | [5 empty] | 20.8987 | 4.6441 | 7.2415 | 3.9906e-05 | 7.6271 | 2.5678e-05 | 1.2731 | 1.4526e-05 | -5.1034e-07 |  |  |  | 1 | 0 | 0 | 20.8985 | 4.644 | 7.2414 | 3.0064e-05 | 7.627 | 2.3356e-05 | 1.273 | 6.4507e-06 | 4.716e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7551 | 20.8985 | 4.644 | 7.2414 | 7.627 | 1.273 | 0.22222 | 0.34651 | 0.36495 | 0.060913 | 1.5593 | 1.6423 | 0.27411 | 1.0532 | 0.17579 | 0.1669 | -1.5041 | -1.0599 | -1.008 | -2.7983 | [10 empty] | 0 | 1.7551 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014065 | -1.8133 | -1.8274 | [14 empty] | IPGP | Lab 1 | Jun 21, 2022 | 3 | 106.5% | 3.6176 | 3.2778 | -2.4129 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.2% | 4.9972 | 0.24743 | -11.5161 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_205 | iRM_12 |  |  | 0.00025002 | 6.4112e-05 | 9.633e-05 | 1.2267e-05 | 8.7921e-05 | 2.0202e-06 | 0.00010086 | 1.3561e-05 | -4.6349e-06 | [5 empty] | 19.5492 | 4.3443 | 6.7744 | 5.1758e-05 | 7.1355 | 2.4684e-05 | 1.1911 | 1.7429e-05 | -2.4599e-08 |  |  |  | 1 | 0 | 0 | 19.549 | 4.3443 | 6.7743 | 3.9491e-05 | 7.1354 | 2.2664e-05 | 1.191 | 3.868e-06 | 4.6103e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.758 | 19.549 | 4.3443 | 6.7743 | 7.1354 | 1.191 | 0.22223 | 0.34653 | 0.365 | 0.060924 | 1.5594 | 1.6425 | 0.27415 | 1.0533 | 0.17581 | 0.16691 | -1.5041 | -1.0598 | -1.0079 | -2.7981 | [10 empty] | 0 | 1.758 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43911 | -1.3882 | 0.014063 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.43% | 13.1667 | 5.4013 | -11.6894 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.5% | -1.9367 | -0.79692 | 16.4123 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_210 | iRM_UR |  |  | 0.00026567 | 6.0832e-05 | 9.8925e-05 | 6.939e-06 | 6.9082e-05 | -2.8722e-06 | 8.6804e-05 | 1.1356e-05 | -2.5403e-06 | [5 empty] | 20.923 | 4.6495 | 7.2502 | 3.5881e-05 | 7.6365 | 2.4684e-05 | 1.2747 | 1.5286e-05 | -4.1772e-06 |  |  |  | 1 | 0 | 0 | 20.9227 | 4.6494 | 7.2501 | 2.8942e-05 | 7.6364 | 2.7556e-05 | 1.2746 | 3.9294e-06 | -1.6369e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7568 | 20.9227 | 4.6494 | 7.2501 | 7.6364 | 1.2746 | 0.22222 | 0.34652 | 0.36498 | 0.060919 | 1.5593 | 1.6424 | 0.27414 | 1.0533 | 0.1758 | 0.16691 | -1.5041 | -1.0598 | -1.0079 | -2.7982 | [10 empty] | 0 | 1.7568 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014072 | -1.8133 | -1.8274 | [14 empty] | iRM_UR | Lab 1 | Jun 21, 2022 | 3 | 106.8% | -10.1708 | -8.498 | -10.9677 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.53% | 5.1762 | 2.8216 | -13.1988 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_215 | iRM_12 |  |  | 0.0002685 | 6.0948e-05 | 8.8768e-05 | 1.1811e-05 | 7.2506e-05 | -2.5058e-07 | 9.2354e-05 | 9.1399e-06 | -1.2068e-06 | [5 empty] | 19.6291 | 4.3621 | 6.8021 | 4.9273e-05 | 7.1648 | 2.358e-05 | 1.196 | 1.8551e-05 | -2.3904e-07 |  |  |  | 1 | 0 | 0 | 19.6289 | 4.362 | 6.8021 | 3.7463e-05 | 7.1648 | 2.3831e-05 | 1.1959 | 9.4111e-06 | 9.6773e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7587 | 19.6289 | 4.362 | 6.8021 | 7.1648 | 1.1959 | 0.22222 | 0.34653 | 0.36501 | 0.060928 | 1.5594 | 1.6425 | 0.27417 | 1.0533 | 0.17582 | 0.16692 | -1.5041 | -1.0598 | -1.0078 | -2.7981 | [10 empty] | 0 | 1.7587 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014064 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100% | -11.0139 | 0.42583 | 25.8053 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.7% | -3.6455 | -5.2016 | 2.3958 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_220 | SRM |  |  | 0.0016862 | 0.00038174 | 0.00058742 | 6.2096e-06 | 0.00060002 | 1.4829e-05 | 0.0001803 | 1.1956e-05 | -4.6059e-06 | [5 empty] | 20.4275 | 4.5395 | 7.0788 | 5.2406e-05 | 7.4562 | 3.9564e-05 | 1.2447 | 1.5336e-05 | -2.3684e-07 |  |  |  | 1 | 0 | 0 | 20.4258 | 4.5391 | 7.0782 | 4.6196e-05 | 7.4556 | 2.4735e-05 | 1.2445 | 3.3804e-06 | 4.369e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7586 | 20.4258 | 4.5391 | 7.0782 | 7.4556 | 1.2445 | 0.22223 | 0.34653 | 0.36501 | 0.060927 | 1.5594 | 1.6425 | 0.27417 | 1.0533 | 0.17582 | 0.16692 | -1.5041 | -1.0598 | -1.0078 | -2.7981 | [10 empty] | 0 | 1.7586 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 103.9% | 3.9051 | -2.9282 | 4.7037 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.08% | 3.964 | 4.0478 | 2.8033 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_225 | iRM_12 |  |  | 0.00026828 | 5.8603e-05 | 9.0803e-05 | 1.2996e-05 | 8.6502e-05 | 1.2397e-06 | 9.3678e-05 | 1.6182e-05 | -4.0779e-06 | [5 empty] | 19.6963 | 4.377 | 6.8255 | 5.3217e-05 | 7.1895 | 2.4796e-05 | 1.2001 | 1.7304e-05 | -1.4039e-06 |  |  |  | 1 | 0 | 0 | 19.696 | 4.377 | 6.8254 | 4.022e-05 | 7.1894 | 2.3557e-05 | 1.2 | 1.122e-06 | 2.674e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7591 | 19.696 | 4.377 | 6.8254 | 7.1894 | 1.2 | 0.22223 | 0.34654 | 0.36502 | 0.060927 | 1.5594 | 1.6425 | 0.27417 | 1.0533 | 0.17582 | 0.16692 | -1.5041 | -1.0598 | -1.0078 | -2.7981 | [10 empty] | 0 | 1.7591 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3883 | 0.014066 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | 4.1844 | 1.1252 | -24.4888 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 101.1% | -5.7801 | -2.035 | -4.7261 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_230 | alfaA |  |  | 0.00026426 | 5.6135e-05 | 9.3902e-05 | 9.7935e-06 | 6.8118e-05 | 1.318e-06 | 9.352e-05 | 1.7663e-05 | -4.7023e-06 | [5 empty] | 20.4173 | 4.5367 | 7.0735 | 3.916e-05 | 7.4489 | 2.7495e-05 | 1.2432 | 1.9458e-05 | 9.4011e-08 |  |  |  | 1 | 0 | 0 | 20.417 | 4.5366 | 7.0734 | 2.9367e-05 | 7.4488 | 2.6177e-05 | 1.2431 | 1.7949e-06 | 4.7963e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7475 | 20.417 | 4.5366 | 7.0734 | 7.4488 | 1.2431 | 0.2222 | 0.34644 | 0.36483 | 0.060883 | 1.5592 | 1.6419 | 0.274 | 1.0531 | 0.17574 | 0.16688 | -1.5042 | -1.06 | -1.0083 | -2.7988 | [10 empty] | 0 | 1.7475 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014072 | -1.8133 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 21, 2022 | 3 | 103.8% | 1.8771 | -4.8042 | 20.9231 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.1% | 4.1641 | 7.4771 | 12.2294 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_235 | iRM_12 |  |  | 0.00041174 | 9.4828e-05 | 0.00013824 | 9.4064e-06 | 0.00013754 | 1.4001e-06 | 9.9546e-05 | 1.4338e-05 | -8.8019e-06 | [5 empty] | 19.6251 | 4.3613 | 6.8009 | 5.1186e-05 | 7.1638 | 2.4461e-05 | 1.1959 | 2.0715e-05 | -5.7238e-07 |  |  |  | 1 | 0 | 0 | 19.6247 | 4.3612 | 6.8008 | 4.178e-05 | 7.1636 | 2.3061e-05 | 1.1958 | 6.3775e-06 | 8.2296e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7601 | 19.6247 | 4.3612 | 6.8008 | 7.1636 | 1.1958 | 0.22223 | 0.34654 | 0.36503 | 0.060932 | 1.5594 | 1.6426 | 0.27419 | 1.0534 | 0.17583 | 0.16692 | -1.504 | -1.0598 | -1.0078 | -2.798 | [10 empty] | 0 | 1.76 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.99% | 1.3538 | -1.198 | 9.2785 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.2% | -8.8991 | -11.6156 | -8.1548 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_240 | Ames |  |  | 0.00025787 | 6.36e-05 | 9.415e-05 | 1.0718e-05 | 7.1099e-05 | 3.8562e-06 | 9.3795e-05 | 1.0868e-05 | -5.514e-06 | [5 empty] | 21.0259 | 4.6729 | 7.2876 | 4.3586e-05 | 7.6776 | 3.1911e-05 | 1.2819 | 2.3708e-05 | 8.6315e-06 |  |  |  | 1 | 0 | 0 | 21.0256 | 4.6728 | 7.2875 | 3.2868e-05 | 7.6775 | 2.8055e-05 | 1.2818 | 1.2841e-05 | 1.4146e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7676 | 21.0256 | 4.6728 | 7.2875 | 7.6775 | 1.2818 | 0.22225 | 0.3466 | 0.36515 | 0.060961 | 1.5595 | 1.643 | 0.2743 | 1.0535 | 0.17588 | 0.16695 | -1.504 | -1.0596 | -1.0074 | -2.7975 | [10 empty] | 0 | 1.7675 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [14 empty] | Ames | Lab 1 | Jun 21, 2022 | 3 | 107.3% | -8.9372 | 1.7677 | -0.70539 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.1% | 10.5562 | 6.1203 | -14.1507 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_245 | iRM_12 |  |  | 0.00026493 | 5.6805e-05 | 8.6858e-05 | 2.7999e-05 | 6.8289e-05 | 2.7811e-06 | 9.656e-05 | 1.051e-05 | -4.7673e-06 | [5 empty] | 19.5568 | 4.3462 | 6.7775 | 4.9226e-05 | 7.1393 | 2.4514e-05 | 1.1918 | 2.2886e-05 | -3.7191e-07 |  |  |  | 1 | 0 | 0 | 19.5566 | 4.3461 | 6.7774 | 2.1228e-05 | 7.1392 | 2.1733e-05 | 1.1917 | 1.2375e-05 | 4.3954e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7615 | 19.5566 | 4.3461 | 6.7774 | 7.1392 | 1.1917 | 0.22223 | 0.34655 | 0.36506 | 0.060938 | 1.5594 | 1.6427 | 0.27421 | 1.0534 | 0.17584 | 0.16693 | -1.504 | -1.0597 | -1.0077 | -2.7979 | [10 empty] | 0 | 1.7615 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.72% | -2.2327 | -0.61844 | 0.11641 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.8% | -4.5712 | -2.1478 | 14.9932 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_250 | IPGP |  |  | 0.00024953 | 5.0909e-05 | 8.5006e-05 | 9.2022e-06 | 6.6071e-05 | 1.7858e-06 | 0.00010245 | 2.2225e-05 | -4.9431e-06 | [5 empty] | 20.9124 | 4.6473 | 7.2469 | 3.5373e-05 | 7.6336 | 2.6107e-05 | 1.2743 | 2.3404e-05 | 9.1228e-07 |  |  |  | 1 | 0 | 0 | 20.9122 | 4.6473 | 7.2469 | 2.6171e-05 | 7.6335 | 2.4321e-05 | 1.2742 | 1.1789e-06 | 5.8554e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7598 | 20.9122 | 4.6473 | 7.2469 | 7.6335 | 1.2742 | 0.22223 | 0.34654 | 0.36503 | 0.060931 | 1.5594 | 1.6426 | 0.27418 | 1.0534 | 0.17583 | 0.16692 | -1.504 | -1.0598 | -1.0078 | -2.798 | [10 empty] | 0 | 1.7597 | 0.21794 | 0.33337 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8274 | [14 empty] | IPGP | Lab 1 | Jun 21, 2022 | 3 | 106.8% | 5.9616 | -4.5258 | 1.476 | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.7% | 0.97368 | 1.9467 | 6.4441 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_255 | iRM_12 |  |  | 0.0002864 | 5.9004e-05 | 0.00010457 | 2.0533e-05 | 7.9247e-05 | -3.1079e-07 | 9.6832e-05 | 1.4397e-05 | -6.057e-06 | [5 empty] | 19.5968 | 4.3551 | 6.7914 | 5.252e-05 | 7.1541 | 2.2552e-05 | 1.1943 | 2.3124e-05 | 6.8684e-07 |  |  |  | 1 | 0 | 0 | 19.5965 | 4.355 | 6.7913 | 3.1987e-05 | 7.1541 | 2.2863e-05 | 1.1942 | 8.7264e-06 | 6.7439e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7623 | 19.5965 | 4.355 | 6.7913 | 7.1541 | 1.1942 | 0.22223 | 0.34656 | 0.36507 | 0.06094 | 1.5594 | 1.6427 | 0.27422 | 1.0534 | 0.17584 | 0.16693 | -1.504 | -1.0597 | -1.0077 | -2.7979 | [10 empty] | 0 | 1.7623 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.95% | 3.9704 | 0.8925 | -12.5314 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | -10.4279 | 5.0601 | -39.5082 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_260 | iRM_UR |  |  | 0.0049635 | 0.001099 | 0.0017253 | 1.4218e-05 | 0.0017935 | 1.0497e-06 | 0.00038294 | 1.7333e-05 | -1.6929e-06 | [5 empty] | 20.8576 | 4.6353 | 7.2283 | 3.4055e-05 | 7.6144 | 2.2941e-05 | 1.2711 | 1.4248e-05 | -4.0407e-06 |  |  |  | 1 | 0 | 0 | 20.8527 | 4.6342 | 7.2266 | 1.9837e-05 | 7.6126 | 2.1891e-05 | 1.2708 | -3.0848e-06 | -2.3479e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7621 | 20.8527 | 4.6342 | 7.2266 | 7.6126 | 1.2708 | 0.22223 | 0.34656 | 0.36507 | 0.06094 | 1.5594 | 1.6427 | 0.27422 | 1.0534 | 0.17584 | 0.16693 | -1.504 | -1.0597 | -1.0077 | -2.7979 | [10 empty] | 0 | 1.7621 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014074 | -1.8133 | -1.8274 | [14 empty] | iRM_UR | Lab 1 | Jun 21, 2022 | 3 | 106.3% | 4.7491 | -1.4878 | -5.4991 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.57% | 6.9685 | -0.088558 | 31.3635 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_265 | iRM_12 |  |  | 0.00027151 | 6.7049e-05 | 9.3387e-05 | 9.4334e-06 | 8.599e-05 | -4.8554e-07 | 8.5991e-05 | 1.6167e-05 | 1.6162e-07 | [5 empty] | 19.655 | 4.368 | 6.8116 | 4.9236e-05 | 7.1755 | 2.3159e-05 | 1.1979 | 1.9431e-05 | -7.8298e-07 |  |  |  | 1 | 0 | 0 | 19.6548 | 4.3679 | 6.8115 | 3.9803e-05 | 7.1754 | 2.3645e-05 | 1.1978 | 3.2639e-06 | -9.446e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7625 | 19.6548 | 4.3679 | 6.8115 | 7.1754 | 1.1978 | 0.22223 | 0.34656 | 0.36507 | 0.060943 | 1.5594 | 1.6427 | 0.27423 | 1.0534 | 0.17585 | 0.16693 | -1.504 | -1.0597 | -1.0077 | -2.7978 | [10 empty] | 0 | 1.7625 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014073 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.7% | -8.061 | -3.5431 | 19.8195 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.46% | 0.27098 | -1.3111 | 6.4189 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_270 | SRM |  |  | 0.00027133 | 6.2207e-05 | 8.294e-05 | 1.2909e-05 | 8.1198e-05 | 8.8027e-07 | 8.094e-05 | 1.5483e-05 | -2.9964e-06 | [5 empty] | 20.6314 | 4.585 | 7.1501 | 5.4822e-05 | 7.5321 | 2.5693e-05 | 1.2575 | 1.868e-05 | -5.7469e-07 |  |  |  | 1 | 0 | 0 | 20.6311 | 4.585 | 7.15 | 4.1913e-05 | 7.532 | 2.4813e-05 | 1.2574 | 3.1966e-06 | 2.4217e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7631 | 20.6311 | 4.585 | 7.15 | 7.532 | 1.2574 | 0.22224 | 0.34656 | 0.36508 | 0.060946 | 1.5594 | 1.6428 | 0.27424 | 1.0534 | 0.17586 | 0.16694 | -1.504 | -1.0597 | -1.0076 | -2.7978 | [10 empty] | 0 | 1.7631 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014074 | -1.8133 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 105.5% | 3.8216 | -3.7751 | 22.1838 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.5% | -2.9043 | 4.2256 | -20.8563 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_275 | iRM_12 |  |  | 0.00026639 | 6.0704e-05 | 0.00010533 | 2.4868e-05 | 7.3623e-05 | 5.7086e-06 | 8.7109e-05 | 2.305e-05 | -2.7111e-06 | [5 empty] | 19.4505 | 4.3226 | 6.7409 | 5.2634e-05 | 7.1011 | 2.1169e-05 | 1.1855 | 2.1758e-05 | 1.002e-06 |  |  |  | 1 | 0 | 0 | 19.4502 | 4.3226 | 6.7408 | 2.7766e-05 | 7.1011 | 1.5461e-05 | 1.1854 | -1.2914e-06 | 3.713e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7636 | 19.4502 | 4.3226 | 6.7408 | 7.1011 | 1.1854 | 0.22224 | 0.34657 | 0.36509 | 0.060946 | 1.5595 | 1.6428 | 0.27424 | 1.0534 | 0.17585 | 0.16693 | -1.504 | -1.0597 | -1.0076 | -2.7978 | [10 empty] | 0 | 1.7636 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.5% | 4.1712 | 5.9696 | -10.5606 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.2% | 3.8579 | -5.1702 | -5.2927 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_280 | alfaA |  |  | 0.0030738 | 0.00067615 | 0.0010714 | 6.4137e-06 | 0.0011148 | 1.0149e-07 | 0.00025844 | 1.0354e-05 | -1.9092e-06 | [5 empty] | 20.4532 | 4.5448 | 7.0865 | 4.2248e-05 | 7.4632 | 2.4293e-05 | 1.2456 | 2.0479e-05 | 4.9954e-07 |  |  |  | 1 | 0 | 0 | 20.4501 | 4.5442 | 7.0854 | 3.5835e-05 | 7.4621 | 2.4192e-05 | 1.2454 | 1.0124e-05 | 2.4088e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7512 | 20.4501 | 4.5442 | 7.0854 | 7.4621 | 1.2454 | 0.22221 | 0.34647 | 0.36489 | 0.060898 | 1.5592 | 1.6421 | 0.27406 | 1.0532 | 0.17577 | 0.16689 | -1.5041 | -1.06 | -1.0082 | -2.7986 | [10 empty] | 0 | 1.7512 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 21, 2022 | 3 | 105.2% | -3.1881 | -4.6042 | 16.8728 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.4% | -2.9433 | 3.9138 | 23.3935 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_285 | iRM_12 |  |  | 0.00027278 | 5.9978e-05 | 9.1484e-05 | 1.8732e-05 | 7.232e-05 | -2.5961e-06 | 8.9392e-05 | 9.2839e-06 | -2.5182e-06 | [5 empty] | 19.4406 | 4.3205 | 6.7378 | 5.2461e-05 | 7.0981 | 2.366e-05 | 1.185 | 2.1277e-05 | 2.1915e-08 |  |  |  | 1 | 0 | 0 | 19.4404 | 4.3205 | 6.7377 | 3.3728e-05 | 7.098 | 2.6256e-05 | 1.185 | 1.1993e-05 | 2.5401e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7654 | 19.4404 | 4.3205 | 6.7377 | 7.098 | 1.185 | 0.22224 | 0.34658 | 0.36512 | 0.060953 | 1.5595 | 1.6429 | 0.27426 | 1.0535 | 0.17587 | 0.16694 | -1.504 | -1.0596 | -1.0075 | -2.7976 | [10 empty] | 0 | 1.7654 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 99.78% | 1.6991 | -7.5514 | -1.9192 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.2% | 1.3433 | -6.4545 | -13.1573 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_290 | Ames |  |  | 0.00025194 | 5.2528e-05 | 8.4154e-05 | 1.6286e-05 | 7.7163e-05 | 1.8672e-07 | 8.6167e-05 | 1.401e-05 | -5.5858e-06 | [5 empty] | 20.9624 | 4.659 | 7.2662 | 4.5942e-05 | 7.6559 | 2.8966e-05 | 1.2783 | 2.1529e-05 | 1.001e-05 |  |  |  | 1 | 0 | 0 | 20.9621 | 4.659 | 7.2661 | 2.9656e-05 | 7.6558 | 2.8779e-05 | 1.2782 | 7.5186e-06 | 1.5596e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7719 | 20.9621 | 4.659 | 7.2661 | 7.6558 | 1.2782 | 0.22226 | 0.34663 | 0.36522 | 0.060978 | 1.5596 | 1.6432 | 0.27436 | 1.0536 | 0.17591 | 0.16696 | -1.5039 | -1.0595 | -1.0073 | -2.7972 | [10 empty] | 0 | 1.7719 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5514 | 0.24951 | 1.0142 | 0.16311 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.01407 | -1.8133 | -1.8274 | [14 empty] | Ames | Lab 1 | Jun 21, 2022 | 3 | 107.6% | -7.0882 | 1.0943 | -21.9137 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | 3.9851 | 4.2598 | 7.4186 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_295 | iRM_12 |  |  | 0.00025235 | 4.9052e-05 | 9.9193e-05 | 1.5181e-05 | 6.5424e-05 | -9.7546e-07 | 9.6234e-05 | 9.4129e-06 | -8.0144e-07 | [5 empty] | 19.5153 | 4.3371 | 6.7637 | 5.3223e-05 | 7.1253 | 2.3345e-05 | 1.1896 | 1.5404e-05 | 5.1221e-07 |  |  |  | 1 | 0 | 0 | 19.5151 | 4.3371 | 6.7636 | 3.8042e-05 | 7.1252 | 2.4321e-05 | 1.1895 | 5.9914e-06 | 1.3136e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.765 | 19.5151 | 4.3371 | 6.7636 | 7.1252 | 1.1895 | 0.22224 | 0.34658 | 0.36511 | 0.060952 | 1.5595 | 1.6429 | 0.27426 | 1.0535 | 0.17587 | 0.16694 | -1.504 | -1.0596 | -1.0076 | -2.7977 | [10 empty] | 0 | 1.765 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014066 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 21, 2022 | 3 | 100.4% | 1.0968 | 8.5033 | 3.6459 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.78% | -5.622 | 4.4603 | -3.1619 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_300 | IPGP |  |  | 0.00024868 | 5.7561e-05 | 9.3947e-05 | 1.2912e-05 | 6.5864e-05 | -2.279e-06 | 8.0748e-05 | 7.9043e-06 | 4.4017e-08 | [5 empty] | 20.8631 | 4.6366 | 7.2305 | 3.8768e-05 | 7.6169 | 2.6621e-05 | 1.2716 | 1.8753e-05 | 2.4443e-06 |  |  |  | 1 | 0 | 0 | 20.8629 | 4.6365 | 7.2304 | 2.5856e-05 | 7.6169 | 2.89e-05 | 1.2716 | 1.0849e-05 | 2.4003e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7638 | 20.8629 | 4.6365 | 7.2304 | 7.6169 | 1.2715 | 0.22224 | 0.34657 | 0.36509 | 0.060948 | 1.5594 | 1.6428 | 0.27425 | 1.0534 | 0.17586 | 0.16694 | -1.504 | -1.0597 | -1.0076 | -2.7977 | [10 empty] | 0 | 1.7638 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014074 | -1.8133 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 21, 2022 | 3 | 107.1% | 3.5986 | -3.194 | 10.1371 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.1% | 6.6088 | -9.3927 | 3.0602 | iRM_12
Lab 1 | Jun 21, 2022 | 3 | Zr_305 | iRM_12 |  |  | 0.0011385 | 0.00023801 | 0.00039713 | 1.1381e-05 | 0.0003928 | 1.0762e-06 | 0.00012957 | 1.9496e-05 | -3.0202e-06 | [5 empty] | 19.435 | 4.3193 | 6.7359 | 5.3175e-05 | 7.0962 | 2.3205e-05 | 1.1847 | 2.0221e-05 | 6.7106e-07 |  |  |  | 1 | 0 | 0 | 19.4338 | 4.319 | 6.7355 | 4.1794e-05 | 7.0958 | 2.2129e-05 | 1.1846 | 7.257e-07 | 3.6912e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7659 | 19.4338 | 4.319 | 6.7355 | 7.0958 | 1.1846 | 0.22224 | 0.34658 | 0.36513 | 0.060956 | 1.5595 | 1.6429 | 0.27428 | 1.0535 | 0.17588 | 0.16694 | -1.504 | -1.0596 | -1.0075 | -2.7976 | [10 empty] | 0 | 1.7659 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [22 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.6% | -13.2211 | 7.3686 | -8.8603 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_004 | iRM_12 |  |  | 9.7585e-05 | 2.774e-05 | 3.9918e-05 | 1.2252e-05 | 3.3099e-05 | 4.6896e-06 | 7.7748e-05 | 1.1084e-05 | -5.8377e-06 | [5 empty] | 20.0824 | 4.4631 | 6.96 | 5.5381e-05 | 7.3319 | 2.6196e-05 | 1.224 | 1.4947e-05 | 1.3222e-06 |  |  |  | 1 | 0 | 0 | 20.0823 | 4.4631 | 6.96 | 4.3129e-05 | 7.3319 | 2.1506e-05 | 1.224 | 3.8636e-06 | 7.1599e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7639 | 20.0823 | 4.4631 | 6.96 | 7.3319 | 1.2239 | 0.22224 | 0.34657 | 0.36509 | 0.060947 | 1.5595 | 1.6428 | 0.27424 | 1.0534 | 0.17586 | 0.16693 | -1.504 | -1.0597 | -1.0076 | -2.7978 | [9 empty] | L1 JunS4 | 0 | 1.7639 | 0.21794 | 0.33338 | 0.3381 | 0.054379 | 1.5297 | 1.5513 | 0.24951 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014063 | -1.8133 | -1.8274 | [22 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.34% | 12.0088 | -0.49784 | -1.1533 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_008 | iRM_UR |  |  | 0.00019818 | 3.7936e-05 | 7.744e-05 | 2.1545e-05 | 6.4446e-05 | 3.0464e-06 | 6.5027e-05 | 5.9967e-06 | -3.177e-06 | [5 empty] | 21.123 | 4.6945 | 7.321 | 3.9296e-05 | 7.7125 | 2.6239e-05 | 1.2877 | 1.4086e-05 | -3.0633e-06 |  |  |  | 1 | 0 | 0 | 21.1228 | 4.6944 | 7.3209 | 1.7751e-05 | 7.7125 | 2.3193e-05 | 1.2876 | 8.0893e-06 | 1.1368e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7659 | 21.1228 | 4.6944 | 7.3209 | 7.7125 | 1.2876 | 0.22224 | 0.34659 | 0.36513 | 0.060957 | 1.5595 | 1.6429 | 0.27428 | 1.0535 | 0.17588 | 0.16695 | -1.504 | -1.0596 | -1.0075 | -2.7976 | [10 empty] | 0 | 1.7659 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 4 | 106.4% | 3.0552 | 0.20533 | 29.1913 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | -4.1567 | -4.644 | 16.8741 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_012 | iRM_12 |  |  | 0.0002315 | 3.7401e-05 | 8.3323e-05 | 1.5968e-05 | 6.722e-05 | 4.2253e-06 | 7.5306e-05 | 1.6151e-05 | -5.3874e-06 | [5 empty] | 19.6124 | 4.3588 | 6.7977 | 5.1732e-05 | 7.1618 | 2.5699e-05 | 1.1958 | 1.125e-05 | 9.0239e-07 |  |  |  | 1 | 0 | 0 | 19.6121 | 4.3588 | 6.7976 | 3.5764e-05 | 7.1617 | 2.1474e-05 | 1.1957 | -4.9014e-06 | 6.2898e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7685 | 19.6121 | 4.3588 | 6.7976 | 7.1617 | 1.1957 | 0.22225 | 0.3466 | 0.36517 | 0.060967 | 1.5595 | 1.6431 | 0.27432 | 1.0536 | 0.1759 | 0.16696 | -1.504 | -1.0596 | -1.0074 | -2.7974 | [10 empty] | 0 | 1.7685 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.03% | 1.7231 | -6.403 | 4.264 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.48% | 3.0948 | 5.6061 | -9.0191 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_016 | SRM |  |  | 0.0002932 | 5.5802e-05 | 0.00010036 | 1.07e-05 | 8.7347e-05 | 1.5985e-06 | 7.4962e-05 | 1.8473e-05 | -4.8897e-06 | [5 empty] | 20.8355 | 4.6307 | 7.2218 | 5.2538e-05 | 7.6085 | 2.8318e-05 | 1.2704 | 1.8934e-05 | 2.7426e-06 |  |  |  | 1 | 0 | 0 | 20.8352 | 4.6306 | 7.2217 | 4.1838e-05 | 7.6084 | 2.672e-05 | 1.2703 | 4.6079e-07 | 7.6323e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7688 | 20.8352 | 4.6306 | 7.2217 | 7.6084 | 1.2703 | 0.22225 | 0.34661 | 0.36517 | 0.060968 | 1.5595 | 1.6431 | 0.27432 | 1.0536 | 0.1759 | 0.16696 | -1.504 | -1.0596 | -1.0074 | -2.7974 | [10 empty] | 0 | 1.7688 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014068 | -1.8133 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 4 | 106.5% | 9.7985 | 8.0267 | -4.1873 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | -1.906 | -4.2873 | -2.0858 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_020 | iRM_12 |  |  | 0.0017894 | 0.00040211 | 0.00062126 | 1.8925e-05 | 0.00063502 | 3.4909e-06 | 0.00015926 | 2.0326e-05 | -2.6225e-06 | [5 empty] | 19.5274 | 4.34 | 6.7686 | 5.1245e-05 | 7.1315 | 2.6691e-05 | 1.1908 | 1.9893e-05 | 1.4711e-06 |  |  |  | 1 | 0 | 0 | 19.5256 | 4.3396 | 6.768 | 3.2321e-05 | 7.1309 | 2.32e-05 | 1.1906 | -4.3359e-07 | 4.0936e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7711 | 19.5256 | 4.3396 | 6.768 | 7.1309 | 1.1906 | 0.22225 | 0.34662 | 0.36521 | 0.060978 | 1.5596 | 1.6432 | 0.27436 | 1.0536 | 0.17592 | 0.16697 | -1.5039 | -1.0595 | -1.0073 | -2.7972 | [10 empty] | 0 | 1.7711 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014077 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.1% | -12.6831 | -4.9748 | 20.7813 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | -4.6501 | 4.2325 | -8.4201 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_024 | alfaA |  |  | 0.00024708 | 4.9942e-05 | 8.7698e-05 | 1.0148e-05 | 6.3784e-05 | 3.6906e-06 | 7.3216e-05 | 1.6169e-05 | -2.6174e-06 | [5 empty] | 20.6982 | 4.5996 | 7.1725 | 4.1922e-05 | 7.555 | 2.7839e-05 | 1.2612 | 2.2018e-05 | -1.1822e-06 |  |  |  | 1 | 0 | 0 | 20.698 | 4.5996 | 7.1724 | 3.1774e-05 | 7.5549 | 2.4149e-05 | 1.2611 | 5.8485e-06 | 1.4352e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7585 | 20.698 | 4.5996 | 7.1724 | 7.5549 | 1.2611 | 0.22222 | 0.34653 | 0.36501 | 0.060928 | 1.5594 | 1.6425 | 0.27418 | 1.0533 | 0.17583 | 0.16692 | -1.5041 | -1.0598 | -1.0078 | -2.7981 | [10 empty] | 0 | 1.7585 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014082 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 4 | 106.4% | -0.67379 | -9.76 | 11.8933 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.62% | 6.1272 | -6.7758 | 15.3736 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_028 | iRM_12 |  |  | 0.00025945 | 4.6356e-05 | 8.7191e-05 | 1.1264e-05 | 6.6571e-05 | -7.426e-07 | 7.1121e-05 | 9.9049e-06 | -6.207e-06 | [5 empty] | 19.3825 | 4.3079 | 6.7187 | 4.9966e-05 | 7.079 | 2.5628e-05 | 1.182 | 1.7075e-05 | 4.7193e-07 |  |  |  | 1 | 0 | 0 | 19.3823 | 4.3079 | 6.7186 | 3.8701e-05 | 7.0789 | 2.637e-05 | 1.1819 | 7.1702e-06 | 6.6789e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7721 | 19.3823 | 4.3079 | 6.7186 | 7.0789 | 1.1819 | 0.22226 | 0.34663 | 0.36522 | 0.060981 | 1.5596 | 1.6432 | 0.27437 | 1.0536 | 0.17592 | 0.16697 | -1.5039 | -1.0595 | -1.0072 | -2.7972 | [10 empty] | 0 | 1.7721 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.96% | 9.4255 | 9.1951 | -15.1327 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.72% | -4.9002 | 1.8043 | -16.8466 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_032 | Ames |  |  | 0.0004155 | 8.2443e-05 | 0.00014376 | 1.1099e-05 | 0.00012582 | 8.0092e-07 | 7.842e-05 | 1.8935e-05 | -2.6291e-06 | [5 empty] | 21.0474 | 4.6783 | 7.2969 | 3.9606e-05 | 7.6893 | 3.3139e-05 | 1.2841 | 1.541e-05 | 6.9912e-06 |  |  |  | 1 | 0 | 0 | 21.047 | 4.6783 | 7.2967 | 2.8507e-05 | 7.6892 | 3.2339e-05 | 1.284 | -3.5247e-06 | 9.6203e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7791 | 21.047 | 4.6783 | 7.2967 | 7.6892 | 1.284 | 0.22228 | 0.34669 | 0.36534 | 0.061008 | 1.5597 | 1.6436 | 0.27447 | 1.0538 | 0.17597 | 0.16699 | -1.5038 | -1.0593 | -1.0069 | -2.7967 | [10 empty] | 0 | 1.7791 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014067 | -1.8133 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 4 | 108.9% | 10.8387 | 4.3557 | -1.2916 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | 6.5609 | 9.3334 | 16.0675 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_036 | iRM_12 |  |  | 0.00024598 | 5.2674e-05 | 7.4214e-05 | 7.3352e-06 | 6.4856e-05 | 5.0247e-07 | 7.3832e-05 | -2.4419e-07 | -5.3544e-07 | [5 empty] | 19.2542 | 4.2794 | 6.6743 | 5.3241e-05 | 7.0325 | 2.3976e-05 | 1.1743 | 1.5956e-05 | -4.5629e-07 |  |  |  | 1 | 0 | 0 | 19.254 | 4.2794 | 6.6742 | 4.5906e-05 | 7.0324 | 2.3473e-05 | 1.1742 | 1.62e-05 | 7.9149e-08 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7734 | 19.254 | 4.2794 | 6.6742 | 7.0324 | 1.1742 | 0.22226 | 0.34664 | 0.36525 | 0.060986 | 1.5596 | 1.6433 | 0.27439 | 1.0537 | 0.17593 | 0.16697 | -1.5039 | -1.0595 | -1.0072 | -2.7971 | [10 empty] | 0 | 1.7734 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.65% | -7.4796 | -6.6957 | 3.5884 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.2% | -2.9453 | -10.3973 | -10.4189 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_040 | IPGP |  |  | 0.00024547 | 4.2915e-05 | 7.3885e-05 | 1.3363e-05 | 7.5367e-05 | 4.0617e-06 | 6.2785e-05 | 4.8751e-06 | -5.3215e-06 | [5 empty] | 20.8287 | 4.6293 | 7.2198 | 3.7491e-05 | 7.607 | 2.6507e-05 | 1.2702 | 2.1998e-05 | 9.8173e-07 |  |  |  | 1 | 0 | 0 | 20.8284 | 4.6293 | 7.2198 | 2.4128e-05 | 7.6069 | 2.2446e-05 | 1.2701 | 1.7122e-05 | 6.3032e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7717 | 20.8284 | 4.6293 | 7.2198 | 7.6069 | 1.2701 | 0.22226 | 0.34663 | 0.36522 | 0.060979 | 1.5596 | 1.6432 | 0.27436 | 1.0536 | 0.17592 | 0.16697 | -1.5039 | -1.0595 | -1.0073 | -2.7972 | [10 empty] | 0 | 1.7717 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 4 | 108.2% | 8.6327 | 2.8177 | 5.6446 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.41% | -5.884 | 4.7206 | 4.0243 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_044 | iRM_12 |  |  | 0.0026278 | 0.00057788 | 0.00090748 | 3.28e-06 | 0.00093715 | 2.2737e-06 | 0.00020731 | 1.2586e-05 | -1.583e-06 | [5 empty] | 19.264 | 4.2817 | 6.6778 | 4.7327e-05 | 7.0362 | 2.404e-05 | 1.1749 | 1.8111e-05 | 1.0666e-07 |  |  |  | 1 | 0 | 0 | 19.2613 | 4.2811 | 6.6769 | 4.4047e-05 | 7.0353 | 2.1767e-05 | 1.1747 | 5.5256e-06 | 1.6896e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.774 | 19.2613 | 4.2811 | 6.6769 | 7.0353 | 1.1747 | 0.22226 | 0.34665 | 0.36526 | 0.060988 | 1.5596 | 1.6433 | 0.2744 | 1.0537 | 0.17594 | 0.16697 | -1.5039 | -1.0594 | -1.0072 | -2.7971 | [10 empty] | 0 | 1.774 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014071 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.5% | 13.3519 | 0.70557 | 5.3558 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 101% | 3.5994 | -4.7899 | 8.7539 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_048 | iRM_UR |  |  | 0.0002453 | 6.3332e-05 | 8.1511e-05 | 5.8104e-06 | 7.9764e-05 | 2.9421e-06 | 5.6918e-05 | 1.2151e-05 | -1.7511e-06 | [5 empty] | 20.5098 | 4.5585 | 7.1096 | 3.6013e-05 | 7.4911 | 2.3163e-05 | 1.2509 | 1.3991e-05 | -2.6928e-06 |  |  |  | 1 | 0 | 0 | 20.5095 | 4.5584 | 7.1095 | 3.0202e-05 | 7.491 | 2.0221e-05 | 1.2508 | 1.8404e-06 | -9.4167e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7735 | 20.5095 | 4.5584 | 7.1095 | 7.491 | 1.2508 | 0.22226 | 0.34664 | 0.36525 | 0.060987 | 1.5596 | 1.6433 | 0.27439 | 1.0537 | 0.17594 | 0.16697 | -1.5039 | -1.0595 | -1.0072 | -2.7971 | [10 empty] | 0 | 1.7735 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014072 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 4 | 107% | -0.011583 | -2.0628 | 16.6269 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 98.63% | 7.4383 | 2.7903 | -13.9096 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_052 | iRM_12 |  |  | 0.00078366 | 0.00017446 | 0.00026628 | 5.545e-07 | 0.00025776 | 3.3946e-06 | 9.5833e-05 | 1.1179e-05 | -1.8758e-06 | [5 empty] | 19.081 | 4.241 | 6.6145 | 4.7462e-05 | 6.9697 | 2.3646e-05 | 1.1638 | 1.0444e-05 | 2.0015e-06 |  |  |  | 1 | 0 | 0 | 19.0802 | 4.2408 | 6.6143 | 4.6908e-05 | 6.9694 | 2.0252e-05 | 1.1637 | -7.3565e-07 | 3.8773e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.775 | 19.0802 | 4.2408 | 6.6143 | 6.9694 | 1.1637 | 0.22226 | 0.34666 | 0.36527 | 0.060992 | 1.5597 | 1.6434 | 0.27441 | 1.0537 | 0.17594 | 0.16698 | -1.5039 | -1.0594 | -1.0071 | -2.797 | [10 empty] | 0 | 1.775 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9118 | 0.42507 | 0.43914 | -1.3882 | 0.014068 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.62% | -14.6895 | 0.77946 | -16.2352 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 101% | -13.6222 | 0.63264 | 1.7217 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_056 | SRM |  |  | 0.00041852 | 9.7347e-05 | 0.00014627 | 1.3119e-05 | 0.00013236 | -1.9727e-07 | 7.6654e-05 | 1.9466e-05 | -2.255e-06 | [5 empty] | 20.2219 | 4.4945 | 7.0099 | 4.875e-05 | 7.3862 | 2.5791e-05 | 1.2334 | 1.942e-05 | -1.0609e-07 |  |  |  | 1 | 0 | 0 | 20.2215 | 4.4945 | 7.0097 | 3.5631e-05 | 7.386 | 2.5988e-05 | 1.2333 | -4.5622e-08 | 2.1489e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7741 | 20.2215 | 4.4945 | 7.0097 | 7.386 | 1.2333 | 0.22226 | 0.34665 | 0.36526 | 0.060989 | 1.5596 | 1.6434 | 0.2744 | 1.0537 | 0.17594 | 0.16698 | -1.5039 | -1.0594 | -1.0072 | -2.7971 | [10 empty] | 0 | 1.7741 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014074 | -1.8133 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 4 | 106.1% | -1.9784 | -7.1062 | 4.5582 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | 9.9671 | 1.8663 | 3.3676 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_060 | iRM_12 |  |  | 0.00026577 | 7.9287e-05 | 8.2622e-05 | 1.3597e-05 | 8.3168e-05 | 1.9281e-07 | 7.5524e-05 | 1.1794e-05 | -1.3048e-06 | [5 empty] | 19.0454 | 4.2331 | 6.6022 | 4.9269e-05 | 6.9568 | 2.3238e-05 | 1.1617 | 1.2849e-05 | 1.4206e-06 |  |  |  | 1 | 0 | 0 | 19.0451 | 4.2331 | 6.6021 | 3.5672e-05 | 6.9567 | 2.3045e-05 | 1.1616 | 1.0553e-06 | 2.7254e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7752 | 19.0451 | 4.2331 | 6.6021 | 6.9567 | 1.1616 | 0.22227 | 0.34666 | 0.36527 | 0.060994 | 1.5597 | 1.6434 | 0.27442 | 1.0537 | 0.17595 | 0.16698 | -1.5039 | -1.0594 | -1.0071 | -2.797 | [10 empty] | 0 | 1.7752 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014066 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.95% | 12.0145 | 6.188 | 11.6854 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.09% | -0.30235 | 1.3291 | 4.7308 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_064 | alfaA |  |  | 0.00058357 | 0.00014137 | 0.00019201 | 9.5775e-06 | 0.00019029 | 3.3095e-06 | 7.3917e-05 | 1.2001e-05 | -7.8531e-06 | [5 empty] | 20.2888 | 4.5089 | 7.0313 | 3.8174e-05 | 7.4069 | 2.5156e-05 | 1.2366 | 2.0875e-05 | -1.4783e-06 |  |  |  | 1 | 0 | 0 | 20.2882 | 4.5087 | 7.0311 | 2.8596e-05 | 7.4068 | 2.1847e-05 | 1.2365 | 8.8732e-06 | 6.3748e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7628 | 20.2882 | 4.5087 | 7.0311 | 7.4067 | 1.2365 | 0.22223 | 0.34656 | 0.36508 | 0.060945 | 1.5594 | 1.6428 | 0.27424 | 1.0534 | 0.17586 | 0.16694 | -1.504 | -1.0597 | -1.0076 | -2.7978 | [10 empty] | 0 | 1.7628 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014078 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 4 | 106.6% | 2.9121 | -5.881 | 10.0424 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | -4.9799 | -8.4168 | -17.1683 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_068 | iRM_12 |  |  | 0.00086598 | 0.00018941 | 0.00028876 | 1.2078e-05 | 0.0002957 | 1.7051e-06 | 9.7427e-05 | 1.0313e-05 | -2.789e-06 | [5 empty] | 19.0303 | 4.2298 | 6.5971 | 4.7222e-05 | 6.9516 | 2.2887e-05 | 1.1609 | 2.0335e-05 | -9.5232e-07 |  |  |  | 1 | 0 | 0 | 19.0294 | 4.2296 | 6.5968 | 3.5145e-05 | 6.9513 | 2.1182e-05 | 1.1608 | 1.0023e-05 | 1.8367e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7764 | 19.0294 | 4.2296 | 6.5968 | 6.9513 | 1.1608 | 0.22227 | 0.34666 | 0.36529 | 0.060998 | 1.5597 | 1.6435 | 0.27444 | 1.0537 | 0.17596 | 0.16698 | -1.5039 | -1.0594 | -1.0071 | -2.7969 | [10 empty] | 0 | 1.7764 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014077 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.3% | -6.9055 | -5.9719 | 0.32936 | [8 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.1% | 0.53191 | 7.8953 | 11.3902 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_072 | Ames |  |  | 0.00026119 | 5.2737e-05 | 9.4922e-05 | 1.616e-05 | 8.9058e-05 | 2.0118e-06 | 7.1958e-05 | 1.4646e-05 | -8.5988e-07 | [5 empty] | 20.687 | 4.5984 | 7.1726 | 3.7512e-05 | 7.5591 | 3.0526e-05 | 1.2625 | 1.5379e-05 | 7.2737e-06 |  |  |  | 1 | 0 | 0 | 20.6868 | 4.5983 | 7.1725 | 2.1352e-05 | 7.559 | 2.8515e-05 | 1.2624 | 7.3237e-07 | 8.1336e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7833 | 20.6868 | 4.5983 | 7.1725 | 7.559 | 1.2624 | 0.22228 | 0.34672 | 0.3654 | 0.061025 | 1.5598 | 1.6439 | 0.27453 | 1.0539 | 0.17601 | 0.16701 | -1.5038 | -1.0592 | -1.0068 | -2.7965 | [10 empty] | 0 | 1.7833 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.01407 | -1.8133 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 4 | 109.1% | 7.0113 | 6.8713 | -10.8784 | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | 3.2827 | -0.41137 | 10.3538 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_076 | iRM_12 |  |  | 0.0022178 | 0.00048874 | 0.00076463 | 1.2888e-05 | 0.00078216 | 1.2131e-05 | 0.0001835 | -6.3847e-08 | -4.0339e-06 | [5 empty] | 18.8867 | 4.1979 | 6.5474 | 4.8469e-05 | 6.8994 | 3.5785e-05 | 1.1522 | 1.7746e-05 | 1.518e-06 |  |  |  | 1 | 0 | 0 | 18.8845 | 4.1974 | 6.5467 | 3.5581e-05 | 6.8986 | 2.3654e-05 | 1.152 | 1.781e-05 | 5.5519e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7771 | 18.8845 | 4.1974 | 6.5467 | 6.8986 | 1.152 | 0.22227 | 0.34667 | 0.3653 | 0.061001 | 1.5597 | 1.6435 | 0.27445 | 1.0538 | 0.17596 | 0.16699 | -1.5039 | -1.0594 | -1.007 | -2.7969 | [10 empty] | 0 | 1.7771 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.17% | -4.1228 | 0.45344 | 2.9983 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.27% | -18.9051 | -3.2269 | -32.5442 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_080 | IPGP |  |  | 0.00024614 | 2.627e-05 | 9.222e-05 | 4.3576e-06 | 7.8337e-05 | 2.62e-06 | 5.6855e-05 | 6.4948e-06 | -3.3164e-06 | [5 empty] | 20.6735 | 4.595 | 7.1666 | 3.444e-05 | 7.5515 | 2.4564e-05 | 1.261 | 1.481e-05 | 2.3317e-06 |  |  |  | 1 | 0 | 0 | 20.6732 | 4.595 | 7.1665 | 3.0082e-05 | 7.5514 | 2.1944e-05 | 1.2609 | 8.3148e-06 | 5.6481e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7754 | 20.6732 | 4.595 | 7.1665 | 7.5514 | 1.2609 | 0.22227 | 0.34666 | 0.36528 | 0.060994 | 1.5596 | 1.6434 | 0.27442 | 1.0537 | 0.17595 | 0.16698 | -1.5039 | -1.0594 | -1.0071 | -2.797 | [10 empty] | 0 | 1.7754 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43912 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 4 | 109% | 11.425 | -0.43079 | 2.7877 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.9% | 17.4275 | -1.4466 | 30.5231 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_084 | iRM_12 |  |  | 0.00027124 | 4.835e-05 | 0.00010005 | 7.4584e-06 | 8.3012e-05 | -2.6797e-06 | 7.3124e-05 | 1.0988e-05 | -1.0822e-06 | [5 empty] | 19.0542 | 4.2352 | 6.6056 | 4.5608e-05 | 6.9607 | 2.3081e-05 | 1.1624 | 1.5328e-05 | 1.5451e-06 |  |  |  | 1 | 0 | 0 | 19.0539 | 4.2351 | 6.6055 | 3.8149e-05 | 6.9606 | 2.5761e-05 | 1.1623 | 4.3399e-06 | 2.6273e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7776 | 19.0539 | 4.2351 | 6.6055 | 6.9606 | 1.1623 | 0.22227 | 0.34667 | 0.36531 | 0.061002 | 1.5597 | 1.6435 | 0.27445 | 1.0538 | 0.17597 | 0.16699 | -1.5039 | -1.0594 | -1.007 | -2.7968 | [10 empty] | 0 | 1.7776 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.2% | 4.9972 | 0.24743 | -11.5161 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.77% | -4.3115 | 2.1552 | -5.4546 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_088 | iRM_UR |  |  | 0.00026919 | 5.007e-05 | 9.2206e-05 | 1.6685e-05 | 7.9498e-05 | 2.6119e-06 | 7.3793e-05 | 1.8527e-05 | -2.2682e-06 | [5 empty] | 20.4768 | 4.5515 | 7.0989 | 3.1895e-05 | 7.4808 | 2.2612e-05 | 1.2493 | 1.4867e-05 | -4.5495e-06 |  |  |  | 1 | 0 | 0 | 20.4765 | 4.5514 | 7.0989 | 1.521e-05 | 7.4807 | 2e-05 | 1.2492 | -3.6598e-06 | -2.2812e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7788 | 20.4765 | 4.5514 | 7.0989 | 7.4807 | 1.2492 | 0.22227 | 0.34668 | 0.36533 | 0.061007 | 1.5597 | 1.6436 | 0.27447 | 1.0538 | 0.17597 | 0.16699 | -1.5038 | -1.0593 | -1.007 | -2.7968 | [10 empty] | 0 | 1.7788 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 4 | 107.2% | 0.74473 | 0.67792 | -15.8672 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.42% | -4.2643 | 0.33657 | -21.2774 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_092 | iRM_12 |  |  | 0.00035482 | 7.4619e-05 | 0.0001209 | 6.2066e-06 | 0.00010933 | 2.8786e-07 | 6.6037e-05 | 3.7803e-06 | -8.5914e-07 | [5 empty] | 19.1471 | 4.256 | 6.6383 | 5.0332e-05 | 6.9956 | 2.327e-05 | 1.1683 | 2.0897e-05 | 4.4268e-07 |  |  |  | 1 | 0 | 0 | 19.1468 | 4.2559 | 6.6381 | 4.4126e-05 | 6.9955 | 2.2982e-05 | 1.1683 | 1.7117e-05 | 1.3018e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7808 | 19.1468 | 4.2559 | 6.6381 | 6.9955 | 1.1683 | 0.22228 | 0.3467 | 0.36536 | 0.061016 | 1.5597 | 1.6437 | 0.2745 | 1.0538 | 0.17599 | 0.167 | -1.5038 | -1.0593 | -1.0069 | -2.7966 | [10 empty] | 0 | 1.7808 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014076 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.5% | -1.9367 | -0.79692 | 16.4123 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.8% | 0.29366 | 5.2096 | 16.4023 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_096 | SRM |  |  | 0.0002436 | 5.5629e-05 | 7.9988e-05 | 1.2834e-05 | 7.408e-05 | -4.5876e-07 | 6.8117e-05 | 1.3539e-05 | -2.4566e-06 | [5 empty] | 20.7549 | 4.6133 | 7.1956 | 4.5055e-05 | 7.583 | 2.5232e-05 | 1.2664 | 1.0424e-05 | 2.4125e-06 |  |  |  | 1 | 0 | 0 | 20.7546 | 4.6133 | 7.1955 | 3.222e-05 | 7.5829 | 2.5691e-05 | 1.2663 | -3.1149e-06 | 4.8692e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7805 | 20.7546 | 4.6133 | 7.1955 | 7.5829 | 1.2663 | 0.22228 | 0.3467 | 0.36536 | 0.061014 | 1.5597 | 1.6437 | 0.2745 | 1.0538 | 0.17599 | 0.167 | -1.5038 | -1.0593 | -1.0069 | -2.7967 | [10 empty] | 0 | 1.7805 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014077 | -1.8133 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 4 | 108.7% | -9.2237 | -1.3602 | -4.1285 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.1% | -0.81306 | -9.5399 | -12.2388 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_100 | iRM_12 |  |  | 0.00025181 | 5.0879e-05 | 9.1057e-05 | 1.7148e-05 | 7.8923e-05 | 4.1864e-07 | 7.4995e-05 | 1.0802e-05 | 4.2391e-07 | [5 empty] | 19.0535 | 4.2352 | 6.6058 | 5.1843e-05 | 6.9614 | 2.2823e-05 | 1.1626 | 1.8562e-05 | 1.3632e-06 |  |  |  | 1 | 0 | 0 | 19.0532 | 4.2351 | 6.6057 | 3.4695e-05 | 6.9613 | 2.2404e-05 | 1.1625 | 7.7602e-06 | 9.3924e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7806 | 19.0532 | 4.2351 | 6.6057 | 6.9613 | 1.1625 | 0.22228 | 0.3467 | 0.36536 | 0.061014 | 1.5597 | 1.6437 | 0.27449 | 1.0538 | 0.17599 | 0.167 | -1.5038 | -1.0593 | -1.0069 | -2.7966 | [10 empty] | 0 | 1.7806 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.53% | 5.1762 | 2.8216 | -13.1988 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.5% | 3.0466 | 8.7746 | -0.92513 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_104 | alfaA |  |  | 0.00024284 | 5.3781e-05 | 8.9828e-05 | 1.0634e-05 | 7.8373e-05 | 1.005e-06 | 7.0329e-05 | 5.5676e-06 | -6.8277e-06 | [5 empty] | 20.4316 | 4.5408 | 7.0815 | 4.1875e-05 | 7.4605 | 2.5142e-05 | 1.2456 | 1.4459e-05 | -2.7787e-07 |  |  |  | 1 | 0 | 0 | 20.4313 | 4.5407 | 7.0814 | 3.124e-05 | 7.4604 | 2.4137e-05 | 1.2456 | 8.8915e-06 | 6.5498e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7672 | 20.4313 | 4.5407 | 7.0814 | 7.4604 | 1.2456 | 0.22224 | 0.34659 | 0.36515 | 0.060963 | 1.5595 | 1.643 | 0.27431 | 1.0535 | 0.17589 | 0.16695 | -1.504 | -1.0596 | -1.0075 | -2.7975 | [10 empty] | 0 | 1.7672 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014078 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 4 | 107% | -6.2328 | -1.2622 | 22.5365 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.41% | 0.75001 | -3.6706 | 16.7153 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_108 | iRM_12 |  |  | 0.00026809 | 5.4984e-05 | 9.0311e-05 | -1.0213e-06 | 8.2527e-05 | -3.0599e-07 | 6.2594e-05 | 5.643e-06 | -4.9777e-06 | [5 empty] | 19.1409 | 4.2546 | 6.6361 | 4.7244e-05 | 6.9934 | 2.21e-05 | 1.1679 | 1.4673e-05 | 1.9697e-07 |  |  |  | 1 | 0 | 0 | 19.1407 | 4.2546 | 6.636 | 4.8265e-05 | 6.9933 | 2.2406e-05 | 1.1679 | 9.0297e-06 | 5.1747e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7807 | 19.1407 | 4.2546 | 6.636 | 6.9933 | 1.1679 | 0.22228 | 0.3467 | 0.36536 | 0.061015 | 1.5597 | 1.6437 | 0.2745 | 1.0538 | 0.17599 | 0.167 | -1.5038 | -1.0593 | -1.0069 | -2.7966 | [10 empty] | 0 | 1.7807 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014079 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.7% | -3.6455 | -5.2016 | 2.3958 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.55% | -1.1452 | -2.3228 | -10.902 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_112 | Ames |  |  | 0.0002819 | 6.7112e-05 | 9.3902e-05 | 1.3948e-05 | 8.8136e-05 | 1.5371e-06 | 6.8174e-05 | 7.9558e-06 | -1.9359e-06 | [5 empty] | 21.0358 | 4.6763 | 7.2944 | 3.8956e-05 | 7.6885 | 3.0479e-05 | 1.2842 | 2.1368e-05 | 8.3326e-06 |  |  |  | 1 | 0 | 0 | 21.0355 | 4.6762 | 7.2943 | 2.5008e-05 | 7.6884 | 2.8941e-05 | 1.2842 | 1.3412e-05 | 1.0269e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7891 | 21.0355 | 4.6762 | 7.2943 | 7.6884 | 1.2841 | 0.2223 | 0.34676 | 0.3655 | 0.061047 | 1.5599 | 1.6441 | 0.27461 | 1.054 | 0.17605 | 0.16702 | -1.5037 | -1.0591 | -1.0065 | -2.7961 | [10 empty] | 0 | 1.7891 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43912 | -1.3882 | 0.014069 | -1.8133 | -1.8274 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 4 | 110.4% | 7.9553 | 6.9644 | -22.63 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 101% | -1.5363 | -0.39306 | 7.888 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_116 | iRM_12 |  |  | 0.00025101 | 5.4594e-05 | 7.7316e-05 | 7.5394e-06 | 7.4158e-05 | 9.3461e-07 | 5.0988e-05 | 1.7993e-05 | -2.105e-06 | [5 empty] | 18.9548 | 4.2133 | 6.5718 | 4.89e-05 | 6.9257 | 2.2373e-05 | 1.1566 | 1.3818e-05 | 2.315e-06 |  |  |  | 1 | 0 | 0 | 18.9546 | 4.2132 | 6.5717 | 4.1361e-05 | 6.9256 | 2.1438e-05 | 1.1566 | -4.1756e-06 | 4.42e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7818 | 18.9546 | 4.2132 | 6.5717 | 6.9256 | 1.1566 | 0.22228 | 0.34671 | 0.36538 | 0.061019 | 1.5598 | 1.6438 | 0.27451 | 1.0539 | 0.176 | 0.167 | -1.5038 | -1.0593 | -1.0068 | -2.7966 | [10 empty] | 0 | 1.7818 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014073 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.08% | 3.964 | 4.0478 | 2.8033 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 98.87% | 0.14536 | 5.1915 | 0.88912 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_120 | IPGP |  |  | 0.00024953 | 5.7085e-05 | 8.3152e-05 | 6.3566e-06 | 7.4796e-05 | -1.3513e-06 | 6.0902e-05 | 3.7466e-06 | -3.7741e-06 | [5 empty] | 20.6499 | 4.59 | 7.1591 | 3.3857e-05 | 7.5444 | 2.4625e-05 | 1.2599 | 1.7111e-05 | 4.6089e-07 |  |  |  | 1 | 0 | 0 | 20.6497 | 4.5899 | 7.1591 | 2.7501e-05 | 7.5443 | 2.5976e-05 | 1.2599 | 1.3364e-05 | 4.235e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7799 | 20.6497 | 4.5899 | 7.1591 | 7.5443 | 1.2599 | 0.22228 | 0.34669 | 0.36535 | 0.061012 | 1.5597 | 1.6437 | 0.27449 | 1.0538 | 0.17598 | 0.167 | -1.5038 | -1.0593 | -1.0069 | -2.7967 | [10 empty] | 0 | 1.7799 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 4 | 108.5% | -2.8363 | -2.4561 | -4.1916 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100% | 8.3987 | -4.2598 | -1.434 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_124 | iRM_12 |  |  | 0.00027155 | 6.2531e-05 | 8.7153e-05 | 1.2906e-05 | 8.4323e-05 | -3.7335e-06 | 5.4536e-05 | 1.1773e-05 | -2.1785e-06 | [5 empty] | 19.1227 | 4.2506 | 6.6301 | 4.9636e-05 | 6.9874 | 2.1403e-05 | 1.167 | 1.4981e-05 | -2.1208e-06 |  |  |  | 1 | 0 | 0 | 19.1224 | 4.2506 | 6.63 | 3.6729e-05 | 6.9873 | 2.5137e-05 | 1.1669 | 3.2075e-06 | 5.771e-08 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7832 | 19.1224 | 4.2506 | 6.63 | 6.9873 | 1.1669 | 0.22228 | 0.34672 | 0.3654 | 0.061025 | 1.5598 | 1.6439 | 0.27454 | 1.0539 | 0.17601 | 0.16701 | -1.5038 | -1.0592 | -1.0068 | -2.7965 | [10 empty] | 0 | 1.7832 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 101.1% | -5.7801 | -2.035 | -4.7261 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 101.1% | -14.3616 | -0.25284 | -14.8951 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_128 | iRM_UR |  |  | 0.00026085 | 5.6945e-05 | 8.3694e-05 | 2.8257e-05 | 7.2141e-05 | 1.6266e-06 | 6.1596e-05 | 2.6279e-05 | -2.3781e-06 | [5 empty] | 20.3305 | 4.5192 | 7.049 | 3.1247e-05 | 7.4288 | 2.1625e-05 | 1.2407 | 1.8019e-05 | -3.3032e-06 |  |  |  | 1 | 0 | 0 | 20.3303 | 4.5191 | 7.0489 | 2.9908e-06 | 7.4287 | 1.9998e-05 | 1.2406 | -8.2606e-06 | -9.2502e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7832 | 20.3303 | 4.5191 | 7.0489 | 7.4287 | 1.2406 | 0.22229 | 0.34672 | 0.3654 | 0.061023 | 1.5598 | 1.6438 | 0.27453 | 1.0539 | 0.176 | 0.167 | -1.5038 | -1.0592 | -1.0068 | -2.7965 | [10 empty] | 0 | 1.7832 | 0.21794 | 0.33338 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43912 | -1.3882 | 0.014065 | -1.8133 | -1.8274 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 4 | 107% | 7.8537 | 8.7149 | -23.4133 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.2% | 8.1632 | 4.0272 | 15.9972 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_132 | iRM_12 |  |  | 0.00027549 | 6.1013e-05 | 8.3082e-05 | 1.6655e-05 | 8.0878e-05 | 1.9663e-07 | 6.4725e-05 | 1.1651e-05 | -9.35e-06 | [5 empty] | 18.8773 | 4.1961 | 6.5451 | 5.0746e-05 | 6.8978 | 2.2426e-05 | 1.152 | 1.8551e-05 | -9.902e-07 |  |  |  | 1 | 0 | 0 | 18.877 | 4.1961 | 6.545 | 3.4091e-05 | 6.8977 | 2.223e-05 | 1.152 | 6.9004e-06 | 8.3598e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7833 | 18.877 | 4.1961 | 6.545 | 6.8977 | 1.152 | 0.22228 | 0.34672 | 0.3654 | 0.061025 | 1.5598 | 1.6439 | 0.27454 | 1.0539 | 0.17601 | 0.16701 | -1.5038 | -1.0592 | -1.0068 | -2.7965 | [10 empty] | 0 | 1.7833 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014073 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.1% | 4.1641 | 7.4771 | 12.2294 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.84% | -0.7186 | -5.4824 | -8.9986 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_136 | SRM |  |  | 0.00026789 | 4.624e-05 | 8.4039e-05 | 1.7223e-05 | 7.831e-05 | 7.8471e-07 | 6.431e-05 | 9.3644e-06 | -1.0784e-06 | [5 empty] | 20.5436 | 4.5665 | 7.1229 | 4.9189e-05 | 7.5067 | 2.3882e-05 | 1.2538 | 1.6485e-05 | -5.3666e-07 |  |  |  | 1 | 0 | 0 | 20.5433 | 4.5665 | 7.1228 | 3.1967e-05 | 7.5066 | 2.3097e-05 | 1.2537 | 7.1208e-06 | 5.4177e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7834 | 20.5433 | 4.5665 | 7.1228 | 7.5066 | 1.2537 | 0.22228 | 0.34672 | 0.3654 | 0.061026 | 1.5598 | 1.6439 | 0.27454 | 1.0539 | 0.17601 | 0.16701 | -1.5038 | -1.0592 | -1.0067 | -2.7964 | [10 empty] | 0 | 1.7834 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014072 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 4 | 108.5% | 1.403 | 7.1435 | 12.8115 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.7% | 0.04127 | 3.273 | 6.8818 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_140 | iRM_12 |  |  | 0.0019729 | 0.00041545 | 0.00068487 | 1.3567e-05 | 0.00069834 | -3.6797e-06 | 0.00017215 | 5.9282e-06 | -3.1475e-06 | [5 empty] | 18.9774 | 4.2184 | 6.58 | 5.0276e-05 | 6.9349 | 2.286e-05 | 1.1583 | 1.1417e-05 | -5.2011e-07 |  |  |  | 1 | 0 | 0 | 18.9755 | 4.218 | 6.5793 | 3.671e-05 | 6.9342 | 2.6539e-05 | 1.1581 | 5.4892e-06 | 2.6274e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.785 | 18.9755 | 4.218 | 6.5793 | 6.9342 | 1.1581 | 0.22229 | 0.34673 | 0.36543 | 0.061031 | 1.5598 | 1.644 | 0.27456 | 1.0539 | 0.17602 | 0.16701 | -1.5038 | -1.0592 | -1.0067 | -2.7964 | [10 empty] | 0 | 1.785 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014086 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.2% | -8.8991 | -11.6156 | -8.1548 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.77% | 0.77672 | 0.55293 | -1.6803 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_144 | alfaA |  |  | 0.00024869 | 5.5073e-05 | 8.4178e-05 | 1.7097e-05 | 7.5793e-05 | -1.4864e-06 | 6.8037e-05 | 8.813e-06 | -3.559e-06 | [5 empty] | 20.3797 | 4.5295 | 7.0642 | 3.3732e-05 | 7.443 | 2.3797e-05 | 1.2428 | 1.4029e-05 | -1.1625e-06 |  |  |  | 1 | 0 | 0 | 20.3795 | 4.5295 | 7.0641 | 1.6636e-05 | 7.443 | 2.5283e-05 | 1.2428 | 5.2165e-06 | 2.3965e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7717 | 20.3795 | 4.5295 | 7.0641 | 7.443 | 1.2428 | 0.22226 | 0.34663 | 0.36522 | 0.060981 | 1.5596 | 1.6432 | 0.27437 | 1.0536 | 0.17592 | 0.16697 | -1.5039 | -1.0595 | -1.0073 | -2.7972 | [10 empty] | 0 | 1.7717 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014075 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 4 | 107.3% | 0.14868 | 5.6375 | 26.0358 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.36% | 5.9631 | 2.5952 | 1.4463 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_148 | iRM_12 |  |  | 0.00026515 | 6.6025e-05 | 9.4147e-05 | 7.8545e-06 | 8.0701e-05 | -1.6235e-06 | 6.835e-05 | 3.9265e-06 | -5.918e-06 | [5 empty] | 18.9955 | 4.2225 | 6.5864 | 4.8734e-05 | 6.9416 | 2.1467e-05 | 1.1594 | 9.6647e-06 | -8.6679e-07 |  |  |  | 1 | 0 | 0 | 18.9952 | 4.2225 | 6.5863 | 4.0879e-05 | 6.9415 | 2.3091e-05 | 1.1593 | 5.7382e-06 | 5.0512e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7854 | 18.9952 | 4.2225 | 6.5863 | 6.9415 | 1.1593 | 0.22229 | 0.34673 | 0.36544 | 0.061033 | 1.5598 | 1.6439 | 0.27456 | 1.0539 | 0.17602 | 0.16701 | -1.5038 | -1.0592 | -1.0067 | -2.7963 | [10 empty] | 0 | 1.7854 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42504 | 0.43912 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.1% | 10.5562 | 6.1203 | -14.1507 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 101.3% | -14.68 | -11.1745 | 0.082121 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_152 | Ames |  |  | 0.00026045 | 5.2362e-05 | 9.4833e-05 | 1.7851e-06 | 7.9866e-05 | 3.1516e-06 | 6.6919e-05 | 1.2229e-05 | -1.8156e-06 | [5 empty] | 20.9067 | 4.6477 | 7.2502 | 3.7267e-05 | 7.6426 | 3.0571e-05 | 1.2767 | 1.3941e-05 | 6.943e-06 |  |  |  | 1 | 0 | 0 | 20.9064 | 4.6476 | 7.2501 | 3.5482e-05 | 7.6425 | 2.7419e-05 | 1.2766 | 1.7121e-06 | 8.7586e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.793 | 20.9064 | 4.6476 | 7.2501 | 7.6425 | 1.2766 | 0.22231 | 0.34679 | 0.36556 | 0.061063 | 1.56 | 1.6444 | 0.27468 | 1.0541 | 0.17608 | 0.16704 | -1.5037 | -1.059 | -1.0063 | -2.7959 | [10 empty] | 0 | 1.793 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014079 | -1.8133 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 4 | 110.2% | -11.801 | -2.9843 | -8.8294 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.11% | 4.0466 | 12.6723 | -21.0345 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_156 | iRM_12 |  |  | 0.00024916 | 5.9653e-05 | 9.5994e-05 | 9.9076e-06 | 7.1707e-05 | 2.434e-06 | 6.2645e-05 | 1.0055e-05 | -5.8197e-06 | [5 empty] | 18.964 | 4.2155 | 6.5756 | 4.5175e-05 | 6.9304 | 2.2286e-05 | 1.1576 | 1.1451e-05 | -5.4273e-07 |  |  |  | 1 | 0 | 0 | 18.9637 | 4.2155 | 6.5755 | 3.5267e-05 | 6.9303 | 1.9852e-05 | 1.1575 | 1.3963e-06 | 5.277e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7863 | 18.9637 | 4.2155 | 6.5755 | 6.9303 | 1.1575 | 0.22229 | 0.34674 | 0.36545 | 0.061038 | 1.5598 | 1.644 | 0.27458 | 1.054 | 0.17603 | 0.16702 | -1.5038 | -1.0592 | -1.0066 | -2.7963 | [10 empty] | 0 | 1.7863 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014077 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 99.8% | -4.5712 | -2.1478 | 14.9932 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100% | 7.1432 | -5.1927 | 36.9643 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_160 | IPGP |  |  | 0.00026612 | 6.1019e-05 | 9.2868e-05 | 8.4278e-06 | 8.2252e-05 | -2.2877e-06 | 6.0048e-05 | 8.8458e-06 | -2.9684e-06 | [5 empty] | 20.8336 | 4.631 | 7.2235 | 3.1931e-05 | 7.6129 | 2.5226e-05 | 1.2715 | 1.5403e-05 | -4.9412e-07 |  |  |  | 1 | 0 | 0 | 20.8334 | 4.631 | 7.2234 | 2.3503e-05 | 7.6128 | 2.7514e-05 | 1.2714 | 6.5577e-06 | 2.4742e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.784 | 20.8334 | 4.631 | 7.2234 | 7.6128 | 1.2714 | 0.22229 | 0.34672 | 0.36541 | 0.061028 | 1.5598 | 1.6439 | 0.27455 | 1.0539 | 0.17601 | 0.16701 | -1.5038 | -1.0592 | -1.0067 | -2.7964 | [10 empty] | 0 | 1.784 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014078 | -1.8133 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 4 | 109.7% | -1.807 | -2.441 | -14.4865 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.5% | -3.0949 | 1.9911 | -22.0627 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_164 | iRM_12 |  |  | 0.00026398 | 4.2481e-05 | 9.9426e-05 | 3.4143e-07 | 7.5529e-05 | -1.6486e-06 | 6.778e-05 | 6.1168e-06 | -3.4401e-06 | [5 empty] | 19.0084 | 4.2255 | 6.5912 | 4.674e-05 | 6.947 | 2.2048e-05 | 1.1604 | 1.1529e-05 | -1.3127e-06 |  |  |  | 1 | 0 | 0 | 19.0082 | 4.2254 | 6.5911 | 4.6399e-05 | 6.9469 | 2.3697e-05 | 1.1603 | 5.4125e-06 | 2.1274e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7875 | 19.0082 | 4.2254 | 6.5911 | 6.9469 | 1.1603 | 0.2223 | 0.34675 | 0.36547 | 0.061042 | 1.5599 | 1.6441 | 0.2746 | 1.054 | 0.17604 | 0.16702 | -1.5037 | -1.0592 | -1.0066 | -2.7962 | [10 empty] | 0 | 1.7875 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014075 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 4 | 100.7% | 0.97368 | 1.9467 | 6.4441 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.27% | 1.0287 | -1.2047 | 2.2395 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_168 | iRM_UR |  |  | 0.00024825 | 5.29e-05 | 9.2538e-05 | 5.9949e-07 | 8.8717e-05 | -8.2964e-07 | 5.2659e-05 | 2.0986e-05 | 1.0908e-06 | [5 empty] | 20.0562 | 4.4583 | 6.9543 | 2.7423e-05 | 7.3295 | 2.0507e-05 | 1.2242 | 1.0715e-05 | -1.3601e-06 |  |  |  | 1 | 0 | 0 | 20.056 | 4.4583 | 6.9542 | 2.6823e-05 | 7.3295 | 2.1337e-05 | 1.2242 | -1.0271e-05 | -2.4509e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7863 | 20.056 | 4.4583 | 6.9542 | 7.3295 | 1.2242 | 0.22229 | 0.34674 | 0.36545 | 0.061037 | 1.5598 | 1.644 | 0.27458 | 1.054 | 0.17603 | 0.16702 | -1.5038 | -1.0592 | -1.0066 | -2.7963 | [10 empty] | 0 | 1.7863 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.01408 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 4 | 106.1% | -6.4654 | -4.6174 | 0.44618 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.3% | -2.5245 | -2.4203 | 2.7743 | iRM_12
Lab 1 | Jun 22, 2022 | 4 | Zr_172 | iRM_12 |  |  | 0.00028049 | 5.0391e-05 | 9.0238e-05 | 2.9589e-06 | 8.4087e-05 | -4.9528e-06 | 7.1348e-05 | 3.0909e-06 | -4.0441e-06 | [5 empty] | 18.7976 | 4.1786 | 6.5179 | 3.8676e-05 | 6.8697 | 2.1121e-05 | 1.1474 | 1.7087e-05 | 3.8983e-07 |  |  |  | 1 | 0 | 0 | 18.7973 | 4.1785 | 6.5178 | 3.5717e-05 | 6.8696 | 2.6074e-05 | 1.1474 | 1.3996e-05 | 4.434e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7868 | 18.7973 | 4.1785 | 6.5178 | 6.8696 | 1.1474 | 0.22229 | 0.34674 | 0.36546 | 0.061039 | 1.5598 | 1.644 | 0.27459 | 1.054 | 0.17603 | 0.16702 | -1.5038 | -1.0592 | -1.0066 | -2.7962 | [10 empty] | 0 | 1.7868 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5513 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43912 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [22 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.6% | 3.8096 | 4.5273 | -9.4701 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_004 | iRM_12 |  |  | 0.00015558 | 2.6332e-05 | 5.9052e-05 | -2.1738e-08 | 5.5601e-05 | 4.1974e-06 | 4.6804e-05 | 7.8827e-06 | -1.6888e-06 | [5 empty] | 20.7433 | 4.6112 | 7.1931 | 3.9224e-05 | 7.5819 | 2.3159e-05 | 1.2665 | 1.2614e-05 | -3.4695e-06 |  |  |  | 1 | 0 | 0 | 20.7432 | 4.6112 | 7.193 | 3.9246e-05 | 7.5818 | 1.8962e-05 | 1.2664 | 4.7311e-06 | -1.7807e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7899 | 20.7432 | 4.6112 | 7.193 | 7.5818 | 1.2664 | 0.2223 | 0.34676 | 0.36551 | 0.061052 | 1.5599 | 1.6442 | 0.27464 | 1.0541 | 0.17606 | 0.16703 | -1.5037 | -1.0591 | -1.0065 | -2.796 | [9 empty] | L1 JunS5 | 0 | 1.7899 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014083 | -1.8132 | -1.8273 | [22 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 98.57% | -1.5988 | -3.5245 | 20.7172 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_008 | iRM_UR |  |  | 0.0037197 | 0.0008198 | 0.001284 | 6.534e-08 | 0.0013467 | 1.1412e-07 | 0.00027523 | 7.5977e-06 | -1.8679e-06 | [5 empty] | 20.423 | 4.54 | 7.0819 | 2.5713e-05 | 7.4644 | 2.0904e-05 | 1.2468 | 8.4381e-06 | -4.4534e-06 |  |  |  | 1 | 0 | 0 | 20.4193 | 4.5392 | 7.0806 | 2.5648e-05 | 7.4631 | 2.079e-05 | 1.2466 | 8.4044e-07 | -2.5855e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7889 | 20.4193 | 4.5392 | 7.0806 | 7.4631 | 1.2466 | 0.2223 | 0.34676 | 0.36549 | 0.061048 | 1.5599 | 1.6441 | 0.27462 | 1.054 | 0.17605 | 0.16703 | -1.5037 | -1.0591 | -1.0065 | -2.7961 | [10 empty] | 0 | 1.7889 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014078 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 5 | 98.71% | 10.1005 | 0.73033 | 20.5562 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.7% | -4.1566 | -1.2293 | -22.5064 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_012 | iRM_12 |  |  | 0.00025565 | 6.3025e-05 | 8.5144e-05 | -1.2434e-06 | 6.8486e-05 | -1.6045e-06 | 7.3252e-05 | 4.083e-06 | -1.9489e-06 | [5 empty] | 20.6292 | 4.5859 | 7.1538 | 4.6343e-05 | 7.5406 | 2.2739e-05 | 1.2596 | 5.9191e-06 | -8.1344e-07 |  |  |  | 1 | 0 | 0 | 20.629 | 4.5859 | 7.1537 | 4.7587e-05 | 7.5406 | 2.4344e-05 | 1.2595 | 1.8361e-06 | 1.1354e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7915 | 20.629 | 4.5859 | 7.1537 | 7.5406 | 1.2595 | 0.2223 | 0.34678 | 0.36553 | 0.061056 | 1.56 | 1.6443 | 0.27465 | 1.0541 | 0.17607 | 0.16703 | -1.5037 | -1.0591 | -1.0064 | -2.796 | [10 empty] | 0 | 1.7915 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9118 | 0.42507 | 0.43914 | -1.3882 | 0.014074 | -1.8133 | -1.8274 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | -10.4279 | 5.0601 | -39.5082 | [8 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.5% | 4.7905 | 3.3199 | 2.7836 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_016 | SRM |  |  | 0.0008184 | 0.00018362 | 0.00029492 | 1.1853e-05 | 0.0002903 | 3.4962e-07 | 0.0001028 | -5.0511e-07 | -1.0724e-06 | [5 empty] | 20.6054 | 4.5806 | 7.1452 | 4.9963e-05 | 7.5314 | 2.5063e-05 | 1.258 | 3.8362e-06 | 3.0493e-07 |  |  |  | 1 | 0 | 0 | 20.6046 | 4.5804 | 7.1449 | 3.811e-05 | 7.5311 | 2.4713e-05 | 1.2579 | 4.3413e-06 | 1.3774e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7898 | 20.6046 | 4.5804 | 7.1449 | 7.5311 | 1.2579 | 0.2223 | 0.34676 | 0.36551 | 0.061051 | 1.5599 | 1.6442 | 0.27464 | 1.054 | 0.17606 | 0.16703 | -1.5037 | -1.0591 | -1.0065 | -2.796 | [10 empty] | 0 | 1.7898 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014081 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 100.5% | -0.22694 | -6.2667 | 4.6452 | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.39% | 0.00084539 | 0.13301 | 21.1027 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_020 | iRM_12 |  |  | 0.00026646 | 5.2052e-05 | 8.0622e-05 | 2.4275e-06 | 7.8762e-05 | -1.4911e-06 | 6.4313e-05 | 1.532e-05 | -4.5929e-06 | [5 empty] | 20.3645 | 4.5271 | 7.062 | 4.0665e-05 | 7.4439 | 2.4841e-05 | 1.2435 | 7.255e-06 | -9.3879e-07 |  |  |  | 1 | 0 | 0 | 20.3642 | 4.5271 | 7.0619 | 3.8237e-05 | 7.4438 | 2.6332e-05 | 1.2434 | -8.0655e-06 | 3.6541e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7914 | 20.3642 | 4.5271 | 7.0619 | 7.4438 | 1.2434 | 0.2223 | 0.34678 | 0.36553 | 0.061059 | 1.5599 | 1.6443 | 0.27466 | 1.0541 | 0.17607 | 0.16704 | -1.5037 | -1.0591 | -1.0064 | -2.7959 | [10 empty] | 0 | 1.7914 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014076 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.57% | 6.9685 | -0.088558 | 31.3635
Lab 1 | Jun 22, 2022 | 5 | Zr_024 | alfaA |  |  | 0.0014003 | 0.00029966 | 0.00048365 | 4.7449e-06 | 0.00049733 | 1.2514e-06 | 0.00013399 | 5.2412e-06 | -5.1363e-06 | [5 empty] | 20.8315 | 4.6303 | 7.222 | 3.4028e-05 | 7.6106 | 2.3257e-05 | 1.271 | 1.6928e-05 | 7.8385e-07 |  |  |  | 1 | 0 | 0 | 20.8301 | 4.63 | 7.2215 | 2.9284e-05 | 7.6101 | 2.2006e-05 | 1.2709 | 1.1687e-05 | 5.9201e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7795 | 20.8301 | 4.63 | 7.2215 | 7.6101 | 1.2709 | 0.22227 | 0.34669 | 0.36534 | 0.061011 | 1.5597 | 1.6437 | 0.27448 | 1.0538 | 0.17598 | 0.167 | -1.5038 | -1.0593 | -1.0069 | -2.7967 | [10 empty] | 0 | 1.7795 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.01408 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 5 | 102.5% | -0.17264 | -3.9461 | -12.4797 | SRM | Lab 1 | Jun 21, 2022 | 3 | 101.9% | 13.4969 | 0.93421 | -14.066 | iRM_12 | Avg | 1.8116 | -0.38309 | 1.5634
Lab 1 | Jun 22, 2022 | 5 | Zr_028 | iRM_12 |  |  | 0.0002656 | 4.0121e-05 | 8.4649e-05 | 1.0547e-05 | 7.3491e-05 | -3.1351e-07 | 6.1464e-05 | 1.402e-05 | -1.2631e-07 | [5 empty] | 20.2753 | 4.5073 | 7.0312 | 4.2729e-05 | 7.4116 | 2.2319e-05 | 1.2381 | 1.3964e-05 | -8.7378e-07 |  |  |  | 1 | 0 | 0 | 20.275 | 4.5073 | 7.0311 | 3.2182e-05 | 7.4115 | 2.2633e-05 | 1.238 | -5.5414e-08 | -7.4746e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7925 | 20.275 | 4.5073 | 7.0311 | 7.4115 | 1.238 | 0.22231 | 0.34679 | 0.36555 | 0.061062 | 1.5599 | 1.6443 | 0.27468 | 1.0541 | 0.17608 | 0.16704 | -1.5037 | -1.059 | -1.0064 | -2.7959 | [10 empty] | 0 | 1.7925 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014077 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.46% | 0.27098 | -1.3111 | 6.4189 | [8 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 103.2% | -1.625 | -6.0986 | -18.6138 | iRM_12 | 2SD | 12.9638 | 9.7927 | 22.8484
Lab 1 | Jun 22, 2022 | 5 | Zr_032 | Ames |  |  | 0.0023853 | 0.00051458 | 0.00082346 | -1.0884e-07 | 0.00085291 | -1.0619e-06 | 0.00019667 | 1.1881e-05 | 1.6286e-06 | [5 empty] | 21.0593 | 4.6819 | 7.304 | 3.0366e-05 | 7.7002 | 2.9878e-05 | 1.2865 | 1.7987e-05 | 8.5524e-06 |  |  |  | 1 | 0 | 0 | 21.0569 | 4.6814 | 7.3032 | 3.0475e-05 | 7.6993 | 3.0939e-05 | 1.2863 | 6.1055e-06 | 6.9238e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7984 | 21.0569 | 4.6814 | 7.3032 | 7.6993 | 1.2863 | 0.22232 | 0.34683 | 0.36564 | 0.061085 | 1.56 | 1.6447 | 0.27476 | 1.0542 | 0.17612 | 0.16706 | -1.5036 | -1.0589 | -1.0061 | -2.7955 | [10 empty] | 0 | 1.7984 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014076 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 5 | 103.5% | 3.9799 | -0.19577 | 4.0311 | SRM | Lab 1 | Jun 21, 2022 | 3 | 102.5% | -0.2557 | -2.8863 | -18.6487 | iRM_12 | n | 21 | 21 | 21
Lab 1 | Jun 22, 2022 | 5 | Zr_036 | iRM_12 |  |  | 0.00025038 | 5.9043e-05 | 8.0469e-05 | 1.5301e-06 | 7.4921e-05 | 6.0434e-07 | 6.48e-05 | 1.0208e-05 | -5.0761e-06 | [5 empty] | 20.4059 | 4.5364 | 7.0767 | 3.8815e-05 | 7.4599 | 2.373e-05 | 1.2462 | 1.3898e-05 | -2.9179e-07 |  |  |  | 1 | 0 | 0 | 20.4056 | 4.5364 | 7.0767 | 3.7285e-05 | 7.4598 | 2.3125e-05 | 1.2461 | 3.6906e-06 | 4.7843e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7941 | 20.4056 | 4.5364 | 7.0767 | 7.4598 | 1.2461 | 0.22231 | 0.3468 | 0.36557 | 0.061067 | 1.56 | 1.6444 | 0.27469 | 1.0541 | 0.17609 | 0.16704 | -1.5037 | -1.059 | -1.0063 | -2.7958 | [10 empty] | 0 | 1.7941 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42506 | 0.43913 | -1.3882 | 0.014075 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.5% | -2.9043 | 4.2256 | -20.8563 | [8 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 102.3% | 1.3769 | 1.5391 | 12.3557 | iRM_12 | 2SE | 2.8289 | 2.1369 | 4.9859
Lab 1 | Jun 22, 2022 | 5 | Zr_040 | IPGP |  |  | 0.00023989 | 4.2198e-05 | 7.7652e-05 | 7.5034e-06 | 7.2308e-05 | 5.8958e-07 | 7.4215e-05 | 5.0735e-06 | -9.3649e-07 | [5 empty] | 21.0114 | 4.6709 | 7.2863 | 3.0098e-05 | 7.6803 | 2.4076e-05 | 1.2829 | 1.9185e-05 | 1.6907e-06 |  |  |  | 1 | 0 | 0 | 21.0112 | 4.6709 | 7.2862 | 2.2595e-05 | 7.6802 | 2.3486e-05 | 1.2829 | 1.4112e-05 | 2.6272e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7913 | 21.0112 | 4.6709 | 7.2862 | 7.6802 | 1.2829 | 0.2223 | 0.34678 | 0.36553 | 0.061057 | 1.5599 | 1.6443 | 0.27465 | 1.0541 | 0.17607 | 0.16704 | -1.5037 | -1.0591 | -1.0064 | -2.796 | [10 empty] | 0 | 1.7913 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.01408 | -1.8133 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 5 | 103.2% | 2.4189 | -1.6931 | -1.9162 | SRM | Lab 1 | Jun 21, 2022 | 3 | 103.9% | 3.9051 | -2.9282 | 4.7037 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_044 | iRM_12 |  |  | 0.00029336 | 5.4338e-05 | 9.6595e-05 | 1.4192e-06 | 9.2722e-05 | 7.468e-07 | 6.2681e-05 | 7.109e-06 | -1.7508e-06 | [5 empty] | 20.3169 | 4.5166 | 7.0457 | 3.9773e-05 | 7.4271 | 2.2387e-05 | 1.2407 | 1.7997e-05 | -1.5064e-06 |  |  |  | 1 | 0 | 0 | 20.3166 | 4.5166 | 7.0456 | 3.8354e-05 | 7.427 | 2.164e-05 | 1.2406 | 1.0888e-05 | 2.4437e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7934 | 20.3166 | 4.5166 | 7.0456 | 7.427 | 1.2406 | 0.22231 | 0.34679 | 0.36556 | 0.061066 | 1.56 | 1.6444 | 0.27469 | 1.0541 | 0.17609 | 0.16704 | -1.5037 | -1.059 | -1.0063 | -2.7958 | [10 empty] | 0 | 1.7934 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014082 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.2% | 3.8579 | -5.1702 | -5.2927 | [8 empty] | SRM | Lab 1 | Jun 21, 2022 | 3 | 105.5% | 3.8216 | -3.7751 | 22.1838 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_048 | iRM_UR |  |  | 0.0017302 | 0.00038055 | 0.0006086 | 1.523e-07 | 0.00061407 | 7.6098e-07 | 0.00015249 | 1.1278e-05 | -4.5056e-06 | [5 empty] | 20.1615 | 4.4821 | 6.992 | 3.0168e-05 | 7.3705 | 2.082e-05 | 1.2313 | 9.5608e-06 | -1.6486e-06 |  |  |  | 1 | 0 | 0 | 20.1598 | 4.4817 | 6.9914 | 3.0016e-05 | 7.3699 | 2.0059e-05 | 1.2311 | -1.7177e-06 | 2.857e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.794 | 20.1598 | 4.4817 | 6.9914 | 7.3699 | 1.2311 | 0.22231 | 0.3468 | 0.36557 | 0.061067 | 1.56 | 1.6444 | 0.27469 | 1.0541 | 0.17609 | 0.16705 | -1.5037 | -1.059 | -1.0063 | -2.7958 | [10 empty] | 0 | 1.794 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 5 | 99.61% | -1.1521 | 4.7513 | -19.4476 | SRM | Lab 1 | Jun 22, 2022 | 4 | 106.5% | 9.7985 | 8.0267 | -4.1873 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_052 | iRM_12 |  |  | 0.00042877 | 9.6677e-05 | 0.00014226 | 1.0679e-06 | 0.00014681 | 8.2037e-07 | 7.1848e-05 | 2.006e-05 | -3.9189e-07 | [5 empty] | 20.16 | 4.4818 | 6.9914 | 4.0896e-05 | 7.37 | 2.234e-05 | 1.2312 | 1.4556e-05 | -1.7753e-06 |  |  |  | 1 | 0 | 0 | 20.1596 | 4.4817 | 6.9913 | 3.9828e-05 | 7.3698 | 2.152e-05 | 1.2311 | -5.5033e-06 | -1.3834e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7941 | 20.1596 | 4.4817 | 6.9913 | 7.3698 | 1.2311 | 0.22231 | 0.3468 | 0.36557 | 0.06107 | 1.56 | 1.6444 | 0.27471 | 1.0541 | 0.1761 | 0.16705 | -1.5037 | -1.059 | -1.0063 | -2.7957 | [10 empty] | 0 | 1.7941 | 0.21794 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014079 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.4% | -2.9433 | 3.9138 | 23.3935 | [8 empty] | SRM | Lab 1 | Jun 22, 2022 | 4 | 106.1% | -1.9784 | -7.1062 | 4.5582 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_056 | SRM |  |  | 0.00030968 | 3.2002e-05 | 0.00010775 | 6.5848e-06 | 9.9406e-05 | 1.3908e-07 | 6.2229e-05 | 3.9516e-06 | -4.2474e-06 | [5 empty] | 20.7151 | 4.6051 | 7.1838 | 4.4588e-05 | 7.5725 | 2.428e-05 | 1.265 | 1.0418e-05 | 2.3282e-06 |  |  |  | 1 | 0 | 0 | 20.7148 | 4.6051 | 7.1836 | 3.8003e-05 | 7.5724 | 2.4141e-05 | 1.2649 | 6.4669e-06 | 6.5755e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7928 | 20.7148 | 4.6051 | 7.1836 | 7.5724 | 1.2649 | 0.22231 | 0.34679 | 0.36555 | 0.061064 | 1.5599 | 1.6444 | 0.27468 | 1.0541 | 0.17608 | 0.16704 | -1.5037 | -1.059 | -1.0063 | -2.7958 | [10 empty] | 0 | 1.7928 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5513 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43912 | -1.3882 | 0.014079 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 102.5% | 6.1165 | 1.3364 | -8.3884 | SRM | Lab 1 | Jun 22, 2022 | 4 | 108.7% | -9.2237 | -1.3602 | -4.1285 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_060 | iRM_12 |  |  | 0.00028962 | 4.6817e-05 | 9.38e-05 | 3.5622e-06 | 8.9181e-05 | 1.207e-07 | 5.8912e-05 | 1.1362e-05 | -3.049e-06 | [5 empty] | 20.2476 | 4.5014 | 7.0222 | 4.0705e-05 | 7.4027 | 2.1239e-05 | 1.2367 | 1.622e-05 | -1.7777e-06 |  |  |  | 1 | 0 | 0 | 20.2473 | 4.5013 | 7.0221 | 3.7143e-05 | 7.4026 | 2.1118e-05 | 1.2367 | 4.8579e-06 | 1.2713e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7964 | 20.2473 | 4.5013 | 7.0221 | 7.4026 | 1.2367 | 0.22232 | 0.34681 | 0.36561 | 0.061078 | 1.56 | 1.6446 | 0.27473 | 1.0542 | 0.17611 | 0.16706 | -1.5037 | -1.059 | -1.0062 | -2.7956 | [10 empty] | 0 | 1.7964 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014083 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.2% | 1.3433 | -6.4545 | -13.1573 | [8 empty] | SRM | Lab 1 | Jun 22, 2022 | 4 | 108.5% | 1.403 | 7.1435 | 12.8115 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_064 | alfaA |  |  | 0.0021225 | 0.00047657 | 0.00073406 | 6.0985e-06 | 0.00076399 | 9.7465e-07 | 0.00018341 | 1.1389e-05 | 3.1387e-07 | [5 empty] | 20.6254 | 4.5846 | 7.1509 | 3.3902e-05 | 7.5361 | 2.2989e-05 | 1.2586 | 1.6927e-05 | -8.6796e-07 |  |  |  | 1 | 0 | 0 | 20.6233 | 4.5841 | 7.1502 | 2.7804e-05 | 7.5353 | 2.2014e-05 | 1.2584 | 5.5377e-06 | -1.1818e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7818 | 20.6233 | 4.5841 | 7.1502 | 7.5353 | 1.2584 | 0.22228 | 0.3467 | 0.36538 | 0.06102 | 1.5598 | 1.6438 | 0.27452 | 1.0539 | 0.176 | 0.16701 | -1.5038 | -1.0593 | -1.0068 | -2.7965 | [10 empty] | 0 | 1.7818 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014077 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 5 | 101.8% | -9.9331 | 1.8304 | 2.1322 | SRM | Lab 1 | Jun 22, 2022 | 5 | 100.5% | -0.22694 | -6.2667 | 4.6452 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_068 | iRM_12 |  |  | 0.00026467 | 4.2689e-05 | 8.0286e-05 | -5.9079e-06 | 6.9681e-05 | 2.1858e-06 | 7.3294e-05 | 9.5058e-06 | 3.9118e-06 | [5 empty] | 20.2717 | 4.5067 | 7.0305 | 4.0025e-05 | 7.4114 | 2.2761e-05 | 1.2382 | 1.0365e-05 | -1.0877e-06 |  |  |  | 1 | 0 | 0 | 20.2714 | 4.5067 | 7.0304 | 4.5933e-05 | 7.4113 | 2.0575e-05 | 1.2381 | 8.5884e-07 | -4.9996e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.796 | 20.2714 | 4.5067 | 7.0304 | 7.4113 | 1.2381 | 0.22232 | 0.34681 | 0.36561 | 0.061077 | 1.56 | 1.6445 | 0.27473 | 1.0542 | 0.17611 | 0.16706 | -1.5037 | -1.059 | -1.0062 | -2.7956 | [10 empty] | 0 | 1.796 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014074 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | 3.9851 | 4.2598 | 7.4186 | [8 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 102.5% | 6.1165 | 1.3364 | -8.3884 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_072 | Ames |  |  | 0.00084425 | 0.00018117 | 0.00028872 | -4.6896e-07 | 0.0002931 | -2.0655e-06 | 0.0001048 | 1.7925e-05 | 1.2102e-06 | [5 empty] | 21.1892 | 4.7111 | 7.3499 | 3.447e-05 | 7.7494 | 2.9682e-05 | 1.2948 | 1.8258e-05 | 9.3172e-06 |  |  |  | 1 | 0 | 0 | 21.1884 | 4.7109 | 7.3496 | 3.4939e-05 | 7.7491 | 3.1748e-05 | 1.2947 | 3.3375e-07 | 8.107e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8034 | 21.1884 | 4.7109 | 7.3496 | 7.7491 | 1.2947 | 0.22233 | 0.34687 | 0.36572 | 0.061104 | 1.5601 | 1.6449 | 0.27483 | 1.0544 | 0.17616 | 0.16708 | -1.5036 | -1.0588 | -1.0059 | -2.7952 | [10 empty] | 0 | 1.8034 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42505 | 0.43913 | -1.3882 | 0.014076 | -1.8133 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 5 | 104.7% | 4.4483 | -1.6179 | -20.9423 | SRM | Lab 1 | Jun 22, 2022 | 5 | 102.9% | -0.94789 | -7.6022 | 2.8503 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_076 | iRM_12 |  |  | 0.0030874 | 0.00067057 | 0.0010619 | 1.4977e-05 | 0.0011079 | -1.4372e-06 | 0.0002465 | 1.2492e-05 | -2.1002e-06 | [5 empty] | 20.1973 | 4.4901 | 7.0047 | 4.1182e-05 | 7.3843 | 2.2504e-05 | 1.2336 | 1.4742e-05 | -1.6221e-06 |  |  |  | 1 | 0 | 0 | 20.1942 | 4.4895 | 7.0037 | 2.6205e-05 | 7.3832 | 2.3942e-05 | 1.2334 | 2.2496e-06 | 4.781e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7963 | 20.1942 | 4.4895 | 7.0037 | 7.3832 | 1.2334 | 0.22231 | 0.34682 | 0.36561 | 0.061077 | 1.56 | 1.6446 | 0.27473 | 1.0542 | 0.17611 | 0.16706 | -1.5037 | -1.059 | -1.0062 | -2.7956 | [10 empty] | 0 | 1.7963 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014074 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.78% | -5.622 | 4.4603 | -3.1619 | [8 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 102.9% | 8.8925 | 3.8862 | 6.1875 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_080 | IPGP |  |  | 0.00034911 | 5.4203e-05 | 0.00012141 | 2.3525e-06 | 0.00012079 | -6.5046e-07 | 6.8895e-05 | 1.2325e-05 | -5.6738e-06 | [5 empty] | 21.0831 | 4.6869 | 7.3116 | 2.9764e-05 | 7.7075 | 2.5602e-05 | 1.2876 | 1.757e-05 | 1.6241e-06 |  |  |  | 1 | 0 | 0 | 21.0827 | 4.6869 | 7.3115 | 2.7411e-05 | 7.7074 | 2.6252e-05 | 1.2875 | 5.2455e-06 | 7.2979e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7944 | 21.0827 | 4.6869 | 7.3115 | 7.7074 | 1.2875 | 0.22231 | 0.3468 | 0.36558 | 0.06107 | 1.56 | 1.6445 | 0.27471 | 1.0542 | 0.1761 | 0.16705 | -1.5037 | -1.059 | -1.0063 | -2.7957 | [10 empty] | 0 | 1.7944 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.01408 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 5 | 104.4% | -6.9346 | -1.4783 | 9.5049 | SRM | Lab 1 | Jun 22, 2022 | 5 | 103.9% | 13.1982 | -3.1513 | 1.034 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_084 | iRM_12 |  |  | 0.0016631 | 0.0003643 | 0.00056873 | -5.7578e-06 | 0.00058561 | -1.4931e-06 | 0.00014251 | 5.5765e-06 | -4.8681e-06 | [5 empty] | 20.2089 | 4.4928 | 7.009 | 4.2565e-05 | 7.3892 | 2.2662e-05 | 1.2345 | 1.5388e-05 | -9.2396e-07 |  |  |  | 1 | 0 | 0 | 20.2072 | 4.4925 | 7.0085 | 4.8323e-05 | 7.3886 | 2.4155e-05 | 1.2344 | 9.8111e-06 | 3.9442e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7984 | 20.2072 | 4.4925 | 7.0085 | 7.3886 | 1.2344 | 0.22232 | 0.34683 | 0.36564 | 0.061085 | 1.56 | 1.6447 | 0.27476 | 1.0542 | 0.17612 | 0.16706 | -1.5036 | -1.0589 | -1.0061 | -2.7955 | [10 empty] | 0 | 1.7984 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014083 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.1% | 6.6088 | -9.3927 | 3.0602 | [8 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 103.9% | 5.424 | 8.3171 | 20.3895 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_088 | iRM_UR |  |  | 0.00026216 | 5.1549e-05 | 8.8555e-05 | 5.3061e-06 | 7.7708e-05 | 1.1376e-06 | 6.2773e-05 | 1.7666e-05 | 7.7624e-07 | [5 empty] | 19.7447 | 4.3896 | 6.8478 | 2.7677e-05 | 7.219 | 2.0568e-05 | 1.206 | 1.5064e-05 | -7.5529e-07 |  |  |  | 1 | 0 | 0 | 19.7444 | 4.3895 | 6.8477 | 2.2371e-05 | 7.2189 | 1.9431e-05 | 1.206 | -2.6027e-06 | -1.5315e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7967 | 19.7444 | 4.3895 | 6.8477 | 7.2189 | 1.206 | 0.22232 | 0.34682 | 0.36562 | 0.061078 | 1.56 | 1.6446 | 0.27473 | 1.0542 | 0.17611 | 0.16706 | -1.5037 | -1.059 | -1.0062 | -2.7956 | [10 empty] | 0 | 1.7967 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014078 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 5 | 97.73% | 10.0747 | 0.22552 | -3.0958 | SRM | Lab 1 | Jun 23, 2022 | 6 | 106.7% | -3.725 | 5.811 | -12.4275 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_092 | iRM_12 |  |  | 0.00098866 | 0.00021159 | 0.0003355 | -6.5893e-06 | 0.0003423 | -1.5118e-06 | 0.0001022 | 1.8161e-05 | -1.3663e-06 | [5 empty] | 20.2002 | 4.4909 | 7.0061 | 4.1687e-05 | 7.386 | 2.177e-05 | 1.234 | 1.8169e-05 | -3.0745e-07 |  |  |  | 1 | 0 | 0 | 20.1992 | 4.4906 | 7.0057 | 4.8276e-05 | 7.3857 | 2.3282e-05 | 1.2339 | 7.382e-09 | 1.0589e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7984 | 20.1992 | 4.4906 | 7.0057 | 7.3857 | 1.2339 | 0.22232 | 0.34683 | 0.36564 | 0.061085 | 1.5601 | 1.6447 | 0.27476 | 1.0542 | 0.17612 | 0.16706 | -1.5036 | -1.0589 | -1.0061 | -2.7955 | [10 empty] | 0 | 1.7984 | 0.21793 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43915 | -1.3882 | 0.014073 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.6% | -13.2211 | 7.3686 | -8.8603 | [8 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 108.4% | 5.8243 | 1.1404 | 4.2503 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_096 | SRM |  |  | 0.00024332 | 5.7611e-05 | 8.7374e-05 | 3.2591e-06 | 6.0935e-05 | -2.3099e-06 | 6.4301e-05 | 7.1152e-06 | -3.2512e-06 | [5 empty] | 20.6629 | 4.5937 | 7.1664 | 4.6475e-05 | 7.555 | 2.6505e-05 | 1.2622 | 7.6094e-06 | -5.8001e-07 |  |  |  | 1 | 0 | 0 | 20.6627 | 4.5937 | 7.1663 | 4.3216e-05 | 7.5549 | 2.8814e-05 | 1.2621 | 4.942e-07 | 2.6712e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7978 | 20.6627 | 4.5937 | 7.1663 | 7.5549 | 1.2621 | 0.22232 | 0.34682 | 0.36563 | 0.061083 | 1.56 | 1.6446 | 0.27475 | 1.0542 | 0.17612 | 0.16706 | -1.5036 | -1.0589 | -1.0061 | -2.7955 | [10 empty] | 0 | 1.7978 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014084 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 102.9% | -0.94789 | -7.6022 | 2.8503 | SRM | Lab 1 | Jun 23, 2022 | 6 | 109.1% | -1.2537 | -0.027791 | 11.7485 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_100 | iRM_12 |  |  | 0.00025587 | 3.9668e-05 | 8.4976e-05 | 8.0195e-06 | 7.8495e-05 | 2.5861e-06 | 6.4713e-05 | 6.7647e-06 | -2.0676e-06 | [5 empty] | 19.9502 | 4.4354 | 6.9195 | 4.0699e-05 | 7.2949 | 2.2349e-05 | 1.2188 | 1.8319e-05 | -2.3695e-07 |  |  |  | 1 | 0 | 0 | 19.95 | 4.4353 | 6.9194 | 3.268e-05 | 7.2949 | 1.9763e-05 | 1.2187 | 1.1555e-05 | 1.8307e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7994 | 19.95 | 4.4353 | 6.9194 | 7.2949 | 1.2187 | 0.22232 | 0.34684 | 0.36566 | 0.061089 | 1.5601 | 1.6447 | 0.27478 | 1.0543 | 0.17613 | 0.16707 | -1.5036 | -1.0589 | -1.0061 | -2.7954 | [10 empty] | 0 | 1.7994 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014078 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.34% | 12.0088 | -0.49784 | -1.1533 | [8 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 109.1% | -6.0461 | -1.3364 | 1.8274 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_104 | alfaA |  |  | 0.00025478 | 5.0269e-05 | 7.8292e-05 | -1.2644e-06 | 7.3593e-05 | -1.4886e-06 | 7.7344e-05 | 6.3083e-06 | -4.4093e-07 | [5 empty] | 20.6563 | 4.5916 | 7.1621 | 3.2805e-05 | 7.5484 | 2.3223e-05 | 1.2608 | 7.5828e-06 | -6.0819e-07 |  |  |  | 1 | 0 | 0 | 20.656 | 4.5916 | 7.162 | 3.407e-05 | 7.5484 | 2.4711e-05 | 1.2607 | 1.2745e-06 | -1.6726e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7851 | 20.656 | 4.5916 | 7.162 | 7.5484 | 1.2607 | 0.22229 | 0.34673 | 0.36543 | 0.061033 | 1.5598 | 1.644 | 0.27457 | 1.0539 | 0.17603 | 0.16702 | -1.5038 | -1.0592 | -1.0067 | -2.7963 | [10 empty] | 0 | 1.7851 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014084 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 5 | 103.5% | -5.4638 | -3.8158 | 1.5433 | SRM | Lab 1 | Jun 23, 2022 | 6 | 109.6% | -9.932 | -3.6404 | 3.7467 | iRM_12
Lab 1 | Jun 22, 2022 | 5 | Zr_108 | iRM_12 |  |  | 0.0002656 | 5.024e-05 | 9.7376e-05 | 1.8513e-06 | 7.2317e-05 | 1.8579e-06 | 6.5051e-05 | -1.1893e-06 | -2.0933e-06 | [5 empty] | 19.9664 | 4.439 | 6.9253 | 4.1246e-05 | 7.3013 | 2.1674e-05 | 1.2199 | 1.8679e-05 | -1.9519e-06 |  |  |  | 1 | 0 | 0 | 19.9661 | 4.439 | 6.9252 | 3.9395e-05 | 7.3012 | 1.9816e-05 | 1.2198 | 1.9869e-05 | 1.4143e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8007 | 19.9661 | 4.439 | 6.9252 | 7.3012 | 1.2198 | 0.22233 | 0.34685 | 0.36568 | 0.061095 | 1.5601 | 1.6448 | 0.2748 | 1.0543 | 0.17614 | 0.16707 | -1.5036 | -1.0589 | -1.006 | -2.7953 | [10 empty] | 0 | 1.8007 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014082 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | -4.1567 | -4.644 | 16.8741 | [8 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 112.3% | -0.25803 | 4.6887 | 26.2398 | iRM_12 | Rejected for intensity mismatch > 12%
Lab 1 | Jun 22, 2022 | 5 | Zr_112 | Ames |  |  | 0.00048212 | 9.2306e-05 | 0.00015697 | -2.225e-06 | 0.00015799 | 1.451e-06 | 8.5355e-05 | 6.6896e-06 | 1.2717e-06 | [5 empty] | 20.7211 | 4.6071 | 7.1879 | 2.8135e-05 | 7.579 | 2.9046e-05 | 1.2664 | 1.6864e-05 | 9.0206e-06 |  |  |  | 1 | 0 | 0 | 20.7207 | 4.607 | 7.1877 | 3.036e-05 | 7.5789 | 2.7595e-05 | 1.2663 | 1.0174e-05 | 7.7488e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.806 | 20.7207 | 4.607 | 7.1877 | 7.5789 | 1.2663 | 0.22234 | 0.34689 | 0.36576 | 0.061114 | 1.5602 | 1.6451 | 0.27487 | 1.0544 | 0.17618 | 0.16709 | -1.5036 | -1.0588 | -1.0058 | -2.795 | [10 empty] | 0 | 1.806 | 0.21794 | 0.33337 | 0.3381 | 0.05438 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9118 | 0.42505 | 0.43914 | -1.3882 | 0.014083 | -1.8133 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 5 | 104.1% | 0.46935 | -3.1964 | -22.9137
Lab 1 | Jun 22, 2022 | 5 | Zr_116 | iRM_12 |  |  | 0.00024473 | 3.8609e-05 | 7.8059e-05 | -2.7983e-06 | 6.5915e-05 | -5.3416e-06 | 6.396e-05 | 5.1871e-06 | -8.0482e-07 | [5 empty] | 19.852 | 4.4136 | 6.8856 | 3.8097e-05 | 7.2594 | 1.9422e-05 | 1.2129 | 1.2024e-05 | -1.871e-07 |  |  |  | 1 | 0 | 0 | 19.8517 | 4.4135 | 6.8856 | 4.0896e-05 | 7.2594 | 2.4763e-05 | 1.2128 | 6.8366e-06 | 6.1772e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8007 | 19.8517 | 4.4135 | 6.8856 | 7.2594 | 1.2128 | 0.22233 | 0.34685 | 0.36568 | 0.061094 | 1.5601 | 1.6448 | 0.2748 | 1.0543 | 0.17614 | 0.16707 | -1.5036 | -1.0589 | -1.006 | -2.7953 | [10 empty] | 0 | 1.8007 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014077 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.48% | 3.0948 | 5.6061 | -9.0191 | [8 empty] | Bottle 1 | Lab 2 | May 28, 2020 | 1 | 102.3% | 3.4224 | 0.60407 | -2.1549 | SRM 3169 | Avg | -2.4926 | -0.016189 | 5.2006
Lab 1 | Jun 22, 2022 | 5 | Zr_120 | IPGP |  |  | 0.00024237 | 4.8537e-05 | 7.5793e-05 | 2.5087e-06 | 7.1419e-05 | 2.0476e-06 | 6.9391e-05 | 7.3619e-06 | -4.0965e-06 | [5 empty] | 20.8808 | 4.6422 | 7.2422 | 2.8747e-05 | 7.6351 | 2.3922e-05 | 1.2756 | 1.5997e-05 | 2.1987e-06 |  |  |  | 1 | 0 | 0 | 20.8806 | 4.6422 | 7.2421 | 2.6238e-05 | 7.635 | 2.1875e-05 | 1.2756 | 8.6349e-06 | 6.2952e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7989 | 20.8806 | 4.6422 | 7.2421 | 7.635 | 1.2756 | 0.22232 | 0.34683 | 0.36565 | 0.061088 | 1.5601 | 1.6447 | 0.27478 | 1.0542 | 0.17613 | 0.16707 | -1.5036 | -1.0589 | -1.0061 | -2.7954 | [10 empty] | 0 | 1.7989 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014078 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 5 | 104.9% | 2.8394 | 1.8575 | 30.2867 | bottle 1 | Lab 2 | May 28, 2020 | 1 | 105.1% | -0.69285 | 5.3125 | 20.2282 | SRM 3169 | 2SD | 7.2001 | 6.7427 | 24.6243
Lab 1 | Jun 22, 2022 | 5 | Zr_124 | iRM_12 |  |  | 0.00024286 | 5.3912e-05 | 8.2463e-05 | 1.5242e-06 | 6.2721e-05 | 1.3056e-06 | 5.3197e-05 | 1.0849e-05 | -3.6306e-06 | [5 empty] | 19.9471 | 4.4347 | 6.9187 | 4.5712e-05 | 7.2944 | 2.2618e-05 | 1.2187 | 1.3683e-05 | -2.0204e-06 |  |  |  | 1 | 0 | 0 | 19.9468 | 4.4347 | 6.9186 | 4.4188e-05 | 7.2944 | 2.1312e-05 | 1.2187 | 2.8332e-06 | 1.6102e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8014 | 19.9468 | 4.4347 | 6.9186 | 7.2944 | 1.2187 | 0.22233 | 0.34685 | 0.36569 | 0.061096 | 1.5601 | 1.6448 | 0.2748 | 1.0543 | 0.17614 | 0.16707 | -1.5036 | -1.0589 | -1.006 | -2.7953 | [10 empty] | 0 | 1.8014 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014083 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | -1.906 | -4.2873 | -2.0858 | [8 empty] | bottle 1 | Lab 2 | May 28, 2020 | 1 | 107% | -3.894 | -4.8872 | -0.10741 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 22, 2022 | 5 | Zr_128 | iRM_UR |  |  | 0.00025629 | 4.2287e-05 | 8.0728e-05 | -7.6668e-06 | 6.7289e-05 | 1.2413e-06 | 6.0242e-05 | -4.8153e-06 | -2.7586e-06 | [5 empty] | 19.8228 | 4.4071 | 6.8755 | 2.7987e-05 | 7.2488 | 1.7945e-05 | 1.2111 | 1.2563e-05 | -5.7815e-06 |  |  |  | 1 | 0 | 0 | 19.8226 | 4.4071 | 6.8754 | 3.5654e-05 | 7.2487 | 1.6703e-05 | 1.211 | 1.7378e-05 | -3.023e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8006 | 19.8226 | 4.4071 | 6.8754 | 7.2487 | 1.211 | 0.22233 | 0.34685 | 0.36568 | 0.061094 | 1.5601 | 1.6448 | 0.2748 | 1.0543 | 0.17614 | 0.16707 | -1.5036 | -1.0589 | -1.006 | -2.7953 | [10 empty] | 0 | 1.8006 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014083 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 5 | 99.45% | 7.8604 | -1.5128 | 15.6931 | bottle 1 | Lab 2 | May 28, 2020 | 1 | 108.6% | -2.7444 | -1.6295 | 20.9709 | SRM 3169 | 2SE | 2.9394 | 2.7527 | 10.0528
Lab 1 | Jun 22, 2022 | 5 | Zr_132 | iRM_12 |  |  | 0.0041067 | 0.00090747 | 0.0014246 | 1.5331e-06 | 0.0014884 | 1.2409e-06 | 0.00029039 | 4.9504e-06 | 1.25e-06 | [5 empty] | 19.9232 | 4.4295 | 6.9107 | 4.2018e-05 | 7.2862 | 2.0989e-05 | 1.2174 | 1.3462e-05 | -5.1291e-07 |  |  |  | 1 | 0 | 0 | 19.9191 | 4.4286 | 6.9093 | 4.0485e-05 | 7.2847 | 1.9748e-05 | 1.2171 | 8.5113e-06 | -1.7629e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8029 | 19.9191 | 4.4286 | 6.9093 | 7.2847 | 1.2171 | 0.22233 | 0.34687 | 0.36572 | 0.061102 | 1.5601 | 1.6449 | 0.27483 | 1.0543 | 0.17616 | 0.16708 | -1.5036 | -1.0588 | -1.0059 | -2.7952 | [10 empty] | 0 | 1.8029 | 0.21793 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43914 | -1.3882 | 0.01408 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | -4.6501 | 4.2325 | -8.4201 | [8 empty] | bottle 1 | Lab 2 | May 28, 2020 | 1 | 110.6% | -3.7659 | -0.63668 | -8.1506 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_136 | SRM |  |  | 0.00024617 | 4.8165e-05 | 8.1582e-05 | 2.1127e-06 | 7.1028e-05 | -1.6869e-06 | 6.1174e-05 | 7.2896e-06 | -2.6562e-06 | [5 empty] | 20.3973 | 4.5348 | 7.0747 | 4.4509e-05 | 7.4588 | 2.4334e-05 | 1.2462 | 1.2997e-05 | 8.8728e-07 |  |  |  | 1 | 0 | 0 | 20.397 | 4.5348 | 7.0747 | 4.2396e-05 | 7.4587 | 2.6021e-05 | 1.2461 | 5.7071e-06 | 3.5435e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8005 | 20.397 | 4.5348 | 7.0746 | 7.4587 | 1.2461 | 0.22233 | 0.34685 | 0.36568 | 0.061094 | 1.5601 | 1.6448 | 0.27479 | 1.0543 | 0.17614 | 0.16707 | -1.5036 | -1.0589 | -1.006 | -2.7953 | [10 empty] | 0 | 1.8005 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43913 | -1.3882 | 0.014079 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 102.9% | 8.8925 | 3.8862 | 6.1875 | bottle 1 | Lab 2 | May 28, 2020 | 1 | 106.8% | -7.2807 | 1.1397 | 0.41744 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_140 | iRM_12 |  |  | 0.00024835 | 4.8207e-05 | 8.1031e-05 | 7.7974e-06 | 6.9579e-05 | -3.9052e-06 | 5.9874e-05 | 1.3638e-05 | -5.8426e-06 | [5 empty] | 19.7357 | 4.3878 | 6.8455 | 4.2673e-05 | 7.2175 | 2.1012e-05 | 1.2059 | 1.4194e-05 | -1.6792e-06 |  |  |  | 1 | 0 | 0 | 19.7354 | 4.3878 | 6.8455 | 3.4875e-05 | 7.2175 | 2.4918e-05 | 1.2059 | 5.5594e-07 | 4.1634e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8027 | 19.7354 | 4.3878 | 6.8455 | 7.2175 | 1.2059 | 0.22233 | 0.34686 | 0.36571 | 0.061102 | 1.5601 | 1.6449 | 0.27483 | 1.0543 | 0.17616 | 0.16708 | -1.5036 | -1.0588 | -1.0059 | -2.7952 | [10 empty] | 0 | 1.8027 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014086 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.62% | 6.1272 | -6.7758 | 15.3736
Lab 1 | Jun 22, 2022 | 5 | Zr_144 | alfaA |  |  | 0.00094627 | 0.00019903 | 0.0003216 | 1.5661e-06 | 0.00032927 | 3.321e-07 | 0.00010666 | 7.1848e-06 | 1.2093e-06 | [5 empty] | 20.5591 | 4.5702 | 7.1292 | 3.1777e-05 | 7.5144 | 2.3628e-05 | 1.2552 | 1.3666e-05 | -1.0456e-06 |  |  |  | 1 | 0 | 0 | 20.5581 | 4.57 | 7.1288 | 3.0211e-05 | 7.5141 | 2.3296e-05 | 1.2551 | 6.4812e-06 | -2.2549e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7898 | 20.5581 | 4.57 | 7.1288 | 7.5141 | 1.2551 | 0.2223 | 0.34676 | 0.36551 | 0.061052 | 1.5599 | 1.6442 | 0.27464 | 1.054 | 0.17606 | 0.16703 | -1.5037 | -1.0591 | -1.0065 | -2.796 | [10 empty] | 0 | 1.7898 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.01408 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 5 | 104.3% | 0.55742 | 1.4729 | 18.0679 | Bottle 1 | Lab 2 | June 3, 2020 | 2 | 105.6% | -0.6633 | -0.18269 | -11.5675 | SRM 3169 | Avg | 0.4227 | 2.2713 | -2.6492
Lab 1 | Jun 22, 2022 | 5 | Zr_148 | iRM_12 |  |  | 0.0002707 | 5.6424e-05 | 9.1737e-05 | 3.3821e-06 | 7.3677e-05 | 2.247e-06 | 7.4872e-05 | 2.8117e-06 | -3.8014e-06 | [5 empty] | 19.704 | 4.3808 | 6.8348 | 3.6448e-05 | 7.2062 | 2.2113e-05 | 1.2041 | 1.3817e-05 | -2.6201e-06 |  |  |  | 1 | 0 | 0 | 19.7038 | 4.3808 | 6.8347 | 3.3066e-05 | 7.2062 | 1.9866e-05 | 1.204 | 1.1005e-05 | 1.1813e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8036 | 19.7038 | 4.3808 | 6.8347 | 7.2062 | 1.204 | 0.22233 | 0.34687 | 0.36573 | 0.061105 | 1.5602 | 1.645 | 0.27484 | 1.0544 | 0.17616 | 0.16708 | -1.5036 | -1.0588 | -1.0059 | -2.7952 | [10 empty] | 0 | 1.8036 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43914 | -1.3882 | 0.014078 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.72% | -4.9002 | 1.8043 | -16.8466 | [8 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 107% | -0.97925 | 1.764 | 2.3811 | SRM 3169 | 2SD | 8.2589 | 5.9313 | 15.8158
Lab 1 | Jun 22, 2022 | 5 | Zr_152 | Ames |  |  | 0.00025045 | 4.1457e-05 | 7.8831e-05 | -6.431e-07 | 7.139e-05 | 1.863e-06 | 6.4863e-05 | 1.4613e-05 | -2.1905e-06 | [5 empty] | 20.8077 | 4.6265 | 7.2186 | 3.1515e-05 | 7.612 | 3.0262e-05 | 1.272 | 1.8005e-05 | 9.5671e-06 |  |  |  | 1 | 0 | 0 | 20.8075 | 4.6265 | 7.2185 | 3.2158e-05 | 7.6119 | 2.8399e-05 | 1.272 | 3.3919e-06 | 1.1758e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8098 | 20.8075 | 4.6265 | 7.2185 | 7.6119 | 1.2719 | 0.22235 | 0.34692 | 0.36582 | 0.061129 | 1.5602 | 1.6453 | 0.27493 | 1.0545 | 0.17621 | 0.1671 | -1.5035 | -1.0587 | -1.0056 | -2.7948 | [10 empty] | 0 | 1.8098 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24952 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014079 | -1.8133 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 5 | 105.4% | 5.0933 | -3.0196 | -13.2439 | bottle 1 | Lab 2 | June 3, 2020 | 2 | 107.7% | -0.5958 | 1.2993 | 3.0076 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 22, 2022 | 5 | Zr_156 | iRM_12 |  |  | 0.00063805 | 0.00013846 | 0.00021 | -3.1945e-06 | 0.00022681 | -1.9862e-06 | 7.3896e-05 | 9.1847e-06 | -5.8999e-07 | [5 empty] | 19.785 | 4.3989 | 6.8631 | 4.2975e-05 | 7.2364 | 2.0801e-05 | 1.2091 | 9.4e-06 | -7.5668e-07 |  |  |  | 1 | 0 | 0 | 19.7843 | 4.3988 | 6.8629 | 4.6169e-05 | 7.2361 | 2.2787e-05 | 1.2091 | 2.1535e-07 | -1.6669e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8052 | 19.7843 | 4.3988 | 6.8629 | 7.2361 | 1.2091 | 0.22234 | 0.34688 | 0.36575 | 0.061113 | 1.5602 | 1.645 | 0.27487 | 1.0544 | 0.17618 | 0.16709 | -1.5036 | -1.0588 | -1.0058 | -2.795 | [10 empty] | 0 | 1.8052 | 0.21794 | 0.33338 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014073 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | 6.5609 | 9.3334 | 16.0675 | [8 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 108.6% | -5.4689 | 2.4144 | 6.3336 | SRM 3169 | 2SE | 3.3717 | 2.4214 | 6.4568
Lab 1 | Jun 22, 2022 | 5 | Zr_160 | IPGP |  |  | 0.00023481 | 4.3561e-05 | 8.393e-05 | -7.0611e-07 | 7.2808e-05 | -9.5559e-07 | 6.2884e-05 | 3.7332e-06 | -2.5785e-06 | [5 empty] | 20.6285 | 4.5863 | 7.1551 | 2.3814e-05 | 7.5438 | 2.774e-05 | 1.2604 | 1.3044e-05 | 5.5312e-08 |  |  |  | 1 | 0 | 0 | 20.6282 | 4.5862 | 7.1551 | 2.452e-05 | 7.5437 | 2.8695e-05 | 1.2604 | 9.3108e-06 | 2.6338e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.802 | 20.6282 | 4.5862 | 7.1551 | 7.5437 | 1.2604 | 0.22233 | 0.34686 | 0.3657 | 0.0611 | 1.5601 | 1.6449 | 0.27482 | 1.0543 | 0.17615 | 0.16708 | -1.5036 | -1.0588 | -1.0059 | -2.7952 | [10 empty] | 0 | 1.802 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014081 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 5 | 104.4% | -0.28776 | -0.76496 | 3.8615 | bottle 1 | Lab 2 | June 3, 2020 | 2 | 108.8% | 6.1721 | 8.0101 | -3.7042 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_164 | iRM_12 |  |  | 0.0026173 | 0.00058825 | 0.00091006 | 1.293e-06 | 0.00094194 | -2.1417e-06 | 0.00020995 | 4.4016e-06 | 9.7789e-07 | [5 empty] | 19.7248 | 4.3855 | 6.8421 | 3.9808e-05 | 7.2143 | 2.1268e-05 | 1.2055 | 1.0533e-05 | -2.9896e-06 |  |  |  | 1 | 0 | 0 | 19.7222 | 4.3849 | 6.8412 | 3.8515e-05 | 7.2134 | 2.341e-05 | 1.2053 | 6.1313e-06 | -3.9675e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8051 | 19.7222 | 4.3849 | 6.8412 | 7.2134 | 1.2053 | 0.22233 | 0.34688 | 0.36575 | 0.061111 | 1.5602 | 1.645 | 0.27486 | 1.0544 | 0.17618 | 0.16709 | -1.5036 | -1.0588 | -1.0058 | -2.7951 | [10 empty] | 0 | 1.8051 | 0.21793 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3882 | 0.014088 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.2% | -2.9453 | -10.3973 | -10.4189 | [8 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 109.5% | 4.0714 | 0.32291 | -12.3461 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_168 | iRM_UR |  |  | 0.0002412 | 4.5448e-05 | 8.4263e-05 | 2.6857e-06 | 8.0815e-05 | 4.2981e-06 | 6.7098e-05 | 8.8628e-07 | -1.1699e-06 | [5 empty] | 19.378 | 4.3084 | 6.7218 | 2.497e-05 | 7.0872 | 1.8941e-05 | 1.1842 | 7.0392e-06 | -5.5754e-06 |  |  |  | 1 | 0 | 0 | 19.3777 | 4.3083 | 6.7217 | 2.2284e-05 | 7.0872 | 1.4643e-05 | 1.1841 | 6.1529e-06 | -4.4055e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8043 | 19.3777 | 4.3083 | 6.7217 | 7.0872 | 1.1841 | 0.22233 | 0.34688 | 0.36574 | 0.061109 | 1.5602 | 1.645 | 0.27485 | 1.0544 | 0.17617 | 0.16708 | -1.5036 | -1.0588 | -1.0058 | -2.7951 | [10 empty] | 0 | 1.8043 | 0.21794 | 0.33338 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014074 | -1.8133 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 5 | 98.56% | 12.8042 | 11.0442 | -5.5402
Lab 1 | Jun 22, 2022 | 5 | Zr_172 | iRM_12 |  |  | 0.00025874 | 5.5936e-05 | 8.3706e-05 | -1.0992e-06 | 7.4239e-05 | -8.4047e-07 | 5.9862e-05 | 2.3565e-05 | -3.7608e-06 | [5 empty] | 19.6011 | 4.358 | 6.7992 | 4.0504e-05 | 7.1691 | 2.1846e-05 | 1.1979 | 1.5462e-05 | -2.4422e-06 |  |  |  | 1 | 0 | 0 | 19.6008 | 4.3579 | 6.7992 | 4.1603e-05 | 7.169 | 2.2687e-05 | 1.1978 | -8.1035e-06 | 1.3186e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8051 | 19.6008 | 4.3579 | 6.7992 | 7.169 | 1.1978 | 0.22233 | 0.34688 | 0.36575 | 0.061112 | 1.5602 | 1.645 | 0.27487 | 1.0544 | 0.17618 | 0.16709 | -1.5036 | -1.0588 | -1.0058 | -2.795 | [10 empty] | 0 | 1.8051 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43915 | -1.3882 | 0.014081 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.41% | -5.884 | 4.7206 | 4.0243 | [8 empty] | Bottle 2 | Lab 2 | May 28, 2020 | 1 | 100.9% | 1.4674 | -1.4703 | 2.6604 | SRM 3169 | Avg | 2.3114 | 1.6696 | -6.6929
Lab 1 | Jun 22, 2022 | 5 | Zr_176 | SRM |  |  | 0.00023745 | 5.3299e-05 | 8.9563e-05 | 2.4905e-06 | 8.004e-05 | -3.0259e-06 | 5.9587e-05 | 1.4934e-05 | -4.9212e-06 | [5 empty] | 20.4334 | 4.543 | 7.0876 | 3.8736e-05 | 7.4728 | 2.3231e-05 | 1.2486 | 1.1994e-05 | 8.8546e-07 |  |  |  | 1 | 0 | 0 | 20.4331 | 4.5429 | 7.0875 | 3.6245e-05 | 7.4727 | 2.6257e-05 | 1.2485 | -2.9396e-06 | 5.8066e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.803 | 20.4331 | 4.5429 | 7.0875 | 7.4727 | 1.2485 | 0.22233 | 0.34686 | 0.36572 | 0.061104 | 1.5601 | 1.6449 | 0.27483 | 1.0544 | 0.17616 | 0.16708 | -1.5036 | -1.0588 | -1.0059 | -2.7952 | [10 empty] | 0 | 1.803 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42504 | 0.43913 | -1.3882 | 0.014086 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 103.9% | 13.1982 | -3.1513 | 1.034 | bottle 2 | Lab 2 | May 28, 2020 | 1 | 103% | -6.5635 | 1.963 | -4.7327 | SRM 3169 | 2SD | 12.8166 | 4.7906 | 17.2269
Lab 1 | Jun 22, 2022 | 5 | Zr_180 | iRM_12 |  |  | 0.0016907 | 0.00036047 | 0.00058971 | 4.0184e-06 | 0.0005982 | -1.9903e-06 | 0.00015828 | 9.4459e-06 | -2.9595e-06 | [5 empty] | 19.7143 | 4.3832 | 6.8387 | 4.116e-05 | 7.2108 | 2.0229e-05 | 1.2049 | 1.7889e-05 | -2.6005e-06 |  |  |  | 1 | 0 | 0 | 19.7126 | 4.3829 | 6.8381 | 3.7141e-05 | 7.2102 | 2.222e-05 | 1.2048 | 8.4435e-06 | 3.5898e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8061 | 19.7126 | 4.3829 | 6.8381 | 7.2102 | 1.2048 | 0.22234 | 0.34689 | 0.36577 | 0.061116 | 1.5602 | 1.6451 | 0.27488 | 1.0544 | 0.17618 | 0.16709 | -1.5036 | -1.0588 | -1.0058 | -2.795 | [10 empty] | 0 | 1.8061 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014084 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 101% | 3.5994 | -4.7899 | 8.7539 | [8 empty] | bottle 2 | Lab 2 | May 28, 2020 | 1 | 104.9% | -3.4635 | 0.011026 | -7.8787 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 22, 2022 | 5 | Zr_184 | alfaA |  |  | 0.00026148 | 4.8987e-05 | 8.5854e-05 | 1.7132e-06 | 7.7666e-05 | -6.5666e-07 | 5.8765e-05 | 9.7577e-06 | -4.0073e-06 | [5 empty] | 20.4409 | 4.5441 | 7.0885 | 3.2001e-05 | 7.472 | 2.2813e-05 | 1.2482 | 1.2317e-05 | 8.6077e-07 |  |  |  | 1 | 0 | 0 | 20.4407 | 4.544 | 7.0884 | 3.0288e-05 | 7.4719 | 2.347e-05 | 1.2481 | 2.5594e-06 | 4.868e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.792 | 20.4407 | 4.544 | 7.0884 | 7.4719 | 1.2481 | 0.2223 | 0.34678 | 0.36554 | 0.06106 | 1.5599 | 1.6443 | 0.27467 | 1.0541 | 0.17608 | 0.16704 | -1.5037 | -1.0591 | -1.0064 | -2.7959 | [10 empty] | 0 | 1.792 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014083 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 5 | 104.4% | -5.1178 | -1.8854 | 10.1725 | bottle 2 | Lab 2 | May 28, 2020 | 1 | 105.6% | 4.6574 | 1.7132 | 2.6216 | SRM 3169 | 2SE | 5.2323 | 1.9558 | 7.0329
Lab 1 | Jun 22, 2022 | 5 | Zr_188 | iRM_12 |  |  | 0.00026931 | 5.4993e-05 | 9.2459e-05 | 2.5454e-07 | 7.1854e-05 | 3.1223e-07 | 5.54e-05 | 1.557e-05 | -3.6266e-06 | [5 empty] | 19.4473 | 4.3239 | 6.7461 | 4.2793e-05 | 7.1132 | 1.9987e-05 | 1.1886 | 1.2161e-05 | -1.4324e-06 |  |  |  | 1 | 0 | 0 | 19.447 | 4.3238 | 6.746 | 4.2538e-05 | 7.1131 | 1.9674e-05 | 1.1885 | -3.4083e-06 | 2.1942e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8063 | 19.447 | 4.3238 | 6.746 | 7.1131 | 1.1885 | 0.22234 | 0.34689 | 0.36577 | 0.061116 | 1.5602 | 1.6451 | 0.27488 | 1.0544 | 0.17618 | 0.16709 | -1.5036 | -1.0587 | -1.0058 | -2.795 | [10 empty] | 0 | 1.8063 | 0.21794 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014078 | -1.8133 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 98.63% | 7.4383 | 2.7903 | -13.9096 | [8 empty] | bottle 2 | Lab 2 | May 28, 2020 | 1 | 107.7% | 8.3517 | 2.163 | -17.5245 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_192 | Ames |  |  | 0.0043274 | 0.00095743 | 0.001491 | -1.2915e-06 | 0.0015672 | -1.6069e-06 | 0.00031032 | 9.0921e-06 | -4.539e-06 | [5 empty] | 20.8392 | 4.6336 | 7.2298 | 3.9175e-05 | 7.6241 | 3.0305e-05 | 1.2741 | 1.7619e-05 | 9.1897e-06 |  |  |  | 1 | 0 | 0 | 20.8349 | 4.6327 | 7.2283 | 4.0467e-05 | 7.6225 | 3.1912e-05 | 1.2738 | 8.5267e-06 | 1.3729e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8117 | 20.8349 | 4.6327 | 7.2283 | 7.6225 | 1.2738 | 0.22235 | 0.34693 | 0.36585 | 0.061138 | 1.5603 | 1.6454 | 0.27496 | 1.0545 | 0.17623 | 0.16711 | -1.5035 | -1.0586 | -1.0055 | -2.7946 | [10 empty] | 0 | 1.8117 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014081 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 5 | 106.4% | -0.57739 | -3.7109 | 10.5593 | bottle 2 | Lab 2 | May 28, 2020 | 1 | 104.6% | 9.4189 | 5.6377 | -15.3036 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_196 | iRM_12 |  |  | 0.0018194 | 0.00040761 | 0.00063325 | 1.3579e-05 | 0.00065542 | -1.0432e-06 | 0.00015737 | 1.4508e-05 | -4.3871e-06 | [5 empty] | 19.7219 | 4.385 | 6.8416 | 3.8807e-05 | 7.214 | 2.0429e-05 | 1.2055 | 1.4601e-05 | -2.6517e-06 |  |  |  | 1 | 0 | 0 | 19.7201 | 4.3846 | 6.8409 | 2.5229e-05 | 7.2134 | 2.1472e-05 | 1.2053 | 9.2839e-08 | 1.7354e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8075 | 19.7201 | 4.3846 | 6.8409 | 7.2134 | 1.2053 | 0.22234 | 0.3469 | 0.36579 | 0.061121 | 1.5602 | 1.6452 | 0.2749 | 1.0544 | 0.17619 | 0.1671 | -1.5036 | -1.0587 | -1.0057 | -2.7949 | [10 empty] | 0 | 1.8075 | 0.21793 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43915 | -1.3882 | 0.014077 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 101% | -13.6222 | 0.63264 | 1.7217
Lab 1 | Jun 22, 2022 | 5 | Zr_200 | IPGP |  |  | 0.000242 | 5.1954e-05 | 7.8262e-05 | 6.0805e-06 | 8.5083e-05 | 1.877e-06 | 6.7893e-05 | 8.0838e-07 | -5.6504e-06 | [5 empty] | 20.8077 | 4.6263 | 7.218 | 2.7554e-05 | 7.6108 | 2.3549e-05 | 1.2718 | 1.6938e-05 | 4.5359e-08 |  |  |  | 1 | 0 | 0 | 20.8075 | 4.6263 | 7.2179 | 2.1474e-05 | 7.6107 | 2.1672e-05 | 1.2717 | 1.6129e-05 | 5.6957e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8062 | 20.8075 | 4.6263 | 7.2179 | 7.6107 | 1.2717 | 0.22234 | 0.34689 | 0.36577 | 0.061117 | 1.5602 | 1.6451 | 0.27489 | 1.0544 | 0.17619 | 0.16709 | -1.5036 | -1.0588 | -1.0058 | -2.795 | [10 empty] | 0 | 1.8062 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3882 | 0.014082 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 5 | 105.8% | -2.6551 | -5.0104 | 16.5397 | Bottle 2 | Lab 2 | June 3, 2020 | 2 | 105.3% | 5.9912 | 0.96502 | 7.093 | SRM 3169 | Avg | 0.045558 | 3.6719 | 0.88254
Lab 1 | Jun 22, 2022 | 5 | Zr_204 | iRM_12 |  |  | 0.00023477 | 4.5299e-05 | 7.8457e-05 | 5.3871e-06 | 7.6663e-05 | 1.3551e-06 | 6.4052e-05 | 6.7404e-06 | -2.9927e-06 | [5 empty] | 19.61 | 4.3603 | 6.8032 | 4.4949e-05 | 7.174 | 2.143e-05 | 1.1989 | 1.0666e-05 | -1.4742e-06 |  |  |  | 1 | 0 | 0 | 19.6098 | 4.3602 | 6.8031 | 3.9561e-05 | 7.1739 | 2.0075e-05 | 1.1988 | 3.926e-06 | 1.5185e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8103 | 19.6098 | 4.3602 | 6.8031 | 7.1739 | 1.1988 | 0.22235 | 0.34692 | 0.36583 | 0.061133 | 1.5603 | 1.6453 | 0.27494 | 1.0545 | 0.17621 | 0.16711 | -1.5035 | -1.0587 | -1.0056 | -2.7947 | [10 empty] | 0 | 1.8103 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014077 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | 9.9671 | 1.8663 | 3.3676 | [8 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 106.4% | -0.4044 | 4.8708 | 27.7052 | SRM 3169 | 2SD | 7.0874 | 4.9305 | 35.2629
Lab 1 | Jun 22, 2022 | 5 | Zr_208 | iRM_UR |  |  | 0.0031147 | 0.00068528 | 0.0010778 | 1.4434e-05 | 0.0011075 | 1.6815e-06 | 0.00023388 | 4.5875e-06 | -2.2061e-06 | [5 empty] | 19.2785 | 4.2864 | 6.6878 | 2.6641e-05 | 7.0519 | 2.013e-05 | 1.1784 | 5.6206e-06 | -3.644e-06 |  |  |  | 1 | 0 | 0 | 19.2754 | 4.2858 | 6.6867 | 1.2207e-05 | 7.0508 | 1.8448e-05 | 1.1782 | 1.0331e-06 | -1.4379e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8079 | 19.2754 | 4.2858 | 6.6867 | 7.0508 | 1.1782 | 0.22234 | 0.3469 | 0.36579 | 0.061123 | 1.5602 | 1.6452 | 0.27491 | 1.0545 | 0.1762 | 0.1671 | -1.5035 | -1.0587 | -1.0057 | -2.7949 | [10 empty] | 0 | 1.8079 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42506 | 0.43913 | -1.3882 | 0.014076 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 22, 2022 | 5 | 98.84% | 3.6285 | 2.9989 | 3.2024 | bottle 2 | Lab 2 | June 3, 2020 | 2 | 106.6% | -0.0079351 | 1.4716 | -9.0703 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 22, 2022 | 5 | Zr_212 | iRM_12 |  |  | 0.00059124 | 0.00013363 | 0.00019962 | 1.0364e-05 | 0.00019665 | 1.8023e-06 | 7.8427e-05 | 3.8073e-06 | -4.712e-06 | [5 empty] | 19.3944 | 4.3122 | 6.728 | 4.2359e-05 | 7.0945 | 2.3711e-05 | 1.1855 | 1.2013e-05 | -2.4843e-06 |  |  |  | 1 | 0 | 0 | 19.3938 | 4.3121 | 6.7278 | 3.1995e-05 | 7.0943 | 2.1908e-05 | 1.1855 | 8.2054e-06 | 2.2278e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8085 | 19.3938 | 4.3121 | 6.7278 | 7.0943 | 1.1855 | 0.22234 | 0.34691 | 0.3658 | 0.061126 | 1.5602 | 1.6452 | 0.27492 | 1.0545 | 0.1762 | 0.1671 | -1.5035 | -1.0587 | -1.0057 | -2.7948 | [10 empty] | 0 | 1.8085 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014081 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 99.09% | -0.30235 | 1.3291 | 4.7308 | [8 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 108.1% | -5.1477 | 2.176 | 11.1908 | SRM 3169 | 2SE | 2.8934 | 2.0129 | 14.396
Lab 1 | Jun 22, 2022 | 5 | Zr_216 | SRM |  |  | 0.00024708 | 5.976e-05 | 7.8165e-05 | -1.6787e-06 | 7.0282e-05 | 4.8684e-07 | 5.5914e-05 | 1.0858e-05 | -1.7357e-06 | [5 empty] | 20.228 | 4.4975 | 7.0169 | 4.0978e-05 | 7.3987 | 2.6262e-05 | 1.2363 | 1.3144e-05 | 1.1697e-06 |  |  |  | 1 | 0 | 0 | 20.2278 | 4.4974 | 7.0168 | 4.2656e-05 | 7.3987 | 2.5775e-05 | 1.2363 | 2.2863e-06 | 2.9054e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8062 | 20.2278 | 4.4974 | 7.0168 | 7.3987 | 1.2363 | 0.22234 | 0.34689 | 0.36577 | 0.061118 | 1.5602 | 1.6451 | 0.27489 | 1.0544 | 0.17619 | 0.16709 | -1.5036 | -1.0587 | -1.0058 | -2.795 | [10 empty] | 0 | 1.8062 | 0.21794 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014076 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 22, 2022 | 5 | 103.9% | 5.424 | 8.3171 | 20.3895 | bottle 2 | Lab 2 | June 3, 2020 | 2 | 108% | 0.26107 | 5.5659 | -11.7135 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_220 | iRM_12 |  |  | 0.0047139 | 0.0010484 | 0.0016398 | 9.3283e-06 | 0.0017051 | 4.2717e-07 | 0.00033281 | 9.8057e-06 | -3.8803e-06 | [5 empty] | 19.54 | 4.3447 | 6.7788 | 3.9117e-05 | 7.1485 | 2.0784e-05 | 1.1946 | 5.0897e-06 | -2.6736e-07 |  |  |  | 1 | 0 | 0 | 19.5352 | 4.3436 | 6.7772 | 2.9789e-05 | 7.1467 | 2.0357e-05 | 1.1943 | -4.716e-06 | 3.6129e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8107 | 19.5352 | 4.3436 | 6.7772 | 7.1467 | 1.1943 | 0.22235 | 0.34692 | 0.36584 | 0.061134 | 1.5603 | 1.6453 | 0.27495 | 1.0545 | 0.17622 | 0.16711 | -1.5035 | -1.0587 | -1.0056 | -2.7947 | [10 empty] | 0 | 1.8107 | 0.21793 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3882 | 0.014088 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.4% | -4.9799 | -8.4168 | -17.1683 | [8 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 109.1% | -0.41889 | 6.9819 | -19.9101 | SRM 3169
Lab 1 | Jun 22, 2022 | 5 | Zr_224 | alfaA |  |  | 0.00046279 | 9.565e-05 | 0.00015751 | 5.8404e-06 | 0.00014915 | -3.399e-06 | 6.6871e-05 | 1.4964e-05 | -3.0327e-06 | [5 empty] | 20.6622 | 4.5935 | 7.166 | 2.9124e-05 | 7.5545 | 2.2597e-05 | 1.2621 | 1.6762e-05 | 1.5247e-06 |  |  |  | 1 | 0 | 0 | 20.6617 | 4.5934 | 7.1659 | 2.3284e-05 | 7.5544 | 2.5996e-05 | 1.262 | 1.7981e-06 | 4.5573e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.7971 | 20.6617 | 4.5934 | 7.1659 | 7.5544 | 1.262 | 0.22231 | 0.34682 | 0.36562 | 0.061081 | 1.56 | 1.6446 | 0.27475 | 1.0542 | 0.17612 | 0.16706 | -1.5037 | -1.059 | -1.0062 | -2.7956 | [10 empty] | 0 | 1.7971 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014086 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 22, 2022 | 5 | 105.8% | -0.47715 | -2.8373 | -3.4785
Lab 1 | Jun 22, 2022 | 5 | Zr_228 | iRM_12 |  |  | 0.0002845 | 5.4275e-05 | 9.9391e-05 | 4.9639e-06 | 9.5727e-05 | 4.3944e-06 | 5.7913e-05 | 1.2847e-05 | -3.3598e-06 | [5 empty] | 19.5272 | 4.3419 | 6.7747 | 4.1198e-05 | 7.1443 | 1.9519e-05 | 1.194 | 9.6213e-06 | -1.5482e-06 |  |  |  | 1 | 0 | 0 | 19.527 | 4.3419 | 6.7746 | 3.6234e-05 | 7.1442 | 1.5124e-05 | 1.1939 | -3.2256e-06 | 1.8116e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8122 | 19.527 | 4.3419 | 6.7746 | 7.1442 | 1.1939 | 0.22235 | 0.34694 | 0.36586 | 0.061141 | 1.5603 | 1.6454 | 0.27498 | 1.0546 | 0.17623 | 0.16712 | -1.5035 | -1.0586 | -1.0055 | -2.7946 | [10 empty] | 0 | 1.8122 | 0.21794 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.014078 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.1% | 0.53191 | 7.8953 | 11.3902 | [8 empty] | Bottle 3 | Lab 2 | May 28, 2020 | 1 | 102% | -5.9992 | -0.77187 | -10.1495 | SRM 3169 | Avg | -1.2489 | -0.86351 | -7.2463
Lab 1 | Jun 22, 2022 | 5 | Zr_232 | Ames |  |  | 0.00024038 | 3.7231e-05 | 8.1267e-05 | 1.1303e-05 | 6.5275e-05 | 9.5091e-07 | 4.8615e-05 | 4.8149e-06 | -2.2585e-06 | [5 empty] | 21.0362 | 4.6778 | 7.2992 | 3.4971e-05 | 7.6984 | 2.9921e-05 | 1.2867 | 1.5694e-05 | 8.4231e-06 |  |  |  | 1 | 0 | 0 | 21.036 | 4.6778 | 7.2991 | 2.3667e-05 | 7.6983 | 2.897e-05 | 1.2867 | 1.0879e-05 | 1.0682e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8182 | 21.036 | 4.6778 | 7.2991 | 7.6983 | 1.2867 | 0.22237 | 0.34698 | 0.36596 | 0.061164 | 1.5604 | 1.6457 | 0.27506 | 1.0547 | 0.17628 | 0.16713 | -1.5034 | -1.0585 | -1.0052 | -2.7942 | [10 empty] | 0 | 1.8182 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014082 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 22, 2022 | 5 | 107.8% | 2.9241 | -1.327 | -15.1418 | bottle 3 | Lab 2 | May 28, 2020 | 1 | 103.5% | -0.44047 | -0.10692 | -7.0333 | SRM 3169 | 2SD | 6.3921 | 2.4244 | 20.6823
Lab 1 | Jun 22, 2022 | 5 | Zr_236 | iRM_12 |  |  | 0.00038706 | 7.5748e-05 | 0.00013408 | 1.1585e-05 | 0.00012586 | 3.2389e-07 | 7.3979e-05 | 1.5096e-05 | 1.1897e-06 | [5 empty] | 19.4887 | 4.3334 | 6.7614 | 3.7152e-05 | 7.1303 | 2.2015e-05 | 1.1917 | 1.544e-05 | -1.0776e-06 |  |  |  | 1 | 0 | 0 | 19.4884 | 4.3333 | 6.7613 | 2.5567e-05 | 7.1302 | 2.1691e-05 | 1.1916 | 3.4373e-07 | -2.2673e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8126 | 19.4884 | 4.3333 | 6.7613 | 7.1302 | 1.1916 | 0.22235 | 0.34694 | 0.36587 | 0.061143 | 1.5603 | 1.6454 | 0.27498 | 1.0546 | 0.17624 | 0.16712 | -1.5035 | -1.0586 | -1.0055 | -2.7945 | [10 empty] | 0 | 1.8126 | 0.21794 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014084 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 22, 2022 | 5 | 100.3% | 3.2827 | -0.41137 | 10.3538 | [8 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 105.2% | -0.56221 | -1.0899 | 6.9535 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 22, 2022 | 5 | Zr_240 | IPGP |  |  | 0.00023757 | 4.4996e-05 | 8.0704e-05 | -4.6143e-06 | 6.702e-05 | -8.3448e-07 | 6.0023e-05 | 6.7196e-06 | -1.3655e-06 | [5 empty] | 20.762 | 4.6164 | 7.2027 | 2.8963e-05 | 7.5953 | 2.6292e-05 | 1.2693 | 1.3935e-05 | 1.8697e-06 |  |  |  | 1 | 0 | 0 | 20.7618 | 4.6163 | 7.2026 | 3.3577e-05 | 7.5953 | 2.7126e-05 | 1.2692 | 7.2153e-06 | 3.2353e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8101 | 20.7618 | 4.6163 | 7.2026 | 7.5953 | 1.2692 | 0.22235 | 0.34692 | 0.36583 | 0.061133 | 1.5603 | 1.6453 | 0.27494 | 1.0545 | 0.17622 | 0.16711 | -1.5035 | -1.0587 | -1.0056 | -2.7947 | [10 empty] | 0 | 1.8101 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43914 | -1.3882 | 0.01409 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 22, 2022 | 5 | 106.9% | -7.287 | -3.4048 | -0.24618 | bottle 3 | Lab 2 | May 28, 2020 | 1 | 107.1% | 3.3375 | -3.1097 | 1.0962 | SRM 3169 | 2SE | 2.6096 | 0.98976 | 8.4435
Lab 1 | Jun 22, 2022 | 5 | Zr_244 | iRM_12 |  |  | 0.00031073 | 6.2722e-05 | 0.00010886 | 2.8448e-06 | 8.6691e-05 | 3.4987e-06 | 5.7501e-05 | 8.3542e-06 | -1.0389e-06 | [5 empty] | 19.3493 | 4.3024 | 6.7131 | 3.6918e-05 | 7.0794 | 2.0227e-05 | 1.1831 | 1.1482e-05 | -1.98e-07 |  |  |  | 1 | 0 | 0 | 19.349 | 4.3024 | 6.7129 | 3.4073e-05 | 7.0794 | 1.6729e-05 | 1.1831 | 3.1281e-06 | 8.409e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8131 | 19.349 | 4.3024 | 6.7129 | 7.0794 | 1.1831 | 0.22236 | 0.34694 | 0.36588 | 0.061144 | 1.5603 | 1.6455 | 0.27498 | 1.0546 | 0.17624 | 0.16712 | -1.5035 | -1.0586 | -1.0055 | -2.7945 | [10 empty] | 0 | 1.8131 | 0.21794 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42505 | 0.43914 | -1.3882 | 0.014089 | -1.8132 | -1.8273 | [22 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 107.8% | -0.26196 | -0.46977 | -22.5545 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_004 | iRM_12 |  |  | 0.0001262 | 1.1667e-05 | 4.1726e-05 | -3.6542e-06 | 5.8409e-05 | -2.5123e-06 | 4.7763e-05 | 7.546e-06 | -3.3584e-06 | [5 empty] | 20.2534 | 4.5034 | 7.0269 | 3.3872e-05 | 7.4106 | 2.0972e-05 | 1.2385 | 1.024e-05 | -2.3842e-06 |  |  |  | 1 | 0 | 0 | 20.2533 | 4.5034 | 7.0268 | 3.7526e-05 | 7.4106 | 2.3484e-05 | 1.2385 | 2.6938e-06 | 9.7421e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8141 | 20.2533 | 4.5034 | 7.0268 | 7.4106 | 1.2385 | 0.22235 | 0.34695 | 0.36589 | 0.06115 | 1.5603 | 1.6455 | 0.27501 | 1.0546 | 0.17625 | 0.16712 | -1.5035 | -1.0586 | -1.0054 | -2.7944 | [9 empty] | L1 JunS6 | 0 | 1.8141 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3881 | 0.01409 | -1.8132 | -1.8273 | [22 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 105.8% | -3.5671 | 0.36716 | -11.7904 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_008 | iRM_UR |  |  | 0.0020543 | 0.00045224 | 0.00070544 | -3.4411e-06 | 0.00074584 | 4.7104e-07 | 0.00017709 | 3.8423e-06 | -6.2338e-06 | [5 empty] | 19.8097 | 4.4048 | 6.873 | 1.6106e-05 | 7.2484 | 1.9283e-05 | 1.2114 | 4.8057e-06 | -4.066e-06 |  |  |  | 1 | 0 | 0 | 19.8076 | 4.4044 | 6.8723 | 1.9547e-05 | 7.2477 | 1.8812e-05 | 1.2113 | 9.6343e-07 | 2.1678e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8147 | 19.8076 | 4.4044 | 6.8723 | 7.2477 | 1.2113 | 0.22236 | 0.34695 | 0.3659 | 0.061152 | 1.5603 | 1.6456 | 0.27502 | 1.0546 | 0.17625 | 0.16713 | -1.5035 | -1.0586 | -1.0054 | -2.7944 | [10 empty] | 0 | 1.8147 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 23, 2022 | 6 | 98.79% | 9.6158 | -1.0341 | 8.8773
Lab 1 | Jun 23, 2022 | 6 | Zr_012 | iRM_12 |  |  | 0.00023909 | 4.6433e-05 | 7.5126e-05 | -1.5561e-06 | 8.784e-05 | 2.117e-06 | 6.0011e-05 | 8.9684e-07 | -1.8223e-06 | [5 empty] | 19.847 | 4.4132 | 6.8863 | 3.2621e-05 | 7.263 | 2.1188e-05 | 1.2139 | 4.3191e-06 | -2.9698e-06 |  |  |  | 1 | 0 | 0 | 19.8468 | 4.4131 | 6.8863 | 3.4177e-05 | 7.2629 | 1.9071e-05 | 1.2138 | 3.4223e-06 | -1.1474e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8175 | 19.8468 | 4.4131 | 6.8863 | 7.2629 | 1.2138 | 0.22236 | 0.34697 | 0.36595 | 0.061161 | 1.5604 | 1.6457 | 0.27505 | 1.0547 | 0.17627 | 0.16713 | -1.5035 | -1.0585 | -1.0053 | -2.7942 | [10 empty] | 0 | 1.8175 | 0.21793 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3882 | 0.014094 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.27% | -18.9051 | -3.2269 | -32.5442 | [8 empty] | Bottle 3 | Lab 2 | June 3, 2020 | 2 | 104.4% | -2.1335 | -1.5104 | 8.0981 | SRM 3169 | Avg | -2.3408 | -0.46757 | 2.1217
Lab 1 | Jun 23, 2022 | 6 | Zr_016 | SRM |  |  | 0.00023181 | 5.6546e-05 | 7.0289e-05 | -2.1654e-06 | 9.1856e-05 | 1.5012e-06 | 6.3878e-05 | -5.4013e-06 | -2.5175e-06 | [5 empty] | 21.1107 | 4.6941 | 7.3246 | 3.8832e-05 | 7.7248 | 2.355e-05 | 1.2911 | 1.3189e-05 | 3.7324e-06 |  |  |  | 1 | 0 | 0 | 21.1105 | 4.6941 | 7.3245 | 4.0998e-05 | 7.7247 | 2.2049e-05 | 1.291 | 1.859e-05 | 6.2499e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8156 | 21.1105 | 4.6941 | 7.3245 | 7.7247 | 1.291 | 0.22236 | 0.34696 | 0.36592 | 0.061154 | 1.5604 | 1.6456 | 0.27503 | 1.0546 | 0.17626 | 0.16712 | -1.5035 | -1.0585 | -1.0053 | -2.7944 | [10 empty] | 0 | 1.8156 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3882 | 0.014087 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 106.7% | -3.725 | 5.811 | -12.4275 | bottle 3 | Lab 2 | June 3, 2020 | 2 | 105.2% | -7.7928 | -1.855 | 11.9504 | SRM 3169 | 2SD | 9.1318 | 4.1749 | 36.6757
Lab 1 | Jun 23, 2022 | 6 | Zr_020 | iRM_12 |  |  | 0.00025117 | 3.5181e-05 | 7.9e-05 | 6.8093e-06 | 8.6098e-05 | -2.1017e-06 | 5.1675e-05 | 9.5363e-07 | -3.9073e-06 | [5 empty] | 19.7318 | 4.3877 | 6.8466 | 3.2219e-05 | 7.2213 | 2.0862e-05 | 1.207 | 8.4389e-06 | -9.6388e-07 |  |  |  | 1 | 0 | 0 | 19.7316 | 4.3877 | 6.8465 | 2.5409e-05 | 7.2212 | 2.2964e-05 | 1.207 | 7.4853e-06 | 2.9435e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.819 | 19.7316 | 4.3877 | 6.8465 | 7.2212 | 1.207 | 0.22237 | 0.34698 | 0.36597 | 0.061169 | 1.5604 | 1.6458 | 0.27508 | 1.0547 | 0.17629 | 0.16714 | -1.5034 | -1.0585 | -1.0052 | -2.7941 | [10 empty] | 0 | 1.819 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3881 | 0.014091 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.9% | 17.4275 | -1.4466 | 30.5231 | [8 empty] | bottle 3 | Lab 2 | June 3, 2020 | 2 | 105.7% | -5.6981 | 0.32831 | 13.2187 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 23, 2022 | 6 | Zr_024 | alfaA |  |  | 0.00024969 | 4.8486e-05 | 8.4559e-05 | -8.4417e-06 | 8.3724e-05 | 1.3677e-06 | 5.2078e-05 | -2.9648e-07 | -2.6389e-06 | [5 empty] | 21.374 | 4.7521 | 7.4142 | 2.8059e-05 | 7.8176 | 2.4781e-05 | 1.3063 | 1.084e-05 | -2.8644e-08 |  |  |  | 1 | 0 | 0 | 21.3737 | 4.752 | 7.4141 | 3.65e-05 | 7.8175 | 2.3413e-05 | 1.3063 | 1.1136e-05 | 2.6103e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8054 | 21.3737 | 4.752 | 7.4141 | 7.8175 | 1.3063 | 0.22233 | 0.34688 | 0.36575 | 0.061116 | 1.5602 | 1.6451 | 0.27489 | 1.0544 | 0.17619 | 0.1671 | -1.5036 | -1.0588 | -1.0058 | -2.795 | [10 empty] | 0 | 1.8054 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43917 | -1.3881 | 0.014095 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 23, 2022 | 6 | 108.5% | -12.1092 | -6.6809 | 23.6381 | bottle 3 | Lab 2 | June 3, 2020 | 2 | 106.6% | -4.0346 | 0.14098 | 20.8489 | SRM 3169 | 2SE | 3.728 | 1.7044 | 14.9728
Lab 1 | Jun 23, 2022 | 6 | Zr_028 | iRM_12 |  |  | 0.00078653 | 0.00017016 | 0.00026068 | 1.9708e-06 | 0.00028416 | -1.9453e-06 | 8.514e-05 | -1.3103e-06 | -3.4741e-06 | [5 empty] | 19.6549 | 4.3706 | 6.8201 | 3.5424e-05 | 7.1935 | 1.9454e-05 | 1.2024 | 1.7361e-06 | -2.0943e-06 |  |  |  | 1 | 0 | 0 | 19.6542 | 4.3705 | 6.8199 | 3.3453e-05 | 7.1932 | 2.1399e-05 | 1.2023 | 3.0464e-06 | 1.3798e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8201 | 19.6542 | 4.3705 | 6.8199 | 7.1932 | 1.2023 | 0.22237 | 0.34699 | 0.36599 | 0.061172 | 1.5605 | 1.6459 | 0.27509 | 1.0547 | 0.17629 | 0.16714 | -1.5034 | -1.0584 | -1.0052 | -2.7941 | [10 empty] | 0 | 1.8201 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3882 | 0.014086 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.77% | -4.3115 | 2.1552 | -5.4546 | [8 empty] | bottle 3 | Lab 2 | June 3, 2020 | 2 | 106.6% | 4.7211 | 2.9817 | -16.5573 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_032 | Ames |  |  | 0.00024535 | 5.121e-05 | 8.4474e-05 | -1.045e-05 | 8.5768e-05 | -2.6643e-06 | 5.773e-05 | -9.7558e-07 | -1.284e-06 | [5 empty] | 21.4698 | 4.7745 | 7.4508 | 2.4806e-05 | 7.8596 | 2.715e-05 | 1.3139 | 1.2924e-05 | 9.4895e-06 |  |  |  | 1 | 0 | 0 | 21.4695 | 4.7744 | 7.4507 | 3.5256e-05 | 7.8595 | 2.9815e-05 | 1.3138 | 1.39e-05 | 1.0773e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8257 | 21.4695 | 4.7744 | 7.4507 | 7.8595 | 1.3138 | 0.22238 | 0.34704 | 0.36608 | 0.061195 | 1.5605 | 1.6462 | 0.27518 | 1.0549 | 0.17634 | 0.16716 | -1.5034 | -1.0583 | -1.0049 | -2.7937 | [10 empty] | 0 | 1.8257 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014091 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 23, 2022 | 6 | 109.2% | 6.2866 | -5.9086 | 31.678 | bottle 3 | Lab 2 | June 3, 2020 | 2 | 107.6% | 0.89306 | -2.8911 | -24.8284 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_036 | iRM_12 |  |  | 0.00035892 | 7.5151e-05 | 0.00011262 | -8.7207e-06 | 0.00012694 | 4.2025e-06 | 6.213e-05 | 9.1174e-06 | -2.9008e-06 | [5 empty] | 19.6681 | 4.3736 | 6.8249 | 3.0401e-05 | 7.1986 | 1.9617e-05 | 1.2032 | 8.5288e-06 | -2.9062e-06 |  |  |  | 1 | 0 | 0 | 19.6678 | 4.3735 | 6.8247 | 3.9122e-05 | 7.1985 | 1.5415e-05 | 1.2032 | -5.886e-07 | -5.4161e-09 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.821 | 19.6678 | 4.3735 | 6.8247 | 7.1985 | 1.2032 | 0.22237 | 0.347 | 0.366 | 0.061174 | 1.5605 | 1.6459 | 0.2751 | 1.0548 | 0.17629 | 0.16714 | -1.5034 | -1.0584 | -1.0051 | -2.794 | [10 empty] | 0 | 1.821 | 0.21793 | 0.33337 | 0.3381 | 0.054381 | 1.5297 | 1.5514 | 0.24953 | 1.0142 | 0.16312 | 0.16084 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3882 | 0.014085 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.42% | -4.2643 | 0.33657 | -21.2774
Lab 1 | Jun 23, 2022 | 6 | Zr_040 | IPGP |  |  | 0.00025136 | 4.8213e-05 | 7.0168e-05 | -3.201e-06 | 8.9459e-05 | -2.148e-06 | 5.1086e-05 | 4.9843e-06 | -2.149e-06 | [5 empty] | 21.3466 | 4.7467 | 7.4068 | 2.4146e-05 | 7.812 | 2.5535e-05 | 1.3057 | 1.4211e-05 | 1.4737e-06 |  |  |  | 1 | 0 | 0 | 21.3463 | 4.7467 | 7.4068 | 2.7347e-05 | 7.8119 | 2.7683e-05 | 1.3057 | 9.2269e-06 | 3.6227e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8184 | 21.3463 | 4.7467 | 7.4068 | 7.8119 | 1.3057 | 0.22236 | 0.34698 | 0.36596 | 0.061166 | 1.5604 | 1.6458 | 0.27507 | 1.0547 | 0.17628 | 0.16714 | -1.5034 | -1.0585 | -1.0052 | -2.7942 | [10 empty] | 0 | 1.8184 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014086 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 23, 2022 | 6 | 107.9% | 8.086 | -2.1051 | 25.6041 | Bottle 4 | Lab 2 | May 28, 2020 | 1 | 101.7% | 2.6211 | 0.038956 | -0.57946 | SRM 3169 | Avg | 1.1912 | 2.339 | 1.4715
Lab 1 | Jun 23, 2022 | 6 | Zr_044 | iRM_12 |  |  | 0.0024054 | 0.00053274 | 0.00082685 | -5.8334e-06 | 0.00087391 | -5.1807e-06 | 0.00018075 | -4.6152e-07 | -3.9301e-06 | [5 empty] | 19.9118 | 4.4279 | 6.9099 | 2.9366e-05 | 7.2888 | 1.7994e-05 | 1.2184 | 1.0269e-05 | -3.8244e-06 |  |  |  | 1 | 0 | 0 | 19.9094 | 4.4274 | 6.9091 | 3.52e-05 | 7.2879 | 2.3175e-05 | 1.2182 | 1.0731e-05 | 1.0569e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8242 | 19.9094 | 4.4274 | 6.9091 | 7.2879 | 1.2182 | 0.22238 | 0.34703 | 0.36605 | 0.061188 | 1.5605 | 1.6461 | 0.27516 | 1.0548 | 0.17632 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7938 | [10 empty] | 0 | 1.8242 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3881 | 0.014084 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.8% | 0.29366 | 5.2096 | 16.4023 | [8 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 102.3% | 7.5247 | 5.6486 | 1.2341 | SRM 3169 | 2SD | 7.6484 | 5.6577 | 11.5762
Lab 1 | Jun 23, 2022 | 6 | Zr_048 | iRM_UR |  |  | 0.00024769 | 5.1522e-05 | 7.7666e-05 | -4.5397e-06 | 7.4532e-05 | 1.1718e-06 | 4.8002e-05 | 6.9763e-06 | -4.6708e-06 | [5 empty] | 19.8356 | 4.411 | 6.8836 | 2.1133e-05 | 7.2611 | 1.9364e-05 | 1.2138 | 7.4494e-06 | -2.0621e-06 |  |  |  | 1 | 0 | 0 | 19.8353 | 4.411 | 6.8835 | 2.5672e-05 | 7.261 | 1.8192e-05 | 1.2138 | 4.7307e-07 | 2.6087e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8249 | 19.8353 | 4.411 | 6.8835 | 7.261 | 1.2138 | 0.22238 | 0.34703 | 0.36607 | 0.061192 | 1.5605 | 1.6461 | 0.27517 | 1.0548 | 0.17633 | 0.16716 | -1.5034 | -1.0583 | -1.0049 | -2.7937 | [10 empty] | 0 | 1.8249 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.014086 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 23, 2022 | 6 | 99.78% | -5.4902 | 3.3508 | 12.4735 | bottle 4 | Lab 2 | May 28, 2020 | 1 | 102.7% | 0.60224 | 1.025 | 7.7373 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 23, 2022 | 6 | Zr_052 | iRM_12 |  |  | 0.00026041 | 5.7579e-05 | 8.3074e-05 | -3.0419e-06 | 8.4565e-05 | 1.8085e-06 | 5.2099e-05 | 3.0686e-06 | -3.9669e-06 | [5 empty] | 19.8474 | 4.4137 | 6.8878 | 3.8468e-05 | 7.2659 | 2.111e-05 | 1.2146 | 1.0967e-05 | -3.6655e-08 |  |  |  | 1 | 0 | 0 | 19.8471 | 4.4137 | 6.8877 | 4.151e-05 | 7.2658 | 1.9301e-05 | 1.2146 | 7.8986e-06 | 3.9302e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8264 | 19.8471 | 4.4137 | 6.8877 | 7.2658 | 1.2146 | 0.22238 | 0.34704 | 0.36609 | 0.061197 | 1.5605 | 1.6462 | 0.27519 | 1.0549 | 0.17634 | 0.16716 | -1.5034 | -1.0583 | -1.0049 | -2.7937 | [10 empty] | 0 | 1.8264 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43916 | -1.3882 | 0.014094 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.1% | -0.81306 | -9.5399 | -12.2388 | [8 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 103.3% | -3.1022 | -0.11765 | -1.9069 | SRM 3169 | 2SE | 3.1224 | 2.3098 | 4.726
Lab 1 | Jun 23, 2022 | 6 | Zr_056 | SRM |  |  | 0.00025962 | 2.5505e-05 | 7.5433e-05 | -2.8318e-06 | 7.8265e-05 | -1.872e-06 | 4.5808e-05 | 1.116e-05 | -3.1153e-06 | [5 empty] | 21.4553 | 4.7711 | 7.4453 | 4.0262e-05 | 7.8533 | 2.3292e-05 | 1.3127 | 1.2912e-05 | 1.9571e-06 |  |  |  | 1 | 0 | 0 | 21.4551 | 4.7711 | 7.4452 | 4.3093e-05 | 7.8532 | 2.5164e-05 | 1.3127 | 1.7517e-06 | 5.0723e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8228 | 21.4551 | 4.7711 | 7.4452 | 7.8532 | 1.3127 | 0.22238 | 0.34701 | 0.36603 | 0.061183 | 1.5605 | 1.646 | 0.27513 | 1.0548 | 0.17631 | 0.16715 | -1.5034 | -1.0584 | -1.005 | -2.7939 | [10 empty] | 0 | 1.8228 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3881 | 0.014088 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 108.4% | 5.8243 | 1.1404 | 4.2503 | bottle 4 | Lab 2 | May 28, 2020 | 1 | 104.1% | 1.7548 | 6.1889 | -6.3245 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_060 | iRM_12 |  |  | 0.00028592 | 5.5025e-05 | 9.6759e-05 | -1.4931e-06 | 0.00010043 | -1.3235e-06 | 5.0324e-05 | -6.3315e-06 | -4.7112e-06 | [5 empty] | 19.7357 | 4.3888 | 6.8487 | 3.3038e-05 | 7.2241 | 2e-05 | 1.2076 | 1.0274e-05 | 1.6978e-06 |  |  |  | 1 | 0 | 0 | 19.7354 | 4.3887 | 6.8486 | 3.4531e-05 | 7.224 | 2.1324e-05 | 1.2075 | 1.6606e-05 | 6.409e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8236 | 19.7354 | 4.3887 | 6.8486 | 7.224 | 1.2075 | 0.22238 | 0.34702 | 0.36604 | 0.061186 | 1.5605 | 1.646 | 0.27515 | 1.0548 | 0.17632 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7938 | [10 empty] | 0 | 1.8236 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43915 | -1.3881 | 0.014085 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.5% | 3.0466 | 8.7746 | -0.92513 | [8 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 104.2% | -2.2537 | 1.2503 | 8.6686 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_064 | alfaA |  |  | 0.00025185 | 5.1579e-05 | 7.3291e-05 | -2.0814e-06 | 7.5152e-05 | -1.2189e-06 | 4.8322e-05 | 4.7723e-07 | -5.333e-06 | [5 empty] | 21.4679 | 4.773 | 7.447 | 2.5858e-05 | 7.8524 | 2.4976e-05 | 1.3122 | 6.6849e-06 | 1.4695e-06 |  |  |  | 1 | 0 | 0 | 21.4676 | 4.773 | 7.4469 | 2.7939e-05 | 7.8523 | 2.6195e-05 | 1.3121 | 6.2077e-06 | 6.8025e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8066 | 21.4676 | 4.773 | 7.4469 | 7.8523 | 1.3121 | 0.22233 | 0.34689 | 0.36577 | 0.061121 | 1.5602 | 1.6452 | 0.27491 | 1.0544 | 0.1762 | 0.1671 | -1.5036 | -1.0588 | -1.0057 | -2.7949 | [10 empty] | 0 | 1.8066 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014092 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 23, 2022 | 6 | 109.6% | -8.9502 | -2.5241 | 11.4765
Lab 1 | Jun 23, 2022 | 6 | Zr_068 | iRM_12 |  |  | 0.00025399 | 4.4305e-05 | 8.8787e-05 | -5.1163e-07 | 7.7268e-05 | 2.0774e-06 | 5.6819e-05 | 4.2196e-06 | -3.5692e-06 | [5 empty] | 19.4238 | 4.3193 | 6.74 | 3.8492e-05 | 7.109 | 2.0248e-05 | 1.1883 | 9.5345e-06 | -1.4915e-06 |  |  |  | 1 | 0 | 0 | 19.4235 | 4.3192 | 6.7399 | 3.9003e-05 | 7.109 | 1.817e-05 | 1.1883 | 5.3149e-06 | 2.0776e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8206 | 19.4235 | 4.3192 | 6.7399 | 7.109 | 1.1883 | 0.22237 | 0.347 | 0.366 | 0.061176 | 1.5604 | 1.6459 | 0.27511 | 1.0548 | 0.1763 | 0.16715 | -1.5034 | -1.0584 | -1.0051 | -2.794 | [10 empty] | 0 | 1.8206 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43916 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.41% | 0.75001 | -3.6706 | 16.7153 | [8 empty] | Bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.8% | 0.046197 | -0.80021 | 3.7165 | SRM 3169 | Avg | -4.5834 | -0.95722 | 5.7156
Lab 1 | Jun 23, 2022 | 6 | Zr_072 | Ames |  |  | 0.00024988 | 5.3305e-05 | 7.5919e-05 | 8.133e-06 | 8.5676e-05 | 1.2469e-06 | 6.0623e-05 | 8.4128e-06 | -1.3952e-06 | [5 empty] | 21.3767 | 4.7537 | 7.4183 | 2.9903e-05 | 7.8251 | 3.2851e-05 | 1.3081 | 1.3736e-05 | 1.1745e-05 |  |  |  | 1 | 0 | 0 | 21.3765 | 4.7536 | 7.4182 | 2.177e-05 | 7.825 | 3.1604e-05 | 1.3081 | 5.3235e-06 | 1.314e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8244 | 21.3765 | 4.7536 | 7.4182 | 7.825 | 1.308 | 0.22238 | 0.34703 | 0.36606 | 0.061191 | 1.5605 | 1.6461 | 0.27517 | 1.0548 | 0.17633 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7938 | [10 empty] | 0 | 1.8244 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3881 | 0.014088 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 23, 2022 | 6 | 110.3% | -7.7178 | 6.2824 | 5.6612 | bottle 4 | Lab 2 | June 3, 2020 | 2 | 104.9% | -11.582 | -4.5879 | 31.6407 | SRM 3169 | 2SD | 12.8138 | 9.9788 | 39.9326
Lab 1 | Jun 23, 2022 | 6 | Zr_076 | iRM_12 |  |  | 0.00024853 | 3.4883e-05 | 8.0874e-05 | 2.1388e-06 | 8.6496e-05 | -1.4902e-06 | 6.3116e-05 | 5.8518e-06 | -3.2097e-06 | [5 empty] | 19.3443 | 4.3015 | 6.7123 | 3.5756e-05 | 7.0798 | 2.023e-05 | 1.1834 | 1.0528e-05 | -8.2355e-07 |  |  |  | 1 | 0 | 0 | 19.344 | 4.3015 | 6.7122 | 3.3617e-05 | 7.0797 | 2.172e-05 | 1.1833 | 4.6765e-06 | 2.3862e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8202 | 19.344 | 4.3015 | 6.7122 | 7.0797 | 1.1833 | 0.22237 | 0.34699 | 0.36599 | 0.061174 | 1.5604 | 1.6459 | 0.2751 | 1.0548 | 0.1763 | 0.16715 | -1.5034 | -1.0585 | -1.0051 | -2.794 | [10 empty] | 0 | 1.8202 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43916 | -1.3881 | 0.014094 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.55% | -1.1452 | -2.3228 | -10.902 | [8 empty] | bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.8% | -8.7607 | 2.0741 | -7.65 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 23, 2022 | 6 | Zr_080 | IPGP |  |  | 0.00024076 | 3.8347e-05 | 7.464e-05 | -7.0669e-07 | 8.5263e-05 | 4.3342e-06 | 4.9762e-05 | -7.5627e-07 | -4.4016e-06 | [5 empty] | 21.0146 | 4.6728 | 7.2913 | 2.8118e-05 | 7.6898 | 2.3422e-05 | 1.2853 | 1.6492e-05 | 7.679e-07 |  |  |  | 1 | 0 | 0 | 21.0144 | 4.6727 | 7.2912 | 2.8825e-05 | 7.6898 | 1.9088e-05 | 1.2852 | 1.7249e-05 | 5.1695e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8163 | 21.0144 | 4.6727 | 7.2912 | 7.6898 | 1.2852 | 0.22236 | 0.34696 | 0.36593 | 0.061158 | 1.5604 | 1.6457 | 0.27504 | 1.0547 | 0.17627 | 0.16713 | -1.5035 | -1.0585 | -1.0053 | -2.7943 | [10 empty] | 0 | 1.8163 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 23, 2022 | 6 | 108.4% | -1.7075 | -0.0031234 | -1.4487 | bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.9% | 4.9061 | 2.795 | -1.2178 | SRM 3169 | 2SE | 5.2312 | 4.0738 | 16.3024
Lab 1 | Jun 23, 2022 | 6 | Zr_084 | iRM_12 |  |  | 0.0014504 | 0.00032614 | 0.00049482 | 8.2081e-06 | 0.00052581 | -2.7555e-06 | 0.00013163 | 3.9263e-06 | -2.2978e-06 | [5 empty] | 19.4396 | 4.3228 | 6.7455 | 3.2353e-05 | 7.115 | 2.1262e-05 | 1.1893 | 8.3332e-06 | -1.1821e-06 |  |  |  | 1 | 0 | 0 | 19.4382 | 4.3225 | 6.7451 | 2.4144e-05 | 7.1145 | 2.4018e-05 | 1.1892 | 4.4069e-06 | 1.1157e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8211 | 19.4382 | 4.3225 | 6.7451 | 7.1145 | 1.1892 | 0.22237 | 0.347 | 0.366 | 0.061177 | 1.5605 | 1.6459 | 0.27511 | 1.0548 | 0.1763 | 0.16715 | -1.5034 | -1.0584 | -1.0051 | -2.794 | [10 empty] | 0 | 1.8211 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014091 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 101% | -1.5363 | -0.39306 | 7.888 | [8 empty] | bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.9% | -2.664 | 3.8111 | -19.3774 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_088 | iRM_UR |  |  | 0.00024322 | 4.5734e-05 | 7.5232e-05 | 1.253e-05 | 7.37e-05 | -2.6854e-07 | 5.5315e-05 | 7.222e-06 | -2.4773e-06 | [5 empty] | 19.0938 | 4.2459 | 6.6254 | 1.8505e-05 | 6.9882 | 1.8656e-05 | 1.1681 | 1.4946e-05 | -2.8037e-06 |  |  |  | 1 | 0 | 0 | 19.0935 | 4.2458 | 6.6253 | 5.9751e-06 | 6.9881 | 1.8925e-05 | 1.168 | 7.7239e-06 | -3.2644e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8205 | 19.0935 | 4.2458 | 6.6253 | 6.9881 | 1.168 | 0.22237 | 0.34699 | 0.36599 | 0.061175 | 1.5604 | 1.6459 | 0.2751 | 1.0548 | 0.1763 | 0.16715 | -1.5034 | -1.0584 | -1.0051 | -2.794 | [10 empty] | 0 | 1.8205 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3881 | 0.014094 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 23, 2022 | 6 | 98.98% | 4.5633 | -4.9401 | 2.1956 | bottle 4 | Lab 2 | June 3, 2020 | 2 | 104.7% | -9.4461 | -9.0355 | 27.1817 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_092 | iRM_12 |  |  | 0.00024029 | 4.7716e-05 | 7.2272e-05 | 3.7447e-06 | 8.429e-05 | -4.2815e-07 | 5.1185e-05 | 1.5988e-05 | -1.174e-06 | [5 empty] | 19.1438 | 4.257 | 6.6429 | 3.4828e-05 | 7.0067 | 1.9343e-05 | 1.1712 | 1.3853e-05 | -9.6888e-07 |  |  |  | 1 | 0 | 0 | 19.1436 | 4.257 | 6.6428 | 3.1083e-05 | 7.0066 | 1.9771e-05 | 1.1711 | -2.1356e-06 | 2.0513e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.821 | 19.1436 | 4.257 | 6.6428 | 7.0066 | 1.1711 | 0.22237 | 0.347 | 0.366 | 0.061177 | 1.5605 | 1.6459 | 0.27511 | 1.0548 | 0.1763 | 0.16715 | -1.5034 | -1.0584 | -1.0051 | -2.794 | [10 empty] | 0 | 1.821 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014086 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 98.87% | 0.14536 | 5.1915 | 0.88912
Lab 1 | Jun 23, 2022 | 6 | Zr_096 | SRM |  |  | 0.00049201 | 8.7905e-05 | 0.00014937 | -1.304e-06 | 0.00015762 | 5.8689e-07 | 7.1411e-05 | 1.2138e-05 | -1.5394e-06 | [5 empty] | 20.9726 | 4.6636 | 7.2771 | 3.6183e-05 | 7.6752 | 2.7382e-05 | 1.2829 | 9.8945e-06 | 3.4433e-06 |  |  |  | 1 | 0 | 0 | 20.9721 | 4.6635 | 7.277 | 3.7487e-05 | 7.6751 | 2.6795e-05 | 1.2828 | -2.2432e-06 | 4.9826e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8187 | 20.9721 | 4.6635 | 7.2769 | 7.6751 | 1.2828 | 0.22237 | 0.34698 | 0.36597 | 0.061168 | 1.5604 | 1.6458 | 0.27508 | 1.0547 | 0.17629 | 0.16714 | -1.5034 | -1.0585 | -1.0052 | -2.7941 | [10 empty] | 0 | 1.8187 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.01409 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 109.1% | -1.2537 | -0.027791 | 11.7485 | Bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.1% | 2.856 | 3.5409 | 4.5114 | SRM 3169 | Avg | 1.2768 | 1.4146 | 5.448
Lab 1 | Jun 23, 2022 | 6 | Zr_100 | iRM_12 |  |  | 0.00033766 | 6.797e-05 | 0.00010635 | 3.8647e-06 | 0.00011729 | -2.6846e-06 | 6.5504e-05 | 1.6651e-05 | -3.0051e-06 | [5 empty] | 19.2872 | 4.289 | 6.6928 | 3.5513e-05 | 7.0595 | 2.0763e-05 | 1.18 | 8.6813e-06 | -5.4677e-07 |  |  |  | 1 | 0 | 0 | 19.2869 | 4.2889 | 6.6927 | 3.1648e-05 | 7.0594 | 2.3448e-05 | 1.18 | -7.9702e-06 | 2.4583e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8222 | 19.2869 | 4.2889 | 6.6927 | 7.0594 | 1.18 | 0.22237 | 0.34701 | 0.36602 | 0.061181 | 1.5605 | 1.646 | 0.27512 | 1.0548 | 0.17631 | 0.16715 | -1.5034 | -1.0584 | -1.0051 | -2.7939 | [10 empty] | 0 | 1.8222 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100% | 8.3987 | -4.2598 | -1.434 | [8 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.2% | 8.6478 | -1.0155 | -6.4352 | SRM 3169 | 2SD | 10.8476 | 3.2489 | 14.1164
Lab 1 | Jun 23, 2022 | 6 | Zr_104 | alfaA |  |  | 0.00216 | 0.00047026 | 0.0007385 | -2.1264e-06 | 0.00078266 | 3.1066e-06 | 0.0001719 | 3.4195e-07 | -6.1768e-06 | [5 empty] | 21.3405 | 4.7447 | 7.4027 | 2.8416e-05 | 7.8056 | 2.2972e-05 | 1.3044 | 1.4053e-05 | 1.4381e-06 |  |  |  | 1 | 0 | 0 | 21.3383 | 4.7442 | 7.402 | 3.0542e-05 | 7.8049 | 1.9865e-05 | 1.3042 | 1.3711e-05 | 7.6149e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8062 | 21.3383 | 4.7442 | 7.402 | 7.8049 | 1.3042 | 0.22233 | 0.34689 | 0.36577 | 0.06112 | 1.5602 | 1.6451 | 0.2749 | 1.0544 | 0.1762 | 0.1671 | -1.5036 | -1.0588 | -1.0058 | -2.7949 | [10 empty] | 0 | 1.8062 | 0.21793 | 0.33337 | 0.3381 | 0.054385 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014091 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 23, 2022 | 6 | 110.3% | -1.0227 | -0.10254 | 40.0003 | bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.6% | 3.4602 | 1.4591 | 10.0732 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 23, 2022 | 6 | Zr_108 | iRM_12 |  |  | 0.00026768 | 5.229e-05 | 7.9813e-05 | -3.429e-06 | 8.4248e-05 | 1.6523e-06 | 5.9915e-05 | 5.5538e-07 | -3.0729e-06 | [5 empty] | 19.4189 | 4.3182 | 6.7386 | 3.4473e-05 | 7.108 | 2.0953e-05 | 1.1882 | 8.4183e-06 | -1.3049e-06 |  |  |  | 1 | 0 | 0 | 19.4186 | 4.3182 | 6.7385 | 3.7902e-05 | 7.1079 | 1.9301e-05 | 1.1881 | 7.863e-06 | 1.768e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.823 | 19.4186 | 4.3182 | 6.7385 | 7.1079 | 1.1881 | 0.22237 | 0.34701 | 0.36603 | 0.061183 | 1.5605 | 1.646 | 0.27514 | 1.0548 | 0.17631 | 0.16715 | -1.5034 | -1.0584 | -1.005 | -2.7939 | [10 empty] | 0 | 1.823 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.01409 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 101.1% | -14.3616 | -0.25284 | -14.8951 | [8 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.9% | 2.5111 | 2.583 | 1.7149 | SRM 3169 | 2SE | 4.4285 | 1.3264 | 5.763
Lab 1 | Jun 23, 2022 | 6 | Zr_112 | Ames |  |  | 0.00024566 | 4.5025e-05 | 7.6561e-05 | 1.0015e-05 | 8.0199e-05 | 5.586e-06 | 6.3044e-05 | 1.2255e-05 | -4.1002e-06 | [5 empty] | 21.5098 | 4.7835 | 7.4649 | 3.2801e-05 | 7.8747 | 3.3858e-05 | 1.3165 | 1.3881e-05 | 8.974e-06 |  |  |  | 1 | 0 | 0 | 21.5096 | 4.7834 | 7.4648 | 2.2786e-05 | 7.8746 | 2.8272e-05 | 1.3164 | 1.6259e-06 | 1.3074e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.827 | 21.5096 | 4.7834 | 7.4648 | 7.8746 | 1.3164 | 0.22239 | 0.34705 | 0.3661 | 0.0612 | 1.5606 | 1.6462 | 0.2752 | 1.0549 | 0.17635 | 0.16717 | -1.5033 | -1.0583 | -1.0049 | -2.7936 | [10 empty] | 0 | 1.827 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014085 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 23, 2022 | 6 | 111.6% | 5.6815 | 3.6854 | 4.7754 | bottle 5 | Lab 2 | May 28, 2020 | 1 | 105.2% | -6.8426 | 1.6624 | 10.8917 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_116 | iRM_12 |  |  | 0.00022189 | 4.1308e-05 | 7.7265e-05 | 2.9523e-06 | 8.7432e-05 | -1.7418e-06 | 5.4649e-05 | 1.0542e-05 | -2.7552e-06 | [5 empty] | 19.144 | 4.2572 | 6.6433 | 3.1453e-05 | 7.0074 | 1.8639e-05 | 1.1714 | 5.6504e-06 | -6.6e-07 |  |  |  | 1 | 0 | 0 | 19.1438 | 4.2571 | 6.6432 | 2.8501e-05 | 7.0073 | 2.038e-05 | 1.1713 | -4.8918e-06 | 2.0952e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8231 | 19.1438 | 4.2571 | 6.6432 | 7.0073 | 1.1713 | 0.22238 | 0.34702 | 0.36604 | 0.061185 | 1.5605 | 1.646 | 0.27514 | 1.0548 | 0.17632 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7939 | [10 empty] | 0 | 1.8231 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43915 | -1.3881 | 0.014087 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.2% | 8.1632 | 4.0272 | 15.9972 | [8 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 105% | -2.9715 | 0.25743 | 11.9319 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_120 | IPGP |  |  | 0.00023021 | 4.7415e-05 | 7.3877e-05 | -2.1804e-06 | 8.8513e-05 | 1.3084e-06 | 6.4174e-05 | 9.0577e-06 | 6.1157e-07 | [5 empty] | 21.2245 | 4.7197 | 7.3649 | 2.6527e-05 | 7.7681 | 2.738e-05 | 1.2985 | 1.4299e-05 | 2.027e-06 |  |  |  | 1 | 0 | 0 | 21.2242 | 4.7196 | 7.3648 | 2.8707e-05 | 7.768 | 2.6071e-05 | 1.2984 | 5.2416e-06 | 1.4154e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8207 | 21.2242 | 4.7196 | 7.3648 | 7.768 | 1.2984 | 0.22237 | 0.347 | 0.366 | 0.061176 | 1.5605 | 1.6459 | 0.27511 | 1.0548 | 0.1763 | 0.16715 | -1.5034 | -1.0584 | -1.0051 | -2.794 | [10 empty] | 0 | 1.8207 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3881 | 0.014083 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 23, 2022 | 6 | 110.8% | -6.1587 | 6.1772 | 9.0001
Lab 1 | Jun 23, 2022 | 6 | Zr_124 | iRM_12 |  |  | 0.00023652 | 3.8621e-05 | 8.7641e-05 | -1.2019e-06 | 9.9022e-05 | -1.9375e-06 | 5.6896e-05 | 1.4942e-05 | -1.8547e-06 | [5 empty] | 19.1788 | 4.265 | 6.6556 | 3.5128e-05 | 7.0206 | 2.0235e-05 | 1.1736 | 6.1358e-06 | 6.487e-08 |  |  |  | 1 | 0 | 0 | 19.1786 | 4.2649 | 6.6555 | 3.633e-05 | 7.0205 | 2.2172e-05 | 1.1736 | -8.8063e-06 | 1.9196e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8247 | 19.1786 | 4.2649 | 6.6555 | 7.0205 | 1.1736 | 0.22238 | 0.34703 | 0.36606 | 0.061191 | 1.5605 | 1.6461 | 0.27516 | 1.0548 | 0.17633 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7938 | [10 empty] | 0 | 1.8246 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43916 | -1.3881 | 0.014092 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.84% | -0.7186 | -5.4824 | -8.9986 | [8 empty] | Bottle 5 | Lab 2 | June 3, 2020 | 2 | 104.2% | -4.9136 | -9.3449 | 34.5917 | SRM 3169 | Avg | -4.754 | -0.76649 | 16.2937
Lab 1 | Jun 23, 2022 | 6 | Zr_128 | iRM_UR |  |  | 0.00023079 | 4.4918e-05 | 7.5109e-05 | 2.9193e-06 | 8.2506e-05 | 2.4986e-06 | 5.6242e-05 | 4.5203e-06 | -3.7428e-06 | [5 empty] | 18.8817 | 4.1988 | 6.5522 | 1.5576e-05 | 6.9113 | 1.982e-05 | 1.1553 | 1.2985e-05 | -2.9634e-06 |  |  |  | 1 | 0 | 0 | 18.8815 | 4.1987 | 6.5521 | 1.2656e-05 | 6.9112 | 1.7321e-05 | 1.1553 | 8.4643e-06 | 7.794e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8228 | 18.8815 | 4.1987 | 6.5521 | 6.9112 | 1.1553 | 0.22237 | 0.34701 | 0.36603 | 0.061185 | 1.5605 | 1.646 | 0.27514 | 1.0548 | 0.17632 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7939 | [10 empty] | 0 | 1.8228 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3881 | 0.014089 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 23, 2022 | 6 | 98.2% | -8.5981 | 0.37901 | 12.6714 | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.4% | -12.1061 | -1.5531 | 31.6137 | SRM 3169 | 2SD | 9.4709 | 10.3456 | 33.0358
Lab 1 | Jun 23, 2022 | 6 | Zr_132 | iRM_12 |  |  | 0.00024159 | 5.1844e-05 | 7.525e-05 | -9.2129e-06 | 9.0357e-05 | 1.6081e-06 | 5.5805e-05 | 7.9781e-06 | -3.8394e-06 | [5 empty] | 19.2757 | 4.2866 | 6.6894 | 2.9149e-05 | 7.0565 | 2.0846e-05 | 1.1797 | 1.5322e-05 | 7.2032e-07 |  |  |  | 1 | 0 | 0 | 19.2754 | 4.2865 | 6.6893 | 3.8362e-05 | 7.0564 | 1.9238e-05 | 1.1796 | 7.3442e-06 | 4.5597e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.826 | 19.2754 | 4.2865 | 6.6893 | 7.0564 | 1.1796 | 0.22238 | 0.34704 | 0.36608 | 0.061197 | 1.5605 | 1.6462 | 0.27519 | 1.0549 | 0.17634 | 0.16717 | -1.5034 | -1.0583 | -1.0049 | -2.7937 | [10 empty] | 0 | 1.826 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014086 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.7% | 0.04127 | 3.273 | 6.8818 | [8 empty] | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.6% | -5.7717 | 2.0264 | 14.2531 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 23, 2022 | 6 | Zr_136 | SRM |  |  | 0.0003031 | 5.3674e-05 | 0.00010296 | 4.4259e-06 | 9.9238e-05 | 1.2991e-06 | 5.8962e-05 | 1.7044e-05 | -3.4514e-06 | [5 empty] | 20.9474 | 4.6581 | 7.2689 | 4.1607e-05 | 7.667 | 2.3005e-05 | 1.2816 | 1.6814e-05 | 1.3699e-06 |  |  |  | 1 | 0 | 0 | 20.9471 | 4.658 | 7.2688 | 3.7181e-05 | 7.6669 | 2.1706e-05 | 1.2815 | -2.3061e-07 | 4.8213e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8217 | 20.9471 | 4.658 | 7.2688 | 7.6669 | 1.2815 | 0.22237 | 0.34701 | 0.36601 | 0.06118 | 1.5605 | 1.646 | 0.27512 | 1.0548 | 0.17631 | 0.16715 | -1.5034 | -1.0584 | -1.0051 | -2.7939 | [10 empty] | 0 | 1.8217 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3881 | 0.014088 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 109.1% | -6.0461 | -1.3364 | 1.8274 | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.6% | -5.7039 | -1.4478 | 18.6152 | SRM 3169 | 2SE | 3.8665 | 4.2236 | 13.4868
Lab 1 | Jun 23, 2022 | 6 | Zr_140 | iRM_12 |  |  | 0.00024256 | 4.8859e-05 | 7.7221e-05 | 5.0683e-06 | 8.2578e-05 | 1.5461e-06 | 5.0686e-05 | 1.2656e-05 | -3.0857e-06 | [5 empty] | 19.1176 | 4.2514 | 6.6345 | 3.2622e-05 | 6.9984 | 1.8394e-05 | 1.1699 | 1.1633e-05 | -2.0187e-06 |  |  |  | 1 | 0 | 0 | 19.1174 | 4.2514 | 6.6344 | 2.7554e-05 | 6.9984 | 1.6848e-05 | 1.1699 | -1.0233e-06 | 1.067e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8254 | 19.1174 | 4.2514 | 6.6344 | 6.9984 | 1.1699 | 0.22238 | 0.34703 | 0.36607 | 0.061194 | 1.5605 | 1.6461 | 0.27518 | 1.0549 | 0.17633 | 0.16716 | -1.5034 | -1.0583 | -1.0049 | -2.7937 | [10 empty] | 0 | 1.8253 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014087 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.77% | 0.77672 | 0.55293 | -1.6803 | [8 empty] | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.6% | 2.3497 | 6.3966 | -10.9165 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_144 | alfaA |  |  | 0.0013578 | 0.00029394 | 0.00046591 | -4.9059e-06 | 0.00050049 | 1.8258e-06 | 0.00012621 | 7.9779e-06 | -7.8771e-06 | [5 empty] | 21.7034 | 4.8256 | 7.5292 | 3.0913e-05 | 7.9397 | 2.5266e-05 | 1.3269 | 1.0111e-05 | -1.8261e-08 |  |  |  | 1 | 0 | 0 | 21.702 | 4.8253 | 7.5287 | 3.5819e-05 | 7.9392 | 2.344e-05 | 1.3267 | 2.1331e-06 | 7.8589e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8099 | 21.702 | 4.8253 | 7.5287 | 7.9392 | 1.3267 | 0.22234 | 0.34691 | 0.36583 | 0.061134 | 1.5603 | 1.6453 | 0.27495 | 1.0545 | 0.17622 | 0.16711 | -1.5035 | -1.0587 | -1.0056 | -2.7947 | [10 empty] | 0 | 1.8099 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43917 | -1.3881 | 0.014091 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 23, 2022 | 6 | 113.7% | -7.4267 | -3.7472 | 17.5265 | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.9% | -2.3783 | -0.67613 | 9.6052 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_148 | iRM_12 |  |  | 0.00023035 | 3.7814e-05 | 7.6399e-05 | -2.4686e-06 | 8.9704e-05 | 7.1708e-07 | 6.936e-05 | 6.9102e-06 | -2.7254e-06 | [5 empty] | 19.0476 | 4.2359 | 6.6103 | 3.2716e-05 | 6.9731 | 2.0254e-05 | 1.1657 | 1.353e-05 | -2.85e-06 |  |  |  | 1 | 0 | 0 | 19.0474 | 4.2358 | 6.6102 | 3.5184e-05 | 6.9731 | 1.9537e-05 | 1.1657 | 6.6201e-06 | -1.2456e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8264 | 19.0474 | 4.2358 | 6.6102 | 6.9731 | 1.1657 | 0.22238 | 0.34704 | 0.36609 | 0.061198 | 1.5606 | 1.6462 | 0.27519 | 1.0549 | 0.17634 | 0.16717 | -1.5034 | -1.0583 | -1.0049 | -2.7936 | [10 empty] | 0 | 1.8264 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014089 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.36% | 5.9631 | 2.5952 | 1.4463
Lab 1 | Jun 23, 2022 | 6 | Zr_152 | Ames |  |  | 0.0024133 | 0.00054087 | 0.00082589 | 1.8027e-06 | 0.00087674 | 1.0804e-06 | 0.00018579 | 4.9999e-06 | -4.5795e-06 | [5 empty] | 21.5851 | 4.8003 | 7.4914 | 3.4501e-05 | 7.9032 | 3.3217e-05 | 1.3213 | 1.4984e-05 | 9.8952e-06 |  |  |  | 1 | 0 | 0 | 21.5827 | 4.7998 | 7.4906 | 3.2699e-05 | 7.9023 | 3.2137e-05 | 1.3211 | 9.9838e-06 | 1.4475e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8297 | 21.5827 | 4.7998 | 7.4906 | 7.9023 | 1.3211 | 0.22239 | 0.34706 | 0.36614 | 0.061211 | 1.5606 | 1.6464 | 0.27524 | 1.055 | 0.17637 | 0.16718 | -1.5033 | -1.0582 | -1.0047 | -2.7934 | [10 empty] | 0 | 1.8297 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43917 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 23, 2022 | 6 | 112.8% | 1.5796 | -0.93415 | 4.7803 | Bottle 6 | Lab 2 | May 28, 2020 | 1 | 102.3% | 6.0491 | 4.6312 | 10.1566 | SRM 3169 | Avg | -1.6342 | -0.45386 | 8.38
Lab 1 | Jun 23, 2022 | 6 | Zr_156 | iRM_12 |  |  | 0.00026272 | 6.0385e-05 | 8.4483e-05 | 3.4625e-06 | 9.7873e-05 | 2.2668e-06 | 5.8863e-05 | 1.2813e-06 | -1.2827e-06 | [5 empty] | 19.2223 | 4.2747 | 6.6711 | 3.5536e-05 | 7.0376 | 2.1868e-05 | 1.1765 | 1.29e-05 | -8.1992e-07 |  |  |  | 1 | 0 | 0 | 19.2221 | 4.2747 | 6.671 | 3.2073e-05 | 7.0375 | 1.9601e-05 | 1.1765 | 1.1619e-05 | 4.6278e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.828 | 19.2221 | 4.2747 | 6.671 | 7.0375 | 1.1765 | 0.22238 | 0.34705 | 0.36611 | 0.061204 | 1.5606 | 1.6463 | 0.27522 | 1.0549 | 0.17635 | 0.16717 | -1.5033 | -1.0583 | -1.0048 | -2.7935 | [10 empty] | 0 | 1.828 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.014096 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 101.3% | -14.68 | -11.1745 | 0.082121 | [8 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 102.8% | -0.38789 | -4.1636 | 1.88 | SRM 3169 | 2SD | 13.2823 | 12.4794 | 39.3828
Lab 1 | Jun 23, 2022 | 6 | Zr_160 | IPGP |  |  | 0.00023009 | 4.41e-05 | 7.1827e-05 | 1.1819e-05 | 8.1231e-05 | 2.3662e-06 | 5.2957e-05 | 1.6393e-05 | -4.7557e-06 | [5 empty] | 20.9808 | 4.6656 | 7.2807 | 2.379e-05 | 7.6797 | 2.8463e-05 | 1.2837 | 1.8341e-05 | 2.3044e-06 |  |  |  | 1 | 0 | 0 | 20.9806 | 4.6655 | 7.2806 | 1.1971e-05 | 7.6796 | 2.6097e-05 | 1.2837 | 1.9478e-06 | 7.0602e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8229 | 20.9806 | 4.6655 | 7.2806 | 7.6796 | 1.2837 | 0.22237 | 0.34702 | 0.36603 | 0.061184 | 1.5605 | 1.646 | 0.27514 | 1.0548 | 0.17632 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7939 | [10 empty] | 0 | 1.8229 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.014087 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 23, 2022 | 6 | 110.1% | 0.19591 | 1.0249 | 12.1395 | bottle 6 | Lab 2 | May 28, 2020 | 1 | 101.7% | -13.1806 | -10.4252 | 41.589 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 23, 2022 | 6 | Zr_164 | iRM_12 |  |  | 0.00023958 | 4.6379e-05 | 7.4625e-05 | 4.0568e-06 | 9.2455e-05 | -3.6013e-06 | 6.3424e-05 | 1.5956e-05 | -2.2092e-06 | [5 empty] | 18.8945 | 4.2019 | 6.5573 | 3.5987e-05 | 6.9173 | 2.004e-05 | 1.1564 | 1.0415e-05 | 4.1233e-07 |  |  |  | 1 | 0 | 0 | 18.8942 | 4.2018 | 6.5572 | 3.193e-05 | 6.9172 | 2.3641e-05 | 1.1563 | -5.5408e-06 | 2.6215e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8272 | 18.8942 | 4.2018 | 6.5572 | 6.9172 | 1.1563 | 0.22239 | 0.34705 | 0.3661 | 0.0612 | 1.5606 | 1.6462 | 0.2752 | 1.0549 | 0.17634 | 0.16717 | -1.5033 | -1.0583 | -1.0048 | -2.7936 | [10 empty] | 0 | 1.8272 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43916 | -1.3881 | 0.01408 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.11% | 4.0466 | 12.6723 | -21.0345 | [8 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 103.5% | 2.4295 | -2.0216 | -19.7319 | SRM 3169 | 2SE | 5.4225 | 5.0947 | 16.078
Lab 1 | Jun 23, 2022 | 6 | Zr_168 | iRM_UR |  |  | 0.00025879 | 4.9713e-05 | 8.2801e-05 | -3.198e-06 | 8.2012e-05 | -2.3129e-06 | 5.3976e-05 | 1.2146e-05 | -3.0285e-06 | [5 empty] | 18.8721 | 4.1969 | 6.5495 | 1.6667e-05 | 6.9092 | 1.8686e-05 | 1.1551 | 1.3723e-05 | -3.3828e-06 |  |  |  | 1 | 0 | 0 | 18.8718 | 4.1968 | 6.5494 | 1.9865e-05 | 6.9091 | 2.0999e-05 | 1.155 | 1.5766e-06 | -3.5424e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8275 | 18.8718 | 4.1968 | 6.5494 | 6.9091 | 1.155 | 0.22239 | 0.34705 | 0.36611 | 0.061203 | 1.5606 | 1.6463 | 0.27521 | 1.0549 | 0.17635 | 0.16717 | -1.5033 | -1.0583 | -1.0048 | -2.7936 | [10 empty] | 0 | 1.8275 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43916 | -1.3881 | 0.014095 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 23, 2022 | 6 | 99.85% | -3.9436 | -9.7305 | 9.2082 | bottle 6 | Lab 2 | May 28, 2020 | 1 | 104.3% | -4.4856 | 3.7589 | 8.4772 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_172 | iRM_12 |  |  | 0.00036815 | 6.4466e-05 | 0.00011941 | 1.7306e-06 | 0.00013135 | 1.0408e-07 | 5.3265e-05 | 1.7188e-06 | 2.5499e-07 | [5 empty] | 18.905 | 4.2043 | 6.5612 | 3.9391e-05 | 6.9217 | 2.132e-05 | 1.1572 | 2.1384e-05 | 4.363e-07 |  |  |  | 1 | 0 | 0 | 18.9047 | 4.2042 | 6.5611 | 3.7661e-05 | 6.9216 | 2.1216e-05 | 1.1572 | 1.9665e-05 | 1.8131e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.829 | 18.9047 | 4.2042 | 6.5611 | 6.9216 | 1.1572 | 0.22239 | 0.34706 | 0.36613 | 0.06121 | 1.5606 | 1.6463 | 0.27523 | 1.0549 | 0.17637 | 0.16718 | -1.5033 | -1.0583 | -1.0048 | -2.7934 | [10 empty] | 0 | 1.829 | 0.21793 | 0.33337 | 0.3381 | 0.054385 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43915 | -1.3881 | 0.01409 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100% | 7.1432 | -5.1927 | 36.9643 | [8 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 104.4% | -0.2295 | 5.4971 | 7.9089 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_176 | SRM |  |  | 0.00023062 | 4.6025e-05 | 8.0659e-05 | -2.9789e-06 | 7.5864e-05 | -1.6037e-06 | 4.5419e-05 | 2.8686e-05 | -2.842e-06 | [5 empty] | 20.7217 | 4.608 | 7.1909 | 4.1885e-05 | 7.5852 | 2.7583e-05 | 1.268 | 1.6673e-05 | 1.3401e-06 |  |  |  | 1 | 0 | 0 | 20.7214 | 4.608 | 7.1908 | 4.4864e-05 | 7.5852 | 2.9186e-05 | 1.2679 | -1.2013e-05 | 4.1821e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8242 | 20.7214 | 4.608 | 7.1908 | 7.5852 | 1.2679 | 0.22238 | 0.34702 | 0.36605 | 0.06119 | 1.5605 | 1.6461 | 0.27516 | 1.0548 | 0.17633 | 0.16716 | -1.5034 | -1.0584 | -1.005 | -2.7938 | [10 empty] | 0 | 1.8242 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43917 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 109.6% | -9.932 | -3.6404 | 3.7467
Lab 1 | Jun 23, 2022 | 6 | Zr_180 | iRM_12 |  |  | 0.00024753 | 4.9927e-05 | 7.5327e-05 | -8.686e-07 | 8.2347e-05 | -1.0266e-06 | 4.8166e-05 | 1.1031e-05 | -2.2494e-06 | [5 empty] | 18.9154 | 4.2066 | 6.5649 | 3.5931e-05 | 6.9258 | 2.0982e-05 | 1.1579 | 1.3232e-05 | -1.1099e-06 |  |  |  | 1 | 0 | 0 | 18.9151 | 4.2066 | 6.5649 | 3.6799e-05 | 6.9257 | 2.2008e-05 | 1.1578 | 2.2002e-06 | 1.1395e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8301 | 18.9151 | 4.2066 | 6.5649 | 6.9257 | 1.1578 | 0.22239 | 0.34707 | 0.36615 | 0.061212 | 1.5606 | 1.6464 | 0.27524 | 1.055 | 0.17637 | 0.16718 | -1.5033 | -1.0582 | -1.0047 | -2.7934 | [10 empty] | 0 | 1.8301 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.01409 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.5% | -3.0949 | 1.9911 | -22.0627 | [8 empty] | Bottle 6 | Lab 2 | June 3, 2020 | 2 | 106% | -5.0128 | -2.1319 | 0.054605 | SRM 3169 | Avg | -3.1847 | 1.623 | 1.7262
Lab 1 | Jun 23, 2022 | 6 | Zr_184 | alfaA |  |  | 0.00022759 | 5.0439e-05 | 6.886e-05 | 6.251e-06 | 8.5643e-05 | -6.9184e-08 | 5.1146e-05 | 3.8988e-06 | -5.3075e-06 | [5 empty] | 21.5881 | 4.8001 | 7.4897 | 2.8181e-05 | 7.8986 | 2.6945e-05 | 1.3201 | 1.5654e-05 | 3.1052e-06 |  |  |  | 1 | 0 | 0 | 21.5878 | 4.8 | 7.4897 | 2.193e-05 | 7.8985 | 2.7014e-05 | 1.32 | 1.1755e-05 | 8.4127e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8131 | 21.5878 | 4.8 | 7.4897 | 7.8985 | 1.32 | 0.22235 | 0.34694 | 0.36588 | 0.061147 | 1.5603 | 1.6455 | 0.27501 | 1.0546 | 0.17625 | 0.16712 | -1.5035 | -1.0586 | -1.0055 | -2.7945 | [10 empty] | 0 | 1.8131 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 23, 2022 | 6 | 114.7% | -14.5602 | -1.1968 | 29.3284 | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.3% | -7.5427 | 5.5312 | 9.4619 | SRM 3169 | 2SD | 7.6768 | 6.8152 | 21.9152
Lab 1 | Jun 23, 2022 | 6 | Zr_188 | iRM_12 |  |  | 0.00023323 | 4.7609e-05 | 7.6944e-05 | 5.0263e-06 | 8.282e-05 | 3.1445e-06 | 5.3988e-05 | 4.3821e-06 | 8.7534e-07 | [5 empty] | 18.7269 | 4.1647 | 6.4995 | 3.8961e-05 | 6.8567 | 2.1253e-05 | 1.1463 | 1.0265e-05 | -7.7833e-07 |  |  |  | 1 | 0 | 0 | 18.7267 | 4.1647 | 6.4994 | 3.3935e-05 | 6.8566 | 1.8108e-05 | 1.1463 | 5.8831e-06 | -1.6537e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8297 | 18.7267 | 4.1647 | 6.4994 | 6.8566 | 1.1463 | 0.22239 | 0.34706 | 0.36614 | 0.061211 | 1.5606 | 1.6464 | 0.27524 | 1.055 | 0.17637 | 0.16718 | -1.5033 | -1.0582 | -1.0047 | -2.7934 | [10 empty] | 0 | 1.8297 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.27% | 1.0287 | -1.2047 | 2.2395 | [8 empty] | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.5% | -3.8999 | 4.5634 | 19.0303 | SRM 3169 | n | 6 | 6 | 6
Lab 1 | Jun 23, 2022 | 6 | Zr_192 | Ames |  |  | 0.0002898 | 5.91e-05 | 8.8227e-05 | 1.335e-05 | 0.00010319 | -3.0145e-06 | 4.7072e-05 | 1.3064e-05 | -2.2123e-06 | [5 empty] | 21.7104 | 4.8284 | 7.5355 | 3.1349e-05 | 7.9504 | 3.2533e-05 | 1.3293 | 1.7326e-05 | 1.0265e-05 |  |  |  | 1 | 0 | 0 | 21.7101 | 4.8284 | 7.5354 | 1.8e-05 | 7.9503 | 3.5547e-05 | 1.3292 | 4.2614e-06 | 1.2478e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8334 | 21.7101 | 4.8284 | 7.5354 | 7.9503 | 1.3292 | 0.2224 | 0.34709 | 0.3662 | 0.061226 | 1.5607 | 1.6466 | 0.27529 | 1.0551 | 0.1764 | 0.16719 | -1.5033 | -1.0582 | -1.0046 | -2.7932 | [10 empty] | 0 | 1.8334 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43916 | -1.3881 | 0.014094 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 23, 2022 | 6 | 115.7% | 3.8837 | 0.14868 | -1.8534 | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.4% | -2.2142 | 2.4611 | -5.1259 | SRM 3169 | 2SE | 3.134 | 2.7823 | 8.9469
Lab 1 | Jun 23, 2022 | 6 | Zr_196 | iRM_12 |  |  | 0.0042754 | 0.00093945 | 0.001482 | 3.5195e-06 | 0.0015515 | -4.2463e-07 | 0.00029736 | 1.2069e-05 | -1.608e-06 | [5 empty] | 18.8163 | 4.1846 | 6.5307 | 3.2341e-05 | 6.89 | 1.8747e-05 | 1.1519 | 1.0399e-05 | 2.0754e-07 |  |  |  | 1 | 0 | 0 | 18.812 | 4.1837 | 6.5293 | 2.8821e-05 | 6.8884 | 1.9172e-05 | 1.1516 | -1.6691e-06 | 1.8156e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8315 | 18.812 | 4.1837 | 6.5293 | 6.8884 | 1.1516 | 0.2224 | 0.34708 | 0.36617 | 0.061218 | 1.5606 | 1.6465 | 0.27527 | 1.055 | 0.17638 | 0.16719 | -1.5033 | -1.0582 | -1.0047 | -2.7933 | [10 empty] | 0 | 1.8315 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014094 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.3% | -2.5245 | -2.4203 | 2.7743 | [8 empty] | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.8% | 3.7989 | 2.0662 | -1.2906 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_200 | IPGP |  |  | 0.0013997 | 0.0002996 | 0.00048495 | -3.9152e-07 | 0.00050687 | -3.3728e-07 | 0.00010967 | 3.0686e-06 | -1.0138e-06 | [5 empty] | 20.8761 | 4.6424 | 7.2447 | 2.5632e-05 | 7.6422 | 2.5108e-05 | 1.2775 | 1.0627e-05 | 1.7699e-06 |  |  |  | 1 | 0 | 0 | 20.8747 | 4.6421 | 7.2442 | 2.6023e-05 | 7.6417 | 2.5446e-05 | 1.2774 | 7.5582e-06 | 2.7837e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8253 | 20.8747 | 4.6421 | 7.2442 | 7.6417 | 1.2774 | 0.22238 | 0.34703 | 0.36607 | 0.061195 | 1.5605 | 1.6462 | 0.27518 | 1.0549 | 0.17634 | 0.16717 | -1.5034 | -1.0583 | -1.0049 | -2.7937 | [10 empty] | 0 | 1.8253 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43917 | -1.3881 | 0.014097 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 23, 2022 | 6 | 111% | -9.4556 | -4.613 | 15.3745 | bottle 6 | Lab 2 | June 3, 2020 | 2 | 107% | -4.2375 | -2.752 | -11.7731 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_204 | iRM_12 |  |  | 0.00023836 | 3.3704e-05 | 8.4421e-05 | 2.6281e-06 | 8.4281e-05 | -2.933e-06 | 4.938e-05 | 1.4708e-05 | 3.3524e-07 | [5 empty] | 18.7941 | 4.1797 | 6.5231 | 3.6089e-05 | 6.8819 | 1.9572e-05 | 1.1506 | 1.2098e-05 | 4.8342e-07 |  |  |  | 1 | 0 | 0 | 18.7938 | 4.1797 | 6.523 | 3.3461e-05 | 6.8818 | 2.2505e-05 | 1.1505 | -2.6096e-06 | 1.4818e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8318 | 18.7938 | 4.1797 | 6.523 | 6.8818 | 1.1505 | 0.2224 | 0.34708 | 0.36618 | 0.061219 | 1.5606 | 1.6465 | 0.27527 | 1.055 | 0.17638 | 0.16719 | -1.5033 | -1.0582 | -1.0046 | -2.7933 | [10 empty] | 0 | 1.8318 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014091 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.6% | 3.8096 | 4.5273 | -9.4701
Lab 1 | Jun 23, 2022 | 6 | Zr_208 | iRM_UR |  |  | 0.0018243 | 0.0003974 | 0.00063155 | 9.8023e-07 | 0.0006544 | 1.8902e-06 | 0.00014482 | 7.7894e-06 | -3.4004e-06 | [5 empty] | 18.563 | 4.1283 | 6.4429 | 2.0356e-05 | 6.7972 | 1.818e-05 | 1.1364 | 1.0767e-05 | -8.8328e-07 |  |  |  | 1 | 0 | 0 | 18.5612 | 4.1279 | 6.4422 | 1.9375e-05 | 6.7966 | 1.629e-05 | 1.1363 | 2.9772e-06 | 2.5171e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8316 | 18.5612 | 4.1279 | 6.4422 | 6.7966 | 1.1363 | 0.2224 | 0.34708 | 0.36617 | 0.061218 | 1.5606 | 1.6465 | 0.27527 | 1.055 | 0.17638 | 0.16719 | -1.5033 | -1.0582 | -1.0047 | -2.7933 | [10 empty] | 0 | 1.8316 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014091 | -1.8132 | -1.8273 | [14 empty] | iRM_UR | Lab 1 | Jun 23, 2022 | 6 | 99.43% | 2.8079 | 2.4614 | -7.8991 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.1% | 2.5461 | -1.8685 | 6.0688 | SRM 3169 | Avg | 0.067358 | 0.031281 | 0.0080002
Lab 1 | Jun 23, 2022 | 6 | Zr_212 | iRM_12 |  |  | 0.00053042 | 0.00011157 | 0.00019071 | 2.9464e-06 | 0.00018808 | -8.8128e-07 | 6.2964e-05 | 1.2854e-05 | -9.9981e-07 | [5 empty] | 18.5416 | 4.1236 | 6.4354 | 3.5471e-05 | 6.7894 | 1.9518e-05 | 1.1351 | 1.3034e-05 | -1.839e-06 |  |  |  | 1 | 0 | 0 | 18.5411 | 4.1234 | 6.4352 | 3.2525e-05 | 6.7892 | 2.0399e-05 | 1.1351 | 1.7946e-07 | -8.3917e-07 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8317 | 18.5411 | 4.1234 | 6.4352 | 6.7892 | 1.1351 | 0.22239 | 0.34708 | 0.36617 | 0.06122 | 1.5606 | 1.6465 | 0.27528 | 1.055 | 0.17639 | 0.16719 | -1.5033 | -1.0582 | -1.0046 | -2.7933 | [10 empty] | 0 | 1.8317 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43917 | -1.3881 | 0.014096 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 98.57% | -1.5988 | -3.5245 | 20.7172 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.77% | -0.14288 | 1.7659 | 5.3228 | SRM 3169 | 2SD | 8.4211 | 6.6273 | 26.2256
Lab 1 | Jun 23, 2022 | 6 | Zr_216 | SRM |  |  | 0.00023205 | 6.2739e-05 | 8.6197e-05 | 4.225e-06 | 8.2135e-05 | 2.335e-06 | 4.8026e-05 | 1.4141e-05 | -1.1158e-06 | [5 empty] | 20.9766 | 4.6649 | 7.28 | 3.4428e-05 | 7.6799 | 2.7907e-05 | 1.2839 | 1.6211e-05 | 4.4881e-06 |  |  |  | 1 | 0 | 0 | 20.9764 | 4.6649 | 7.2799 | 3.0203e-05 | 7.6798 | 2.5572e-05 | 1.2839 | 2.0702e-06 | 5.6038e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8281 | 20.9764 | 4.6649 | 7.2799 | 7.6798 | 1.2839 | 0.22239 | 0.34705 | 0.36612 | 0.061206 | 1.5606 | 1.6463 | 0.27523 | 1.0549 | 0.17636 | 0.16718 | -1.5033 | -1.0583 | -1.0048 | -2.7935 | [10 empty] | 0 | 1.8281 | 0.21793 | 0.33337 | 0.3381 | 0.054385 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.01409 | -1.8132 | -1.8273 | [14 empty] | SRM | Lab 1 | Jun 23, 2022 | 6 | 112.3% | -0.25803 | 4.6887 | 26.2398 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.1% | -3.9202 | -1.1812 | -7.3092 | SRM 3169 | n | 37 | 37 | 37
Lab 1 | Jun 23, 2022 | 6 | Zr_220 | iRM_12 |  |  | 0.0002383 | 6.1849e-05 | 7.6081e-05 | 1.8657e-06 | 8.1844e-05 | -1.8511e-06 | 5.2299e-05 | 1.555e-05 | -2.5054e-06 | [5 empty] | 18.8259 | 4.1869 | 6.5344 | 2.9468e-05 | 6.8943 | 2.0701e-05 | 1.1527 | 1.418e-05 | 1.0785e-06 |  |  |  | 1 | 0 | 0 | 18.8256 | 4.1868 | 6.5344 | 2.7603e-05 | 6.8942 | 2.2552e-05 | 1.1527 | -1.37e-06 | 3.5839e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8342 | 18.8256 | 4.1868 | 6.5344 | 6.8942 | 1.1527 | 0.2224 | 0.3471 | 0.36621 | 0.061228 | 1.5607 | 1.6466 | 0.27531 | 1.0551 | 0.1764 | 0.16719 | -1.5033 | -1.0581 | -1.0045 | -2.7931 | [10 empty] | 0 | 1.8342 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.014094 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.7% | -4.1566 | -1.2293 | -22.5064 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.95% | 4.102 | 4.102 | -9.8968 | SRM 3169 | 2SE | 1.3844 | 1.0895 | 4.3115
Lab 1 | Jun 23, 2022 | 6 | Zr_224 | alfaA |  |  | 0.0018437 | 0.00039942 | 0.00063869 | -3.4562e-06 | 0.00066879 | -9.1734e-07 | 0.00015735 | 7.1954e-06 | -2.516e-06 | [5 empty] | 21.8563 | 4.8599 | 7.5833 | 2.7587e-05 | 7.9978 | 2.6076e-05 | 1.3367 | 2.1343e-05 | 2.4421e-06 |  |  |  | 1 | 0 | 0 | 21.8544 | 4.8595 | 7.5827 | 3.1043e-05 | 7.9971 | 2.6994e-05 | 1.3366 | 1.4148e-05 | 4.9581e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8161 | 21.8544 | 4.8595 | 7.5827 | 7.9971 | 1.3366 | 0.22236 | 0.34696 | 0.36593 | 0.061158 | 1.5604 | 1.6457 | 0.27504 | 1.0547 | 0.17627 | 0.16713 | -1.5035 | -1.0585 | -1.0053 | -2.7943 | [10 empty] | 0 | 1.8161 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [14 empty] | alfaA | Lab 1 | Jun 23, 2022 | 6 | 116% | 4.5396 | -0.64549 | 5.2382 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.75% | 0.95831 | -5.5315 | 24.2923 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_228 | iRM_12 |  |  | 0.00023118 | 4.8225e-05 | 8.2579e-05 | -1.37e-05 | 8.7819e-05 | -4.9173e-06 | 5.0285e-05 | 3.5753e-06 | -2.8017e-06 | [5 empty] | 18.8487 | 4.192 | 6.5425 | 3.2342e-05 | 6.9029 | 1.918e-05 | 1.1542 | 1.7029e-05 | 1.9414e-06 |  |  |  | 1 | 0 | 0 | 18.8485 | 4.192 | 6.5424 | 4.6043e-05 | 6.9028 | 2.4098e-05 | 1.1541 | 1.3454e-05 | 4.7431e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.835 | 18.8485 | 4.192 | 6.5424 | 6.9028 | 1.1541 | 0.2224 | 0.34711 | 0.36623 | 0.061233 | 1.5607 | 1.6467 | 0.27532 | 1.0551 | 0.17641 | 0.1672 | -1.5033 | -1.0581 | -1.0045 | -2.7931 | [10 empty] | 0 | 1.835 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43916 | -1.3881 | 0.01409 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 100.5% | 4.7905 | 3.3199 | 2.7836 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100% | 0.40541 | 2.1912 | -11.9507 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_232 | Ames |  |  | 0.00053761 | 0.00010228 | 0.00018629 | -1.8893e-06 | 0.00018984 | 2.9657e-07 | 6.3003e-05 | -2.0398e-06 | -1.4029e-06 | [5 empty] | 22.0107 | 4.8955 | 7.6407 | 2.912e-05 | 8.0623 | 3.2528e-05 | 1.3481 | 1.6673e-05 | 9.7744e-06 |  |  |  | 1 | 0 | 0 | 22.0102 | 4.8954 | 7.6405 | 3.1009e-05 | 8.0621 | 3.2231e-05 | 1.3481 | 1.8713e-05 | 1.1177e-05 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.839 | 22.0102 | 4.8954 | 7.6405 | 8.0621 | 1.3481 | 0.22242 | 0.34714 | 0.36629 | 0.061247 | 1.5608 | 1.6469 | 0.27537 | 1.0552 | 0.17644 | 0.16721 | -1.5032 | -1.058 | -1.0043 | -2.7928 | [10 empty] | 0 | 1.839 | 0.21793 | 0.33337 | 0.3381 | 0.054382 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42506 | 0.43916 | -1.3881 | 0.014094 | -1.8132 | -1.8273 | [14 empty] | Ames | Lab 1 | Jun 23, 2022 | 6 | 117.2% | 8.9403 | -2.5565 | -30.5683 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.95% | -5.9338 | 0.048176 | 3.2879 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_236 | iRM_12 |  |  | 0.00034712 | 6.7558e-05 | 0.00012327 | -4.3652e-07 | 0.00012859 | -7.5303e-06 | 5.2709e-05 | 2.2917e-06 | -5.0503e-06 | [5 empty] | 18.7006 | 4.1591 | 6.4912 | 3.4401e-05 | 6.8488 | 2.0449e-05 | 1.1452 | 1.1926e-05 | 2.8287e-07 |  |  |  | 1 | 0 | 0 | 18.7003 | 4.159 | 6.491 | 3.4838e-05 | 6.8487 | 2.798e-05 | 1.1451 | 9.6348e-06 | 5.3332e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8355 | 18.7003 | 4.159 | 6.491 | 6.8487 | 1.1451 | 0.2224 | 0.34711 | 0.36623 | 0.061235 | 1.5607 | 1.6467 | 0.27533 | 1.0551 | 0.17642 | 0.1672 | -1.5033 | -1.0581 | -1.0045 | -2.793 | [10 empty] | 0 | 1.8355 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.014093 | -1.8132 | -1.8273 | [6 empty] | iRM_12 | Lab 1 | Jun 23, 2022 | 6 | 99.39% | 0.00084539 | 0.13301 | 21.1027 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | 0.46844 | 0.063249 | -1.5074 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_240 | IPGP |  |  | 0.00086176 | 0.00018906 | 0.00029014 | 1.4067e-05 | 0.00032299 | 1.423e-05 | 8.0377e-05 | 1.7464e-05 | -5.2002e-06 | [5 empty] | 21.0584 | 4.6831 | 7.3085 | 2.3476e-05 | 7.7102 | 3.9483e-05 | 1.289 | 1.3454e-05 | 2.7157e-06 |  |  |  | 1 | 0 | 0 | 21.0575 | 4.683 | 7.3082 | 9.4088e-06 | 7.7099 | 2.5252e-05 | 1.289 | -4.0097e-06 | 7.9159e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8293 | 21.0575 | 4.683 | 7.3082 | 7.7099 | 1.2889 | 0.22239 | 0.34706 | 0.36613 | 0.061211 | 1.5606 | 1.6464 | 0.27524 | 1.055 | 0.17637 | 0.16718 | -1.5033 | -1.0583 | -1.0048 | -2.7934 | [10 empty] | 0 | 1.8293 | 0.21793 | 0.33337 | 0.3381 | 0.054384 | 1.5297 | 1.5514 | 0.24955 | 1.0142 | 0.16314 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42507 | 0.43917 | -1.3881 | 0.014095 | -1.8132 | -1.8273 | [14 empty] | IPGP | Lab 1 | Jun 23, 2022 | 6 | 112.4% | 0.89008 | -1.5951 | 12.8778 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.6% | 2.7905 | -2.2559 | -4.8271 | SRM 3169
Lab 1 | Jun 23, 2022 | 6 | Zr_244 | iRM_12 |  |  | 0.00024759 | 4.0594e-05 | 7.4767e-05 | 1.1993e-06 | 8.3859e-05 | 2.4365e-06 | 5.0321e-05 | 1.1621e-05 | -8.1441e-07 | [5 empty] | 18.782 | 4.1772 | 6.5194 | 3.2721e-05 | 6.8788 | 2.0476e-05 | 1.1502 | 1.2716e-05 | 1.2529e-06 |  |  |  | 1 | 0 | 0 | 18.7817 | 4.1771 | 6.5194 | 3.1522e-05 | 6.8787 | 1.8039e-05 | 1.1501 | 1.0951e-06 | 2.0673e-06 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8361 | 18.7817 | 4.1771 | 6.5194 | 6.8787 | 1.1501 | 0.2224 | 0.34711 | 0.36624 | 0.061236 | 1.5607 | 1.6467 | 0.27534 | 1.0551 | 0.17642 | 0.1672 | -1.5033 | -1.0581 | -1.0045 | -2.793 | [10 empty] | 0 | 1.8361 | 0.21793 | 0.33337 | 0.3381 | 0.054383 | 1.5297 | 1.5514 | 0.24954 | 1.0142 | 0.16313 | 0.16085 | -1.5236 | -1.0985 | -2.9117 | 0.42508 | 0.43917 | -1.3881 | 0.014095 | -1.8132 | -1.8273 | [22 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | 2.4343 | 0.20258 | 10.3596 | SRM 3169
[56 empty] | 0 | 0 | 0 | [67 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.84% | -2.6032 | 2.1483 | -1.8471 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 45 | SRM 3169 |  |  | 0.0060243 | 0.0012854 | 0.0019181 |  | 0.0018299 | 1.1331e-05 | 0.00030125 |  |  | 4.2275e-07 | 1.1594e-05 | 5.665e-06 |  |  | 41.708 | 9.2653 | 14.4431 |  | 15.205 | 1.5984e-05 | 2.5369 |  |  | 3.2489e-07 | 2.2722e-05 | 2.6956e-07 | 1 | 0 | 0 | 41.7019 | 9.264 | 14.4412 | 0 | 15.2032 | 4.6539e-06 | 2.5366 | 0 | 0 | -9.7861e-08 | 1.1128e-05 | -5.3954e-06 | 0 | 0 | 1 | 1 | 1 | 1.7308 | 41.7019 | 9.264 | 14.4412 | 15.2032 | 2.5366 | 0.22215 | 0.3463 | 0.36457 | 0.060828 | 1.5589 | 1.6411 | 0.27381 | 1.0528 | 0.17565 | 0.16685 | -1.5044 | -1.0605 | -1.009 | -2.7997 | [9 empty] | L2 May | 0 | 1.7308 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014134 | -1.813 | -1.8272 | [22 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | -2.8419 | -3.7454 | -5.2241 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 48 | SRM 3169 |  |  | 0.0069515 | 0.0014908 | 0.0022121 |  | 0.0021582 | 1.7524e-05 | 0.00036713 |  |  | -1.3479e-06 | 2.277e-05 | 1.1794e-06 |  |  | 41.7033 | 9.2644 | 14.4418 |  | 15.2039 | 1.1492e-05 | 2.5368 |  |  | 3.3852e-06 | 2.405e-05 | 5.1183e-06 | 1 | 0 | 0 | 41.6963 | 9.2629 | 14.4395 | 0 | 15.2017 | -6.0321e-06 | 2.5365 | 0 | 0 | 4.7331e-06 | 1.2794e-06 | 3.9389e-06 | 0 | 0 | 1 | 1 | 1 | 1.7317 | 41.6963 | 9.2629 | 14.4395 | 15.2017 | 2.5365 | 0.22215 | 0.3463 | 0.36458 | 0.060832 | 1.5589 | 1.6411 | 0.27383 | 1.0528 | 0.17566 | 0.16685 | -1.5044 | -1.0604 | -1.009 | -2.7996 | [10 empty] | 0 | 1.7317 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014136 | -1.813 | -1.8272 | [14 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.1% | 2.5461 | -1.8685 | 6.0688 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.82% | 2.6853 | 3.2157 | 3.1469 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 51 | SRM 3169 |  |  | 0.006788 | 0.0014586 | 0.002166 |  | 0.0020629 | 1.0792e-05 | 0.00034028 |  |  | 1.6108e-06 | 3.0252e-06 | 1.2604e-06 |  |  | 41.6186 | 9.2456 | 14.4127 |  | 15.1734 | 3.0944e-05 | 2.5318 |  |  | 1.564e-06 | 1.165e-05 | 9.3567e-07 | 1 | 0 | 0 | 41.6118 | 9.2441 | 14.4105 | 0 | 15.1714 | 2.0152e-05 | 2.5314 | 0 | 0 | -4.6851e-08 | 8.6246e-06 | -3.2475e-07 | 0 | 0 | 1 | 1 | 1 | 1.7323 | 41.6118 | 9.2441 | 14.4105 | 15.1714 | 2.5314 | 0.22215 | 0.34631 | 0.36459 | 0.060834 | 1.5589 | 1.6412 | 0.27384 | 1.0528 | 0.17566 | 0.16685 | -1.5044 | -1.0604 | -1.009 | -2.7996 | [10 empty] | 0 | 1.7323 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014133 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.77% | -0.14288 | 1.7659 | 5.3228 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.5% | -2.7605 | -2.5207 | 1.8336 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 54 | Bottle 1 |  |  | 0.0073336 | 0.0015692 | 0.002356 |  | 0.0022705 | 1.1422e-06 | 0.00037911 |  |  | 7.7936e-06 | 1.0331e-05 | 7.791e-06 |  |  | 42.6116 | 9.466 | 14.756 |  | 15.5342 | 3.0464e-05 | 2.5919 |  |  | 1.1875e-06 | 1.4677e-05 | 7.6642e-06 | 1 | 0 | 0 | 42.6043 | 9.4644 | 14.7536 | 0 | 15.532 | 2.9321e-05 | 2.5915 | 0 | 0 | -6.6061e-06 | 4.3462e-06 | -1.2686e-07 | 0 | 0 | 1 | 1 | 1 | 1.7305 | 42.6043 | 9.4644 | 14.7536 | 15.532 | 2.5915 | 0.22215 | 0.34629 | 0.36456 | 0.060826 | 1.5588 | 1.6411 | 0.27381 | 1.0528 | 0.17565 | 0.16685 | -1.5044 | -1.0605 | -1.0091 | -2.7997 | [10 empty] | 0 | 1.7305 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014134 | -1.813 | -1.8272 | [14 empty] | Bottle 1 | Lab 2 | May 28, 2020 | 1 | 102.3% | 3.4224 | 0.60407 | -2.1549 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.63% | 3.8399 | 1.4405 | -4.8692 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 57 | SRM 3169 |  |  | 0.0076311 | 0.0016189 | 0.0024353 |  | 0.0023619 | 1.2068e-05 | 0.00038835 |  |  | 4.8468e-06 | 7.1245e-06 | 1.9851e-06 |  |  | 41.722 | 9.2684 | 14.4481 |  | 15.2104 | 3.0092e-05 | 2.5379 |  |  | 1.4924e-06 | 2.7732e-05 | 5.9893e-06 | 1 | 0 | 0 | 41.7144 | 9.2668 | 14.4457 | 0 | 15.2081 | 1.8024e-05 | 2.5375 | 0 | 0 | -3.3544e-06 | 2.0608e-05 | 4.0042e-06 | 0 | 0 | 1 | 1 | 1 | 1.7313 | 41.7144 | 9.2668 | 14.4457 | 15.208 | 2.5375 | 0.22215 | 0.3463 | 0.36458 | 0.060829 | 1.5589 | 1.6411 | 0.27382 | 1.0528 | 0.17566 | 0.16685 | -1.5044 | -1.0605 | -1.009 | -2.7997 | [10 empty] | 0 | 1.7313 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014135 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.1% | -3.9202 | -1.1812 | -7.3092 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.76% | -2.249 | -0.26761 | 2.585 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 60 | Bottle 2 |  |  | 0.0089924 | 0.0019314 | 0.0029094 |  | 0.0028335 | 1.5611e-06 | 0.00047083 |  |  | 3.5495e-06 | 1.9459e-06 | 5.3511e-06 |  |  | 42.1261 | 9.358 | 14.5873 |  | 15.3561 | 2.7473e-05 | 2.5621 |  |  | -9.906e-08 | 1.568e-05 | 9.5494e-06 | 1 | 0 | 0 | 42.1171 | 9.356 | 14.5844 | 0 | 15.3533 | 2.5911e-05 | 2.5616 | 0 | 0 | -3.6486e-06 | 1.3734e-05 | 4.1984e-06 | 0 | 0 | 1 | 1 | 1 | 1.7289 | 42.1171 | 9.356 | 14.5843 | 15.3533 | 2.5616 | 0.22214 | 0.34628 | 0.36454 | 0.06082 | 1.5588 | 1.641 | 0.27379 | 1.0527 | 0.17564 | 0.16684 | -1.5044 | -1.0605 | -1.0091 | -2.7998 | [10 empty] | 0 | 1.7289 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014137 | -1.813 | -1.8272 | [14 empty] | Bottle 2 | Lab 2 | May 28, 2020 | 1 | 100.9% | 1.4674 | -1.4703 | 2.6604 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | 3.0501 | 5.1971 | 2.9058 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 63 | SRM 3169 |  |  | 0.0082643 | 0.0017761 | 0.0026347 |  | 0.002567 | 1.2737e-05 | 0.00043181 |  |  | 5.9702e-06 | 1.5074e-05 | -5.965e-07 |  |  | 41.746 | 9.2736 | 14.4558 |  | 15.2179 | 2.4036e-05 | 2.539 |  |  | -8.882e-07 | 1.3744e-05 | -1.7049e-06 | 1 | 0 | 0 | 41.7378 | 9.2718 | 14.4532 | 0 | 15.2153 | 1.1299e-05 | 2.5386 | 0 | 0 | -6.8584e-06 | -1.3305e-06 | -1.1084e-06 | 0 | 0 | 1 | 1 | 1 | 1.7293 | 41.7378 | 9.2718 | 14.4532 | 15.2153 | 2.5386 | 0.22214 | 0.34629 | 0.36455 | 0.060822 | 1.5588 | 1.641 | 0.2738 | 1.0527 | 0.17564 | 0.16684 | -1.5044 | -1.0605 | -1.0091 | -2.7998 | [10 empty] | 0 | 1.7293 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014135 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.95% | 4.102 | 4.102 | -9.8968 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.76% | -0.39864 | -3.4271 | -4.7422 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 66 | Bottle 3 |  |  | 0.0075906 | 0.0016185 | 0.0024047 |  | 0.002297 | 1.8972e-05 | 0.00037136 |  |  | 4.2579e-06 | 2.2045e-06 | -1.7857e-06 |  |  | 42.5909 | 9.4609 | 14.7475 |  | 15.524 | 3.4694e-05 | 2.59 |  |  | -3.0881e-06 | 1.9585e-05 | 1.5326e-07 | 1 | 0 | 0 | 42.5833 | 9.4593 | 14.745 | 0 | 15.5217 | 1.5722e-05 | 2.5896 | 0 | 0 | -7.346e-06 | 1.7381e-05 | 1.939e-06 | 0 | 0 | 1 | 1 | 1 | 1.7267 | 42.5833 | 9.4593 | 14.745 | 15.5217 | 2.5896 | 0.22214 | 0.34626 | 0.3645 | 0.060812 | 1.5588 | 1.6409 | 0.27376 | 1.0527 | 0.17562 | 0.16684 | -1.5045 | -1.0606 | -1.0092 | -2.8 | [10 empty] | 0 | 1.7267 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.01414 | -1.813 | -1.8272 | [14 empty] | Bottle 3 | Lab 2 | May 28, 2020 | 1 | 102% | -5.9992 | -0.77187 | -10.1495 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.4% | -0.19919 | -0.0061023 | -10.942 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 69 | SRM 3169 |  |  | 0.0076996 | 0.0016358 | 0.0024391 |  | 0.0023366 | 1.0992e-05 | 0.00036554 |  |  | 2.1021e-06 | 1.6628e-05 | 4.5406e-06 |  |  | 41.8067 | 9.2866 | 14.4754 |  | 15.2372 | 2.2826e-05 | 2.5421 |  |  | 2.6894e-06 | 1.291e-05 | -1.006e-06 | 1 | 0 | 0 | 41.799 | 9.285 | 14.4729 | 0 | 15.2349 | 1.1834e-05 | 2.5417 | 0 | 0 | 5.8734e-07 | -3.7176e-06 | -5.5466e-06 | 0 | 0 | 1 | 1 | 1 | 1.7253 | 41.799 | 9.285 | 14.4729 | 15.2349 | 2.5417 | 0.22213 | 0.34625 | 0.36448 | 0.060808 | 1.5587 | 1.6408 | 0.27375 | 1.0526 | 0.17562 | 0.16684 | -1.5045 | -1.0606 | -1.0093 | -2.8 | [10 empty] | 0 | 1.7253 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014143 | -1.813 | -1.8271 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.75% | 0.95831 | -5.5315 | 24.2923 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.66% | -4.8473 | -2.6118 | 33.4628 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 72 | Bottle 4 |  |  | 0.0085318 | 0.0018092 | 0.0027259 |  | 0.0026037 | -1.3443e-06 | 0.0004301 |  |  | 4.4169e-06 | 2.8215e-06 | 4.1327e-06 |  |  | 42.6708 | 9.4782 | 14.7736 |  | 15.5499 | 3.1125e-05 | 2.5941 |  |  | 5.6386e-06 | 6.9809e-06 | 4.6083e-06 | 1 | 0 | 0 | 42.6623 | 9.4764 | 14.7708 | 0 | 15.5473 | 3.2469e-05 | 2.5936 | 0 | 0 | 1.2217e-06 | 4.1593e-06 | 4.7565e-07 | 0 | 0 | 1 | 1 | 1 | 1.722 | 42.6623 | 9.4764 | 14.7708 | 15.5473 | 2.5936 | 0.22213 | 0.34623 | 0.36443 | 0.060795 | 1.5587 | 1.6406 | 0.27369 | 1.0526 | 0.17559 | 0.16682 | -1.5045 | -1.0607 | -1.0094 | -2.8003 | [10 empty] | 0 | 1.722 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014141 | -1.813 | -1.8272 | [14 empty] | Bottle 4 | Lab 2 | May 28, 2020 | 1 | 101.7% | 2.6211 | 0.038956 | -0.57946 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.7% | 2.0254 | -1.4654 | -21.4659 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 75 | SRM 3169 |  |  | 0.0098377 | 0.0021074 | 0.0031586 |  | 0.0030784 | 1.1681e-05 | 0.00049149 |  |  | -2.3675e-07 | 1.6744e-05 | 5.591e-06 |  |  | 42.0818 | 9.3476 | 14.5704 |  | 15.3368 | 3.1964e-05 | 2.5586 |  |  | 1.5863e-06 | 2.0943e-05 | 3.3525e-06 | 1 | 0 | 0 | 42.072 | 9.3455 | 14.5672 | 0 | 15.3337 | 2.0283e-05 | 2.5581 | 0 | 0 | 1.823e-06 | 4.1993e-06 | -2.2385e-06 | 0 | 0 | 1 | 1 | 1 | 1.7243 | 42.072 | 9.3455 | 14.5672 | 15.3337 | 2.5581 | 0.22213 | 0.34624 | 0.36446 | 0.060803 | 1.5587 | 1.6408 | 0.27373 | 1.0526 | 0.17561 | 0.16683 | -1.5045 | -1.0606 | -1.0093 | -2.8001 | [10 empty] | 0 | 1.7243 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.01414 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100% | 0.40541 | 2.1912 | -11.9507 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 98.84% | 6.4259 | 7.1597 | -12.8352 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 78 | Bottle 5 |  |  | 0.0085863 | 0.0018368 | 0.0027199 |  | 0.0026152 | 7.5215e-06 | 0.00043217 |  |  | 1.4266e-06 | 1.7852e-05 | 7.5024e-06 |  |  | 43.5048 | 9.6632 | 15.0617 |  | 15.8527 | 2.705e-05 | 2.6445 |  |  | 8.5453e-06 | 6.3635e-06 | -2.2271e-06 | 1 | 0 | 0 | 43.4962 | 9.6614 | 15.059 | 0 | 15.85 | 1.9528e-05 | 2.644 | 0 | 0 | 7.1187e-06 | -1.1488e-05 | -9.7296e-06 | 0 | 0 | 1 | 1 | 1 | 1.7202 | 43.4962 | 9.6614 | 15.059 | 15.8501 | 2.644 | 0.22212 | 0.34622 | 0.3644 | 0.060788 | 1.5587 | 1.6406 | 0.27367 | 1.0525 | 0.17558 | 0.16682 | -1.5045 | -1.0607 | -1.0095 | -2.8004 | [10 empty] | 0 | 1.7203 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014137 | -1.813 | -1.8272 | [14 empty] | Bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.1% | 2.856 | 3.5409 | 4.5114 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.8% | -9.4802 | -8.0958 | 14.1968 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 81 | SRM 3169 |  |  | 0.0083994 | 0.0017966 | 0.0026712 |  | 0.0025301 | 1.7806e-05 | 0.00041818 |  |  | 2.5911e-06 | 9.5279e-06 | 2.1346e-06 |  |  | 42.3278 | 9.4019 | 14.6548 |  | 15.4249 | 1.8383e-05 | 2.5732 |  |  | 3.355e-06 | 4.8576e-06 | 1.7658e-06 | 1 | 0 | 0 | 42.3194 | 9.4001 | 14.6521 | 0 | 15.4223 | 5.7707e-07 | 2.5728 | 0 | 0 | 7.6394e-07 | -4.6703e-06 | -3.6879e-07 | 0 | 0 | 1 | 1 | 1 | 1.7219 | 42.3194 | 9.4001 | 14.6521 | 15.4224 | 2.5728 | 0.22212 | 0.34623 | 0.36443 | 0.060794 | 1.5587 | 1.6407 | 0.2737 | 1.0526 | 0.17559 | 0.16682 | -1.5045 | -1.0607 | -1.0094 | -2.8003 | [10 empty] | 0 | 1.7219 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014141 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.95% | -5.9338 | 0.048176 | 3.2879 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | 4.4031 | 4.3871 | 11.5719 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 84 | Bottle 6 |  |  | 0.0093457 | 0.0020026 | 0.0029822 |  | 0.0028772 | 9.8091e-06 | 0.00046009 |  |  | 5.1561e-06 | 1.6389e-05 | 3.1964e-07 |  |  | 43.4506 | 9.651 | 15.0423 |  | 15.8315 | 3.1761e-05 | 2.6408 |  |  | 6.3817e-07 | 1.5031e-05 | 2.1213e-06 | 1 | 0 | 0 | 43.4412 | 9.649 | 15.0393 | 0 | 15.8286 | 2.1952e-05 | 2.6404 | 0 | 0 | -4.5179e-06 | -1.3586e-06 | 1.8017e-06 | 0 | 0 | 1 | 1 | 1 | 1.7182 | 43.4412 | 9.649 | 15.0393 | 15.8286 | 2.6404 | 0.22212 | 0.3462 | 0.36437 | 0.06078 | 1.5586 | 1.6404 | 0.27364 | 1.0525 | 0.17556 | 0.16681 | -1.5046 | -1.0607 | -1.0096 | -2.8005 | [10 empty] | 0 | 1.7182 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014137 | -1.813 | -1.8271 | [14 empty] | Bottle 6 | Lab 2 | May 28, 2020 | 1 | 102.3% | 6.0491 | 4.6312 | 10.1566 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.61% | -2.548 | -0.3192 | -22.3548 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 87 | SRM 3169 |  |  | 0.0091396 | 0.0019644 | 0.0028995 |  | 0.0027888 | 1.0336e-05 | 0.00045899 |  |  | 6.5845e-06 | 1.3187e-05 | 5.3759e-06 |  |  | 42.6151 | 9.4657 | 14.7538 |  | 15.5288 | 3.3285e-05 | 2.5905 |  |  | 1.6682e-06 | 2.8279e-05 | 3.1181e-06 | 1 | 0 | 0 | 42.606 | 9.4637 | 14.7509 | 0 | 15.526 | 2.2949e-05 | 2.59 | 0 | 0 | -4.9163e-06 | 1.5092e-05 | -2.2579e-06 | 0 | 0 | 1 | 1 | 1 | 1.7208 | 42.606 | 9.4637 | 14.7509 | 15.526 | 2.59 | 0.22212 | 0.34622 | 0.36441 | 0.060789 | 1.5587 | 1.6406 | 0.27368 | 1.0525 | 0.17558 | 0.16682 | -1.5045 | -1.0607 | -1.0095 | -2.8003 | [10 empty] | 0 | 1.7208 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014142 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | 0.46844 | 0.063249 | -1.5074 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | 3.9651 | -1.9528 | 19.521 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 90 | bottle 1 |  |  | 0.0087855 | 0.0018623 | 0.0027703 |  | 0.0026337 | 1.1676e-05 | 0.00042884 |  |  | 7.0414e-06 | 2.9921e-05 | 8.7137e-06 |  |  | 44.8137 | 9.9537 | 15.514 |  | 16.3277 | 3.4866e-05 | 2.7236 |  |  | 3.4603e-06 | 9.8229e-06 | 6.7714e-06 | 1 | 0 | 0 | 44.8049 | 9.9518 | 15.5112 | 0 | 16.325 | 2.319e-05 | 2.7231 | 0 | 0 | -3.5811e-06 | -2.0099e-05 | -1.9423e-06 | 0 | 0 | 1 | 1 | 1 | 1.7176 | 44.8049 | 9.9518 | 15.5112 | 16.325 | 2.7232 | 0.22211 | 0.34619 | 0.36436 | 0.060778 | 1.5586 | 1.6404 | 0.27363 | 1.0525 | 0.17556 | 0.16681 | -1.5046 | -1.0608 | -1.0096 | -2.8005 | [10 empty] | 0 | 1.7176 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014138 | -1.813 | -1.8271 | [14 empty] | bottle 1 | Lab 2 | May 28, 2020 | 1 | 105.1% | -0.69285 | 5.3125 | 20.2282 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.85% | -2.5014 | 3.1512 | -11.7129 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 93 | SRM 3169 |  |  | 0.0096603 | 0.002059 | 0.0030792 |  | 0.0029382 | 1.6745e-05 | 0.00048713 |  |  | 7.0783e-07 | 1.5645e-05 | 4.4168e-07 |  |  | 42.6392 | 9.4708 | 14.7613 |  | 15.5357 | 2.8765e-05 | 2.5915 |  |  | -8.3091e-07 | 1.2658e-05 | 7.6934e-06 | 1 | 0 | 0 | 42.6295 | 9.4687 | 14.7582 | 0 | 15.5328 | 1.202e-05 | 2.591 | 0 | 0 | -1.5387e-06 | -2.9868e-06 | 7.2517e-06 | 0 | 0 | 1 | 1 | 1 | 1.7181 | 42.6295 | 9.4687 | 14.7582 | 15.5328 | 2.591 | 0.22212 | 0.3462 | 0.36437 | 0.060779 | 1.5586 | 1.6404 | 0.27364 | 1.0525 | 0.17556 | 0.16681 | -1.5046 | -1.0607 | -1.0096 | -2.8005 | [10 empty] | 0 | 1.7181 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014144 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.6% | 2.7905 | -2.2559 | -4.8271 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.3% | 2.533 | -0.49612 | 0.063121 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 96 | bottle 2 |  |  | 0.0096122 | 0.0020486 | 0.0030571 |  | 0.0029387 | 7.4942e-07 | 0.00047254 |  |  | 4.1103e-06 | 1.5258e-05 | -1.2833e-06 |  |  | 44.0907 | 9.7929 | 15.2631 |  | 16.0632 | 2.3979e-05 | 2.6793 |  |  | -3.9593e-07 | 2.32e-05 | 4.938e-06 | 1 | 0 | 0 | 44.0811 | 9.7908 | 15.2601 | 0 | 16.0603 | 2.3229e-05 | 2.6789 | 0 | 0 | -4.5062e-06 | 7.9425e-06 | 6.2213e-06 | 0 | 0 | 1 | 1 | 1 | 1.7161 | 44.0811 | 9.7908 | 15.2601 | 16.0603 | 2.6789 | 0.22211 | 0.34618 | 0.36433 | 0.060771 | 1.5586 | 1.6403 | 0.27361 | 1.0524 | 0.17555 | 0.1668 | -1.5046 | -1.0608 | -1.0097 | -2.8006 | [10 empty] | 0 | 1.7161 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.01414 | -1.813 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | May 28, 2020 | 1 | 103% | -6.5635 | 1.963 | -4.7327 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.4% | -1.5296 | -0.29484 | -2.4415 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 99 | SRM 3169 |  |  | 0.01048 | 0.0022422 | 0.0033433 |  | 0.003218 | 7.8746e-06 | 0.00053428 |  |  | 4.5717e-06 | 1.1971e-05 | 4.5436e-06 |  |  | 43.0085 | 9.5527 | 14.8889 |  | 15.6697 | 3.0248e-05 | 2.6138 |  |  | 5.1403e-06 | 1.0707e-05 | 5.8863e-06 | 1 | 0 | 0 | 42.998 | 9.5505 | 14.8855 | 0 | 15.6665 | 2.2374e-05 | 2.6133 | 0 | 0 | 5.6861e-07 | -1.2635e-06 | 1.3427e-06 | 0 | 0 | 1 | 1 | 1 | 1.7173 | 42.998 | 9.5505 | 14.8855 | 15.6665 | 2.6133 | 0.22211 | 0.34619 | 0.36435 | 0.060776 | 1.5586 | 1.6404 | 0.27363 | 1.0525 | 0.17556 | 0.16681 | -1.5046 | -1.0608 | -1.0096 | -2.8006 | [10 empty] | 0 | 1.7173 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014141 | -1.813 | -1.8271 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | 2.4343 | 0.20258 | 10.3596 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100% | 1.4732 | 0.62772 | 5.1002 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 102 | bottle 3 |  |  | 0.0094937 | 0.0020308 | 0.0030079 |  | 0.0028553 | 1.0894e-05 | 0.00046694 |  |  | 5.5852e-06 | 1.2632e-05 | 2.6817e-06 |  |  | 44.5501 | 9.8948 | 15.4217 |  | 16.2295 | 1.4751e-05 | 2.707 |  |  | 1.0507e-05 | 1.5767e-05 | 5.3317e-06 | 1 | 0 | 0 | 44.5406 | 9.8928 | 15.4187 | 0 | 16.2266 | 3.8563e-06 | 2.7065 | 0 | 0 | 4.9217e-06 | 3.135e-06 | 2.65e-06 | 0 | 0 | 1 | 1 | 1 | 1.7146 | 44.5406 | 9.8928 | 15.4187 | 16.2266 | 2.7065 | 0.22211 | 0.34617 | 0.36431 | 0.060765 | 1.5586 | 1.6403 | 0.27359 | 1.0524 | 0.17554 | 0.16679 | -1.5046 | -1.0608 | -1.0097 | -2.8007 | [10 empty] | 0 | 1.7146 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.01414 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 103.5% | -0.44047 | -0.10692 | -7.0333 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | -2.6095 | 0.70076 | -2.1874 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 105 | SRM 3169 |  |  | 0.0097784 | 0.0020926 | 0.0030878 |  | 0.0029427 | 2.6837e-06 | 0.00047276 |  |  | 2.0137e-06 | 1.2692e-05 | -3.4424e-06 |  |  | 43.0778 | 9.5679 | 14.9125 |  | 15.6941 | 3.68e-05 | 2.6178 |  |  | -3.6995e-07 | 1.6385e-05 | 6.4391e-06 | 1 | 0 | 0 | 43.0681 | 9.5658 | 14.9094 | 0 | 15.6912 | 3.4116e-05 | 2.6173 | 0 | 0 | -2.3836e-06 | 3.6937e-06 | 9.8815e-06 | 0 | 0 | 1 | 1 | 1 | 1.716 | 43.0681 | 9.5658 | 14.9094 | 15.6911 | 2.6173 | 0.22211 | 0.34618 | 0.36433 | 0.060771 | 1.5586 | 1.6403 | 0.27361 | 1.0524 | 0.17555 | 0.1668 | -1.5046 | -1.0608 | -1.0097 | -2.8006 | [10 empty] | 0 | 1.716 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014138 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.84% | -2.6032 | 2.1483 | -1.8471 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | 3.3505 | -1.3965 | 2.9983 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 108 | bottle 4 |  |  | 0.0099462 | 0.0021096 | 0.0031374 |  | 0.0030084 | 1.1048e-05 | 0.00049545 |  |  | 8.0474e-06 | 1.2753e-05 | -1.1165e-06 |  |  | 44.1779 | 9.8119 | 15.2919 |  | 16.0919 | 2.9205e-05 | 2.6839 |  |  | 4.9089e-06 | 6.1201e-06 | -6.7205e-07 | 1 | 0 | 0 | 44.1679 | 9.8098 | 15.2888 | 0 | 16.0889 | 1.8157e-05 | 2.6834 | 0 | 0 | -3.1385e-06 | -6.6329e-06 | 4.4448e-07 | 0 | 0 | 1 | 1 | 1 | 1.7118 | 44.1679 | 9.8098 | 15.2888 | 16.0889 | 2.6834 | 0.2221 | 0.34615 | 0.36427 | 0.060754 | 1.5585 | 1.6401 | 0.27354 | 1.0523 | 0.17551 | 0.16678 | -1.5046 | -1.0609 | -1.0099 | -2.8009 | [10 empty] | 0 | 1.7118 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014133 | -1.813 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 102.3% | 7.5247 | 5.6486 | 1.2341 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.17% | 5.4825 | 3.2857 | -17.9386 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 111 | SRM 3169 |  |  | 0.01163 | 0.0024926 | 0.0037342 |  | 0.003602 | 2.0879e-05 | 0.00058562 |  |  | 2.8293e-06 | 1.5872e-05 | 4.481e-06 |  |  | 43.2905 | 9.615 | 14.9855 |  | 15.7704 | 2.0377e-05 | 2.6304 |  |  | 4.1984e-06 | 2.9672e-05 | 9.9294e-07 | 1 | 0 | 0 | 43.2789 | 9.6125 | 14.9818 | 0 | 15.7668 | -5.0248e-07 | 2.6298 | 0 | 0 | 1.3691e-06 | 1.3799e-05 | -3.488e-06 | 0 | 0 | 1 | 1 | 1 | 1.7143 | 43.2789 | 9.6125 | 14.9818 | 15.7667 | 2.6298 | 0.22211 | 0.34617 | 0.36431 | 0.060764 | 1.5586 | 1.6402 | 0.27358 | 1.0524 | 0.17553 | 0.16679 | -1.5046 | -1.0608 | -1.0098 | -2.8008 | [10 empty] | 0 | 1.7143 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.01414 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | -2.8419 | -3.7454 | -5.2241 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.4% | -6.7098 | -3.1289 | 11.3563 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 114 | bottle 5 |  |  | 0.012068 | 0.0025721 | 0.0038594 |  | 0.0037411 | 1.1534e-05 | 0.00060506 |  |  | 4.0818e-06 | 8.5962e-06 | -1.922e-06 |  |  | 44.7163 | 9.9314 | 15.478 |  | 16.2874 | 2.8351e-05 | 2.7164 |  |  | 5.6597e-07 | 2.3584e-05 | 1.908e-06 | 1 | 0 | 0 | 44.7042 | 9.9288 | 15.4741 | 0 | 16.2837 | 1.6817e-05 | 2.7158 | 0 | 0 | -3.5158e-06 | 1.4988e-05 | 3.83e-06 | 0 | 0 | 1 | 1 | 1 | 1.711 | 44.7042 | 9.9288 | 15.4741 | 16.2837 | 2.7158 | 0.2221 | 0.34614 | 0.36425 | 0.06075 | 1.5585 | 1.64 | 0.27353 | 1.0523 | 0.17551 | 0.16678 | -1.5046 | -1.0609 | -1.0099 | -2.801 | [10 empty] | 0 | 1.711 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014138 | -1.813 | -1.8272 | [14 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.2% | 8.6478 | -1.0155 | -6.4352 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | -6.9724 | -4.1765 | 23.0857 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 117 | SRM 3169 |  |  | 0.011666 | 0.002488 | 0.0037199 |  | 0.0036131 | -4.8511e-06 | 0.0005798 |  |  | 9.0274e-07 | 1.5688e-05 | 2.5981e-06 |  |  | 43.3398 | 9.6259 | 15.0024 |  | 15.7878 | 2.9671e-05 | 2.6332 |  |  | 5.5945e-06 | 1.1501e-05 | 5.108e-06 | 1 | 0 | 0 | 43.3281 | 9.6234 | 14.9987 | 0 | 15.7842 | 3.4522e-05 | 2.6327 | 0 | 0 | 4.6918e-06 | -4.1872e-06 | 2.5099e-06 | 0 | 0 | 1 | 1 | 1 | 1.7136 | 43.3281 | 9.6234 | 14.9987 | 15.7842 | 2.6327 | 0.22211 | 0.34617 | 0.3643 | 0.060761 | 1.5586 | 1.6402 | 0.27357 | 1.0524 | 0.17553 | 0.16679 | -1.5046 | -1.0608 | -1.0098 | -2.8008 | [10 empty] | 0 | 1.7136 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014134 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.82% | 2.6853 | 3.2157 | 3.1469 | [8 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.82% | 10.9992 | 7.2515 | -27.266 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 120 | bottle 6 |  |  | 0.010445 | 0.002235 | 0.0033059 |  | 0.0031547 | 2.234e-05 | 0.0005083 |  |  | 4.6968e-06 | 2.1678e-05 | -5.8929e-06 |  |  | 44.6381 | 9.9138 | 15.4505 |  | 16.2581 | 2.8423e-05 | 2.7115 |  |  | 5.0229e-06 | 1.4264e-05 | 1.0608e-06 | 1 | 0 | 0 | 44.6276 | 9.9116 | 15.4472 | 0 | 16.2549 | 6.0831e-06 | 2.7109 | 0 | 0 | 3.2618e-07 | -7.4137e-06 | 6.9537e-06 | 0 | 0 | 1 | 1 | 1 | 1.7098 | 44.6276 | 9.9116 | 15.4472 | 16.2549 | 2.711 | 0.2221 | 0.34613 | 0.36423 | 0.060746 | 1.5585 | 1.64 | 0.27351 | 1.0523 | 0.1755 | 0.16678 | -1.5046 | -1.0609 | -1.01 | -2.801 | [10 empty] | 0 | 1.7098 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014139 | -1.813 | -1.8272 | [14 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 102.8% | -0.38789 | -4.1636 | 1.88 | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | -3.1984 | -1.0389 | 0.4553 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 123 | SRM 3169 |  |  | 0.010717 | 0.0022848 | 0.0033885 |  | 0.0032332 | 1.527e-05 | 0.00051937 |  |  | 4.3201e-06 | 1.8788e-05 | -2.9151e-07 |  |  | 43.5442 | 9.6712 | 15.0729 |  | 15.8617 | 4.5316e-05 | 2.6455 |  |  | 2.4438e-06 | 1.4773e-05 | 8.5497e-06 | 1 | 0 | 0 | 43.5334 | 9.6689 | 15.0695 | 0 | 15.8585 | 3.0045e-05 | 2.645 | 0 | 0 | -1.8764e-06 | -4.0146e-06 | 8.8412e-06 | 0 | 0 | 1 | 1 | 1 | 1.7128 | 43.5334 | 9.6689 | 15.0695 | 15.8585 | 2.645 | 0.2221 | 0.34616 | 0.36428 | 0.060758 | 1.5585 | 1.6401 | 0.27356 | 1.0524 | 0.17552 | 0.16679 | -1.5046 | -1.0609 | -1.0098 | -2.8009 | [10 empty] | 0 | 1.7128 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014135 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.5% | -2.7605 | -2.5207 | 1.8336
Lab 2 | May 28, 2020 | 1 | 126 | bottle 1 |  |  | 0.012159 | 0.002605 | 0.0038894 |  | 0.0037672 | 7.961e-06 | 0.00059901 |  |  | 4.684e-07 | 9.2166e-06 | 1.7587e-06 |  |  | 46.4446 | 10.3151 | 16.0758 |  | 16.9161 | 2.9966e-05 | 2.8212 |  |  | 7.1058e-06 | 1.2882e-05 | 3.2578e-06 | 1 | 0 | 0 | 46.4324 | 10.3125 | 16.0719 | 0 | 16.9123 | 2.2005e-05 | 2.8206 | 0 | 0 | 6.6374e-06 | 3.6652e-06 | 1.4991e-06 | 0 | 0 | 1 | 1 | 1 | 1.7098 | 46.4324 | 10.3125 | 16.0719 | 16.9123 | 2.8206 | 0.2221 | 0.34614 | 0.36424 | 0.060746 | 1.5585 | 1.64 | 0.27351 | 1.0523 | 0.1755 | 0.16678 | -1.5046 | -1.0609 | -1.01 | -2.8011 | [10 empty] | 0 | 1.7098 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014137 | -1.813 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | May 28, 2020 | 1 | 107% | -3.894 | -4.8872 | -0.10741 | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 100.2% | 5.0631 | -3.0038 | -0.75686 | SRM 3169 | Avg | 0.46094 | -0.14036 | 0.2085
Lab 2 | May 28, 2020 | 1 | 129 | SRM 3169 |  |  | 0.01338 | 0.0028723 | 0.0042887 |  | 0.0041785 | 1.1561e-05 | 0.00066135 |  |  | -2.9015e-06 | 1.4428e-05 | 3.2063e-06 |  |  | 43.2831 | 9.6132 | 14.9822 |  | 15.7659 | 3.4104e-05 | 2.6295 |  |  | 3.2824e-06 | 1.9959e-05 | -2.6058e-07 | 1 | 0 | 0 | 43.2697 | 9.6103 | 14.978 | 0 | 15.7617 | 2.2543e-05 | 2.6288 | 0 | 0 | 6.1839e-06 | 5.531e-06 | -3.4669e-06 | 0 | 0 | 1 | 1 | 1 | 1.7118 | 43.2697 | 9.6103 | 14.9779 | 15.7617 | 2.6288 | 0.2221 | 0.34615 | 0.36427 | 0.060754 | 1.5585 | 1.6401 | 0.27354 | 1.0523 | 0.17551 | 0.16678 | -1.5046 | -1.0609 | -1.0099 | -2.8009 | [10 empty] | 0 | 1.7118 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.01413 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.63% | 3.8399 | 1.4405 | -4.8692 | [8 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.21% | -5.955 | 6.2986 | 15.9334 | SRM 3169 | 2SD | 7.9151 | 9.5532 | 18.6085
Lab 2 | May 28, 2020 | 1 | 132 | bottle 2 |  |  | 0.011074 | 0.0023656 | 0.0034906 |  | 0.003332 | 3.5991e-07 | 0.00053293 |  |  | 2.7426e-06 | 1.6194e-05 | -7.3513e-07 |  |  | 45.4348 | 10.0908 | 15.7262 |  | 16.548 | 2.0918e-05 | 2.7598 |  |  | -9.0016e-07 | 2.3811e-05 | 1.2877e-06 | 1 | 0 | 0 | 45.4237 | 10.0884 | 15.7228 | 0 | 16.5447 | 2.0558e-05 | 2.7592 | 0 | 0 | -3.6428e-06 | 7.6177e-06 | 2.0229e-06 | 0 | 0 | 1 | 1 | 1 | 1.7095 | 45.4237 | 10.0884 | 15.7227 | 16.5446 | 2.7592 | 0.2221 | 0.34614 | 0.36423 | 0.060744 | 1.5585 | 1.64 | 0.2735 | 1.0523 | 0.17549 | 0.16677 | -1.5046 | -1.0609 | -1.01 | -2.8011 | [10 empty] | 0 | 1.7095 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.01413 | -1.813 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | May 28, 2020 | 1 | 104.9% | -3.4635 | 0.011026 | -7.8787 | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 100.5% | -0.16222 | -6.1953 | -11.1069 | SRM 3169 | n | 7 | 7 | 7
Lab 2 | May 28, 2020 | 1 | 135 | SRM 3169 |  |  | 0.012789 | 0.0027442 | 0.0040744 |  | 0.0039545 | 2.0667e-05 | 0.00064197 |  |  | 5.8597e-07 | 9.4057e-06 | -1.6602e-06 |  |  | 43.3438 | 9.6265 | 15.0029 |  | 15.7872 | 2.8604e-05 | 2.633 |  |  | 6.7803e-06 | 1.2247e-05 | 6.926e-06 | 1 | 0 | 0 | 43.3311 | 9.6238 | 14.9988 | 0 | 15.7833 | 7.9365e-06 | 2.6323 | 0 | 0 | 6.1944e-06 | 2.8411e-06 | 8.5862e-06 | 0 | 0 | 1 | 1 | 1 | 1.7107 | 43.3311 | 9.6238 | 14.9988 | 15.7833 | 2.6323 | 0.2221 | 0.34614 | 0.36425 | 0.060749 | 1.5585 | 1.64 | 0.27352 | 1.0523 | 0.1755 | 0.16678 | -1.5046 | -1.0609 | -1.0099 | -2.801 | [10 empty] | 0 | 1.7107 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014129 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.76% | -2.249 | -0.26761 | 2.585 | [8 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.71% | 2.9689 | -1.005 | 2.7341 | SRM 3169 | 2SE | 2.9916 | 3.6108 | 7.0333
Lab 2 | May 28, 2020 | 1 | 138 | bottle 3 |  |  | 0.012507 | 0.0026583 | 0.0039909 |  | 0.0038439 | 1.384e-05 | 0.00061328 |  |  | 1.4875e-06 | 1.6639e-05 | 9.823e-06 |  |  | 45.7497 | 10.1606 | 15.8348 |  | 16.6616 | 2.7537e-05 | 2.7786 |  |  | -4.9696e-07 | 1.1941e-05 | 2.7293e-06 | 1 | 0 | 0 | 45.7372 | 10.1579 | 15.8308 | 0 | 16.6578 | 1.3697e-05 | 2.778 | 0 | 0 | -1.9845e-06 | -4.6978e-06 | -7.0937e-06 | 0 | 0 | 1 | 1 | 1 | 1.708 | 45.7372 | 10.1579 | 15.8308 | 16.6578 | 2.778 | 0.22209 | 0.34612 | 0.36421 | 0.060739 | 1.5585 | 1.6399 | 0.27349 | 1.0522 | 0.17548 | 0.16677 | -1.5047 | -1.061 | -1.01 | -2.8012 | [10 empty] | 0 | 1.708 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014129 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 105.2% | -0.56221 | -1.0899 | 6.9535 | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 100.1% | 2.2062 | 5.2455 | -7.906 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 141 | SRM 3169 |  |  | 0.011785 | 0.0025005 | 0.0037258 |  | 0.0035773 | 2.2761e-05 | 0.00056651 |  |  | 1.9424e-06 | 5.1898e-06 | 4.1526e-06 |  |  | 43.6091 | 9.6853 | 15.0943 |  | 15.883 | 2.6747e-05 | 2.6489 |  |  | 1.8664e-06 | 1.8596e-05 | 6.8621e-06 | 1 | 0 | 0 | 43.5973 | 9.6828 | 15.0906 | 0 | 15.8794 | 3.9859e-06 | 2.6483 | 0 | 0 | -7.5977e-08 | 1.3407e-05 | 2.7094e-06 | 0 | 0 | 1 | 1 | 1 | 1.7094 | 43.5973 | 9.6828 | 15.0906 | 15.8794 | 2.6483 | 0.2221 | 0.34614 | 0.36423 | 0.060744 | 1.5585 | 1.64 | 0.2735 | 1.0523 | 0.17549 | 0.16678 | -1.5046 | -1.0609 | -1.01 | -2.8011 | [10 empty] | 0 | 1.7094 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014127 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | 3.0501 | 5.1971 | 2.9058 | [8 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.99% | -3.6295 | -4.2218 | 7.1394 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 144 | bottle 4 |  |  | 0.011459 | 0.0024332 | 0.0036083 |  | 0.0034531 | 1.2887e-05 | 0.00054497 |  |  | -2.7953e-07 | 1.0078e-05 | 5.3638e-06 |  |  | 44.8291 | 9.9558 | 15.5153 |  | 16.3246 | 2.4143e-05 | 2.7223 |  |  | 1.365e-06 | 1.0461e-05 | 4.3385e-06 | 1 | 0 | 0 | 44.8177 | 9.9534 | 15.5116 | 0 | 16.3212 | 1.1255e-05 | 2.7218 | 0 | 0 | 1.6445e-06 | 3.8338e-07 | -1.0253e-06 | 0 | 0 | 1 | 1 | 1 | 1.7056 | 44.8177 | 9.9534 | 15.5116 | 16.3212 | 2.7218 | 0.22209 | 0.34611 | 0.36417 | 0.06073 | 1.5584 | 1.6398 | 0.27345 | 1.0522 | 0.17547 | 0.16676 | -1.5047 | -1.061 | -1.0101 | -2.8013 | [10 empty] | 0 | 1.7056 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.01413 | -1.813 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 102.7% | 0.60224 | 1.025 | 7.7373 | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.79% | 2.7351 | 1.8993 | -4.5777 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 147 | SRM 3169 |  |  | 0.011481 | 0.002437 | 0.0036364 |  | 0.0034501 | 2.1733e-05 | 0.00054548 |  |  | 1.2117e-06 | 1.7395e-05 | -2.287e-06 |  |  | 43.6762 | 9.7001 | 15.1171 |  | 15.9067 | 3.36e-05 | 2.6527 |  |  | 5.9632e-06 | 2.3314e-05 | 5.5616e-06 | 1 | 0 | 0 | 43.6648 | 9.6976 | 15.1135 | 0 | 15.9032 | 1.1867e-05 | 2.6522 | 0 | 0 | 4.7514e-06 | 5.9197e-06 | 7.8486e-06 | 0 | 0 | 1 | 1 | 1 | 1.7084 | 43.6648 | 9.6976 | 15.1135 | 15.9032 | 2.6522 | 0.22209 | 0.34613 | 0.36421 | 0.06074 | 1.5585 | 1.6399 | 0.27349 | 1.0523 | 0.17549 | 0.16677 | -1.5047 | -1.061 | -1.01 | -2.8012 | [10 empty] | 0 | 1.7084 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014135 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.76% | -0.39864 | -3.4271 | -4.7422
Lab 2 | May 28, 2020 | 1 | 150 | bottle 5 |  |  | 0.011364 | 0.002412 | 0.0035747 |  | 0.0034099 | 2.1041e-05 | 0.00054397 |  |  | 5.419e-06 | 5.3497e-06 | 1.7713e-06 |  |  | 45.3949 | 10.0814 | 15.7108 |  | 16.5301 | 3.0835e-05 | 2.7565 |  |  | 3.0028e-07 | 1.1218e-05 | -3.8004e-06 | 1 | 0 | 0 | 45.3835 | 10.079 | 15.7072 | 0 | 16.5267 | 9.7935e-06 | 2.756 | 0 | 0 | -5.1187e-06 | 5.8687e-06 | -5.5717e-06 | 0 | 0 | 1 | 1 | 1 | 1.7049 | 45.3835 | 10.079 | 15.7072 | 16.5267 | 2.756 | 0.22208 | 0.3461 | 0.36416 | 0.060727 | 1.5584 | 1.6397 | 0.27344 | 1.0522 | 0.17546 | 0.16676 | -1.5047 | -1.061 | -1.0102 | -2.8014 | [10 empty] | 0 | 1.7048 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014134 | -1.813 | -1.8272 | [14 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.6% | 3.4602 | 1.4591 | 10.0732 | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.3% | -3.5107 | -3.6879 | -7.8244 | SRM 3169 | Avg | -1.9867 | -1.0713 | -0.37857
Lab 2 | May 28, 2020 | 1 | 153 | SRM 3169 |  |  | 0.012067 | 0.0025716 | 0.0038181 |  | 0.0036481 | 9.5572e-06 | 0.00058897 |  |  | 4.4279e-06 | -2.7661e-07 | 5.0725e-06 |  |  | 43.9545 | 9.7617 | 15.2131 |  | 16.0073 | 3.6894e-05 | 2.6695 |  |  | -1.3919e-06 | 1.6223e-05 | -6.6921e-06 | 1 | 0 | 0 | 43.9425 | 9.7592 | 15.2092 | 0 | 16.0037 | 2.7336e-05 | 2.6689 | 0 | 0 | -5.8198e-06 | 1.6499e-05 | -1.1765e-05 | 0 | 0 | 1 | 1 | 1 | 1.7074 | 43.9425 | 9.7592 | 15.2092 | 16.0037 | 2.6689 | 0.22209 | 0.34612 | 0.3642 | 0.060736 | 1.5585 | 1.6399 | 0.27347 | 1.0522 | 0.17548 | 0.16677 | -1.5047 | -1.061 | -1.0101 | -2.8012 | [10 empty] | 0 | 1.7073 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014137 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.4% | -0.19919 | -0.0061023 | -10.942 | [8 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 99.9% | -14.2033 | 3.0566 | 0.99206 | SRM 3169 | 2SD | 15.1192 | 7.3545 | 26.533
Lab 2 | May 28, 2020 | 1 | 156 | bottle 6 |  |  | 0.010834 | 0.0023119 | 0.0034302 |  | 0.0032712 | 4.0256e-06 | 0.00053975 |  |  | 2.3847e-06 | 8.476e-06 | 2.3938e-06 |  |  | 44.6955 | 9.9248 | 15.4651 |  | 16.2684 | 2.6545e-05 | 2.7125 |  |  | 3.6334e-06 | 2.0335e-05 | 3.4622e-06 | 1 | 0 | 0 | 44.6847 | 9.9225 | 15.4617 | 0 | 16.2651 | 2.2519e-05 | 2.7119 | 0 | 0 | 1.2487e-06 | 1.1859e-05 | 1.0683e-06 | 0 | 0 | 1 | 1 | 1 | 1.6948 | 44.6847 | 9.9225 | 15.4617 | 16.2651 | 2.7119 | 0.22206 | 0.34602 | 0.364 | 0.06069 | 1.5582 | 1.6392 | 0.27331 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0613 | -1.0106 | -2.802 | [10 empty] | 0 | 1.6948 | 0.21793 | 0.33335 | 0.3381 | 0.054392 | 1.5296 | 1.5514 | 0.24959 | 1.0142 | 0.16317 | 0.16088 | -1.5236 | -1.0986 | -2.9115 | 0.42503 | 0.43918 | -1.3879 | 0.014148 | -1.813 | -1.8271 | [14 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 101.7% | -13.1806 | -10.4252 | 41.589 | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.5% | 6.7219 | -8.3217 | 18.0906 | SRM 3169 | n | 18 | 18 | 18
Lab 2 | May 28, 2020 | 1 | 159 | SRM 3169 |  |  | 0.011587 | 0.0024702 | 0.0036732 |  | 0.0035066 | 1.608e-05 | 0.00056 |  |  | 4.7215e-06 | 1.6835e-05 | -9.1574e-07 |  |  | 43.9239 | 9.7546 | 15.2016 |  | 15.9944 | 3.0034e-05 | 2.6672 |  |  | 3.6472e-06 | 1.2593e-05 | 4.3149e-06 | 1 | 0 | 0 | 43.9123 | 9.7522 | 15.1979 | 0 | 15.9909 | 1.3954e-05 | 2.6667 | 0 | 0 | -1.0743e-06 | -4.2414e-06 | 5.2306e-06 | 0 | 0 | 1 | 1 | 1 | 1.7048 | 43.9123 | 9.7522 | 15.1979 | 15.9909 | 2.6667 | 0.22208 | 0.3461 | 0.36416 | 0.060727 | 1.5584 | 1.6397 | 0.27344 | 1.0522 | 0.17546 | 0.16676 | -1.5047 | -1.061 | -1.0102 | -2.8014 | [10 empty] | 0 | 1.7048 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014138 | -1.813 | -1.8271 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.66% | -4.8473 | -2.6118 | 33.4628 | [8 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.5% | 9.6758 | 1.4344 | -23.5506 | SRM 3169 | 2SE | 3.5636 | 1.7335 | 6.2539
Lab 2 | May 28, 2020 | 1 | 162 | bottle 1 |  |  | 0.011555 | 0.0024535 | 0.0036389 |  | 0.0034556 | 7.7658e-06 | 0.00053868 |  |  | -2.1114e-06 | 1.6267e-05 | 3.308e-06 |  |  | 47.852 | 10.6267 | 16.5601 |  | 17.4229 | 2.5207e-05 | 2.9053 |  |  | 5.457e-06 | 8.8602e-06 | -2.0621e-06 | 1 | 0 | 0 | 47.8404 | 10.6243 | 16.5565 | 0 | 17.4195 | 1.7441e-05 | 2.9047 | 0 | 0 | 7.5684e-06 | -7.4068e-06 | -5.3701e-06 | 0 | 0 | 1 | 1 | 1 | 1.7023 | 47.8404 | 10.6243 | 16.5565 | 17.4195 | 2.9048 | 0.22208 | 0.34608 | 0.36412 | 0.060718 | 1.5584 | 1.6396 | 0.27341 | 1.0521 | 0.17544 | 0.16675 | -1.5047 | -1.0611 | -1.0103 | -2.8015 | [10 empty] | 0 | 1.7023 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014138 | -1.813 | -1.8271 | [14 empty] | bottle 1 | Lab 2 | May 28, 2020 | 1 | 108.6% | -2.7444 | -1.6295 | 20.9709 | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100% | -6.4198 | 0.43341 | -4.8855 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 165 | SRM 3169 |  |  | 0.012221 | 0.0026166 | 0.0038648 |  | 0.0036743 | 1.4141e-05 | 0.00059223 |  |  | 5.4521e-07 | 1.7273e-05 | 1.2298e-06 |  |  | 44.1952 | 9.815 | 15.2957 |  | 16.0935 | 2.9974e-05 | 2.6837 |  |  | 3.4294e-06 | 2.8278e-05 | 5.0109e-07 | 1 | 0 | 0 | 44.183 | 9.8124 | 15.2919 | 0 | 16.0899 | 1.5833e-05 | 2.6831 | 0 | 0 | 2.8842e-06 | 1.1006e-05 | -7.287e-07 | 0 | 0 | 1 | 1 | 1 | 1.7053 | 44.183 | 9.8124 | 15.2918 | 16.0899 | 2.6831 | 0.22209 | 0.3461 | 0.36416 | 0.060727 | 1.5584 | 1.6397 | 0.27344 | 1.0522 | 0.17546 | 0.16676 | -1.5047 | -1.061 | -1.0102 | -2.8014 | [10 empty] | 0 | 1.7053 | 0.21793 | 0.33335 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014134 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.7% | 2.0254 | -1.4654 | -21.4659 | [8 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.4% | 2.397 | -0.27255 | -0.3493 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 168 | bottle 2 |  |  | 0.013118 | 0.002787 | 0.0041634 |  | 0.0039776 | 1.4799e-08 | 0.00063942 |  |  | -1.2517e-06 | 6.7321e-06 | 2.4535e-06 |  |  | 46.4727 | 10.3206 | 16.0831 |  | 16.9209 | 4.1788e-05 | 2.8215 |  |  | 1.4309e-06 | 2.1036e-05 | 1.3173e-06 | 1 | 0 | 0 | 46.4595 | 10.3178 | 16.0789 | 0 | 16.917 | 4.1773e-05 | 2.8209 | 0 | 0 | 2.6826e-06 | 1.4304e-05 | -1.1361e-06 | 0 | 0 | 1 | 1 | 1 | 1.7027 | 46.4595 | 10.3178 | 16.0789 | 16.917 | 2.8209 | 0.22208 | 0.34608 | 0.36412 | 0.060717 | 1.5584 | 1.6396 | 0.2734 | 1.0521 | 0.17544 | 0.16675 | -1.5047 | -1.0611 | -1.0103 | -2.8015 | [10 empty] | 0 | 1.7027 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.014129 | -1.8131 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | May 28, 2020 | 1 | 105.6% | 4.6574 | 1.7132 | 2.6216 | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.4% | -1.1831 | -2.8879 | 8.6975 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 171 | SRM 3169 |  |  | 0.012322 | 0.0026225 | 0.0038751 |  | 0.0036929 | 1.3605e-05 | 0.00060027 |  |  | 1.6537e-06 | 1.6327e-05 | -7.5604e-07 |  |  | 43.8588 | 9.7402 | 15.179 |  | 15.9703 | 3.2411e-05 | 2.6631 |  |  | 5.3293e-06 | 1.999e-05 | 1.4474e-06 | 1 | 0 | 0 | 43.8465 | 9.7376 | 15.1752 | 0 | 15.9666 | 1.8806e-05 | 2.6625 | 0 | 0 | 3.6756e-06 | 3.6629e-06 | 2.2035e-06 | 0 | 0 | 1 | 1 | 1 | 1.7044 | 43.8465 | 9.7376 | 15.1751 | 15.9666 | 2.6625 | 0.22208 | 0.3461 | 0.36415 | 0.060723 | 1.5584 | 1.6397 | 0.27342 | 1.0522 | 0.17545 | 0.16675 | -1.5047 | -1.061 | -1.0102 | -2.8014 | [10 empty] | 0 | 1.7044 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014128 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 98.84% | 6.4259 | 7.1597 | -12.8352 | [8 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 99.99% | -10.6318 | -8.4965 | 3.8123 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 174 | bottle 3 |  |  | 0.012911 | 0.0027528 | 0.0040992 |  | 0.0039249 | 1.7378e-05 | 0.00063356 |  |  | 4.7437e-06 | 1.1294e-05 | 5.681e-06 |  |  | 47.3557 | 10.5166 | 16.3885 |  | 17.2422 | 3.154e-05 | 2.8751 |  |  | 2.6126e-06 | 1.5362e-05 | 4.32e-07 | 1 | 0 | 0 | 47.3428 | 10.5138 | 16.3844 | 0 | 17.2383 | 1.4162e-05 | 2.8745 | 0 | 0 | -2.1312e-06 | 4.0675e-06 | -5.249e-06 | 0 | 0 | 1 | 1 | 1 | 1.7023 | 47.3428 | 10.5138 | 16.3844 | 17.2383 | 2.8745 | 0.22208 | 0.34608 | 0.36412 | 0.060716 | 1.5584 | 1.6396 | 0.2734 | 1.0521 | 0.17544 | 0.16675 | -1.5047 | -1.0611 | -1.0103 | -2.8016 | [10 empty] | 0 | 1.7023 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43916 | -1.388 | 0.014134 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 107.1% | 3.3375 | -3.1097 | 1.0962 | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 99.92% | 2.1509 | 1.022 | -24.568 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 177 | SRM 3169 |  |  | 0.013327 | 0.0028481 | 0.0042432 |  | 0.004071 | -3.2324e-07 | 0.00064932 |  |  | 6.0583e-06 | 2.0699e-05 | 5.6577e-06 |  |  | 44.5528 | 9.8942 | 15.419 |  | 16.2228 | 2.2683e-05 | 2.7052 |  |  | 5.5866e-06 | 2.9686e-05 | 8.374e-06 | 1 | 0 | 0 | 44.5394 | 9.8914 | 15.4147 | 0 | 16.2187 | 2.3006e-05 | 2.7046 | 0 | 0 | -4.7163e-07 | 8.9865e-06 | 2.7163e-06 | 0 | 0 | 1 | 1 | 1 | 1.704 | 44.5394 | 9.8914 | 15.4147 | 16.2187 | 2.7046 | 0.22208 | 0.34609 | 0.36414 | 0.060723 | 1.5584 | 1.6397 | 0.27343 | 1.0522 | 0.17545 | 0.16676 | -1.5047 | -1.0611 | -1.0102 | -2.8014 | [10 empty] | 0 | 1.704 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014135 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.8% | -9.4802 | -8.0958 | 14.1968 | [8 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 96.97% | 10.9909 | -6.2046 | -2.6194 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 180 | bottle 4 |  |  | 0.012429 | 0.0026461 | 0.0039163 |  | 0.0037358 | 9.7936e-06 | 0.00059138 |  |  | 3.419e-06 | 1.2213e-05 | -4.121e-06 |  |  | 46.0492 | 10.2261 | 15.9356 |  | 16.7649 | 3.144e-05 | 2.7954 |  |  | -5.347e-07 | 3.0079e-05 | 7.5387e-07 | 1 | 0 | 0 | 46.0367 | 10.2235 | 15.9317 | 0 | 16.7612 | 2.1647e-05 | 2.7948 | 0 | 0 | -3.9537e-06 | 1.7866e-05 | 4.8749e-06 | 0 | 0 | 1 | 1 | 1 | 1.7002 | 46.0367 | 10.2235 | 15.9317 | 16.7612 | 2.7948 | 0.22207 | 0.34606 | 0.36408 | 0.060708 | 1.5583 | 1.6395 | 0.27337 | 1.0521 | 0.17542 | 0.16674 | -1.5048 | -1.0611 | -1.0104 | -2.8017 | [10 empty] | 0 | 1.7002 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014131 | -1.813 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 103.3% | -3.1022 | -0.11765 | -1.9069 | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.73% | -11.4982 | 2.1533 | -2.4676 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 183 | SRM 3169 |  |  | 0.012822 | 0.0027249 | 0.0040486 |  | 0.0038772 | 6.5622e-06 | 0.00062186 |  |  | 4.4011e-07 | 2.1862e-05 | -2.5857e-07 |  |  | 44.5632 | 9.8966 | 15.4224 |  | 16.226 | 3.7499e-05 | 2.7057 |  |  | 3.8456e-06 | 9.3197e-06 | 2.1638e-06 | 1 | 0 | 0 | 44.5504 | 9.8938 | 15.4184 | 0 | 16.2222 | 3.0937e-05 | 2.7051 | 0 | 0 | 3.4055e-06 | -1.2543e-05 | 2.4224e-06 | 0 | 0 | 1 | 1 | 1 | 1.7032 | 44.5504 | 9.8938 | 15.4184 | 16.2222 | 2.7051 | 0.22208 | 0.34609 | 0.36413 | 0.06072 | 1.5584 | 1.6396 | 0.27341 | 1.0521 | 0.17545 | 0.16675 | -1.5047 | -1.0611 | -1.0102 | -2.8015 | [10 empty] | 0 | 1.7032 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014127 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | 4.4031 | 4.3871 | 11.5719 | [8 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.48% | 5.4921 | -0.36523 | 2.8337 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 186 | bottle 5 |  |  | 0.012505 | 0.0026501 | 0.003927 |  | 0.0037391 | 6.7278e-06 | 0.00058331 |  |  | 3.9878e-06 | -6.6423e-06 | 3.9933e-06 |  |  | 46.1642 | 10.2516 | 15.9751 |  | 16.8059 | 3.3962e-05 | 2.8021 |  |  | 9.9575e-06 | 3.6182e-06 | 3.6399e-06 | 1 | 0 | 0 | 46.1517 | 10.249 | 15.9711 | 0 | 16.8022 | 2.7234e-05 | 2.8015 | 0 | 0 | 5.9697e-06 | 1.0261e-05 | -3.5344e-07 | 0 | 0 | 1 | 1 | 1 | 1.699 | 46.1517 | 10.249 | 15.9711 | 16.8021 | 2.8015 | 0.22207 | 0.34606 | 0.36406 | 0.060703 | 1.5583 | 1.6394 | 0.27335 | 1.052 | 0.17541 | 0.16674 | -1.5048 | -1.0612 | -1.0104 | -2.8018 | [10 empty] | 0 | 1.699 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16315 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014124 | -1.8131 | -1.8272 | [14 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 103.9% | 2.5111 | 2.583 | 1.7149 | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 96.72% | -5.0257 | 2.0391 | 6.3257 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 189 | SRM 3169 |  |  | 0.014165 | 0.0030261 | 0.0044871 |  | 0.0043343 | 4.14e-06 | 0.00070254 |  |  | 4.2387e-06 | 1.7486e-05 | -2.7453e-06 |  |  | 44.2946 | 9.8367 | 15.3287 |  | 16.1266 | 4.1625e-05 | 2.6889 |  |  | -3.494e-06 | 2.5274e-05 | 4.3186e-06 | 1 | 0 | 0 | 44.2805 | 9.8336 | 15.3242 | 0 | 16.1223 | 3.7485e-05 | 2.6882 | 0 | 0 | -7.7327e-06 | 7.7887e-06 | 7.064e-06 | 0 | 0 | 1 | 1 | 1 | 1.7009 | 44.2805 | 9.8336 | 15.3242 | 16.1222 | 2.6882 | 0.22208 | 0.34607 | 0.36409 | 0.060709 | 1.5583 | 1.6395 | 0.27337 | 1.0521 | 0.17542 | 0.16674 | -1.5047 | -1.0611 | -1.0103 | -2.8017 | [10 empty] | 0 | 1.7009 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014127 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.61% | -2.548 | -0.3192 | -22.3548 | [8 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.04% | 3.5086 | 0.41351 | -6.9756 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 192 | bottle 6 |  |  | 0.013721 | 0.0029304 | 0.0043367 |  | 0.0041582 | 1.6999e-05 | 0.00066685 |  |  | 2.4751e-06 | 5.5303e-06 | 5.6599e-06 |  |  | 45.884 | 10.1892 | 15.8773 |  | 16.7023 | 2.3799e-05 | 2.7847 |  |  | 3.218e-06 | 1.5459e-05 | 7.1709e-06 | 1 | 0 | 0 | 45.8703 | 10.1863 | 15.8729 | 0 | 16.6981 | 6.7996e-06 | 2.784 | 0 | 0 | 7.4292e-07 | 9.9284e-06 | 1.5111e-06 | 0 | 0 | 1 | 1 | 1 | 1.6968 | 45.8703 | 10.1863 | 15.8729 | 16.6981 | 2.784 | 0.22207 | 0.34604 | 0.36403 | 0.060693 | 1.5583 | 1.6393 | 0.27331 | 1.052 | 0.17539 | 0.16672 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6968 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.3881 | 0.014129 | -1.8131 | -1.8272 | [14 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 103.5% | 2.4295 | -2.0216 | -19.7319 | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.69% | -9.6749 | 4.3978 | 27.8727 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 195 | SRM 3169 |  |  | 0.013194 | 0.002797 | 0.0041546 |  | 0.0039667 | 9.8359e-06 | 0.00061212 |  |  | 3.19e-06 | 2.3413e-05 | -1.7388e-06 |  |  | 44.3701 | 9.8533 | 15.3542 |  | 16.1527 | 2.2502e-05 | 2.6932 |  |  | -4.5782e-07 | 1.4176e-05 | -6.7543e-07 | 1 | 0 | 0 | 44.3569 | 9.8505 | 15.35 | 0 | 16.1487 | 1.2666e-05 | 2.6926 | 0 | 0 | -3.6478e-06 | -9.2371e-06 | 1.0633e-06 | 0 | 0 | 1 | 1 | 1 | 1.699 | 44.3569 | 9.8505 | 15.35 | 16.1487 | 2.6926 | 0.22207 | 0.34606 | 0.36406 | 0.060703 | 1.5583 | 1.6394 | 0.27335 | 1.052 | 0.17541 | 0.16674 | -1.5048 | -1.0612 | -1.0104 | -2.8018 | [10 empty] | 0 | 1.699 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.014126 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | 3.9651 | -1.9528 | 19.521 | [8 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.54% | -0.85292 | -0.21458 | -7.5691 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 198 | bottle 1 |  |  | 0.013233 | 0.0028041 | 0.0041613 |  | 0.0039703 | -2.3378e-06 | 0.00061844 |  |  | 3.0002e-06 | 1.3708e-05 | 1.1324e-06 |  |  | 49.0187 | 10.8851 | 16.9616 |  | 17.8425 | 2.4986e-05 | 2.9747 |  |  | 4.9816e-06 | 1.4785e-05 | 9.2312e-07 | 1 | 0 | 0 | 49.0055 | 10.8823 | 16.9575 | 0 | 17.8385 | 2.7324e-05 | 2.9741 | 0 | 0 | 1.9814e-06 | 1.0764e-06 | -2.0931e-07 | 0 | 0 | 1 | 1 | 1 | 1.6956 | 49.0055 | 10.8823 | 16.9575 | 17.8385 | 2.9741 | 0.22206 | 0.34603 | 0.36401 | 0.060689 | 1.5583 | 1.6392 | 0.27329 | 1.052 | 0.17538 | 0.16672 | -1.5048 | -1.0612 | -1.0106 | -2.802 | [10 empty] | 0 | 1.6956 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014125 | -1.8131 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | May 28, 2020 | 1 | 110.6% | -3.7659 | -0.63668 | -8.1506 | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 98.01% | -9.5685 | -2.2838 | -10.9981 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 201 | SRM 3169 |  |  | 0.01297 | 0.0027582 | 0.0040786 |  | 0.0038798 | 1.6703e-05 | 0.00060205 |  |  | 1.7204e-06 | 1.9296e-05 | 5.1769e-06 |  |  | 44.2896 | 9.8352 | 15.3259 |  | 16.1223 | 1.9605e-05 | 2.688 |  |  | 3.2273e-06 | 2.1974e-05 | 6.0111e-06 | 1 | 0 | 0 | 44.2767 | 9.8324 | 15.3218 | 0 | 16.1184 | 2.9024e-06 | 2.6874 | 0 | 0 | 1.507e-06 | 2.6779e-06 | 8.3413e-07 | 0 | 0 | 1 | 1 | 1 | 1.6975 | 44.2767 | 9.8324 | 15.3218 | 16.1184 | 2.6874 | 0.22207 | 0.34605 | 0.36404 | 0.060696 | 1.5583 | 1.6393 | 0.27332 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6975 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014122 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.85% | -2.5014 | 3.1512 | -11.7129 | [8 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.09% | -4.1282 | -1.4992 | 16.3688 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 204 | bottle 2 |  |  | 0.013293 | 0.0028177 | 0.0041759 |  | 0.0039801 | 1.5664e-05 | 0.00064116 |  |  | 2.4392e-06 | 1.9407e-05 | -6.9504e-07 |  |  | 47.7243 | 10.5977 | 16.5135 |  | 17.3707 | 3.784e-05 | 2.896 |  |  | 8.0073e-06 | 2.9362e-05 | 4.6396e-06 | 1 | 0 | 0 | 47.711 | 10.5949 | 16.5093 | 0 | 17.3667 | 2.2176e-05 | 2.8953 | 0 | 0 | 5.5681e-06 | 9.9546e-06 | 5.3346e-06 | 0 | 0 | 1 | 1 | 1 | 1.6948 | 47.711 | 10.5949 | 16.5093 | 17.3667 | 2.8953 | 0.22206 | 0.34603 | 0.364 | 0.060684 | 1.5582 | 1.6392 | 0.27327 | 1.0519 | 0.17537 | 0.16672 | -1.5048 | -1.0612 | -1.0106 | -2.8021 | [10 empty] | 0 | 1.6948 | 0.21794 | 0.33336 | 0.3381 | 0.054387 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43914 | -1.3881 | 0.014121 | -1.8131 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | May 28, 2020 | 1 | 107.7% | 8.3517 | 2.163 | -17.5245
Lab 2 | May 28, 2020 | 1 | 207 | SRM 3169 |  |  | 0.013813 | 0.0029453 | 0.00437 |  | 0.0041726 | 1.916e-05 | 0.00066273 |  |  | 2.0148e-06 | 1.8685e-05 | 4.4374e-06 |  |  | 44.3403 | 9.8464 | 15.3431 |  | 16.1402 | 1.8743e-05 | 2.691 |  |  | -1.8229e-06 | 1.1335e-05 | -1.8074e-06 | 1 | 0 | 0 | 44.3265 | 9.8435 | 15.3387 | 0 | 16.1361 | -4.1705e-07 | 2.6903 | 0 | 0 | -3.8377e-06 | -7.3503e-06 | -6.2448e-06 | 0 | 0 | 1 | 1 | 1 | 1.6967 | 44.3265 | 9.8435 | 15.3388 | 16.1361 | 2.6903 | 0.22207 | 0.34604 | 0.36403 | 0.060693 | 1.5583 | 1.6393 | 0.27331 | 1.052 | 0.17539 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6967 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.014124 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.3% | 2.533 | -0.49612 | 0.063121 | [8 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 100.6% | -10.6541 | -10.058 | 3.7301 | SRM 3169 | Avg | 1.4727 | 0.65789 | -11.7489
Lab 2 | May 28, 2020 | 1 | 210 | bottle 3 |  |  | 0.013565 | 0.0028821 | 0.0042647 |  | 0.0040632 | -2.7713e-06 | 0.00065288 |  |  | 4.5703e-06 | 9.3702e-06 | 2.593e-07 |  |  | 48.1662 | 10.6955 | 16.6654 |  | 17.5296 | 2.7691e-05 | 2.9223 |  |  | 1.0749e-06 | 2.9004e-05 | 2.9866e-06 | 1 | 0 | 0 | 48.1526 | 10.6926 | 16.6612 | 0 | 17.5255 | 3.0462e-05 | 2.9217 | 0 | 0 | -3.4954e-06 | 1.9634e-05 | 2.7273e-06 | 0 | 0 | 1 | 1 | 1 | 1.6923 | 48.1526 | 10.6926 | 16.6611 | 17.5255 | 2.9216 | 0.22206 | 0.34601 | 0.36396 | 0.060674 | 1.5582 | 1.639 | 0.27324 | 1.0519 | 0.17536 | 0.16671 | -1.5048 | -1.0613 | -1.0107 | -2.8022 | [10 empty] | 0 | 1.6923 | 0.21793 | 0.33336 | 0.3381 | 0.054387 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.3881 | 0.014125 | -1.8131 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 107.8% | -0.26196 | -0.46977 | -22.5545 | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 102% | 11.2114 | 1.1943 | -7.539 | SRM 3169 | 2SD | 21.0768 | 14.14 | 41.5091
Lab 2 | May 28, 2020 | 1 | 213 | SRM 3169 |  |  | 0.017203 | 0.003689 | 0.0055254 |  | 0.0054121 | 1.6088e-05 | 0.00086601 |  |  | 6.7846e-06 | 1.8049e-05 | -1.0044e-06 |  |  | 45.0221 | 9.9979 | 15.5795 |  | 16.3894 | 4.1274e-05 | 2.7326 |  |  | -4.1667e-06 | 1.1127e-05 | 2.2927e-06 | 1 | 0 | 0 | 45.0049 | 9.9942 | 15.574 | 0 | 16.384 | 2.5186e-05 | 2.7317 | 0 | 0 | -1.0951e-05 | -6.9229e-06 | 3.2971e-06 | 0 | 0 | 1 | 1 | 1 | 1.6981 | 45.0049 | 9.9942 | 15.574 | 16.384 | 2.7317 | 0.22207 | 0.34605 | 0.36405 | 0.060699 | 1.5583 | 1.6393 | 0.27333 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8018 | [10 empty] | 0 | 1.6981 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014125 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.4% | -1.5296 | -0.29484 | -2.4415 | [8 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.45% | 4.1437 | -5.9899 | 11.9243 | SRM 3169 | n | 18 | 18 | 18
Lab 2 | May 28, 2020 | 1 | 216 | bottle 4 |  |  | 0.013663 | 0.0028961 | 0.0043137 |  | 0.0041338 | 1.0463e-05 | 0.00065756 |  |  | 4.1678e-06 | 2.4579e-05 | 2.0013e-06 |  |  | 47.0341 | 10.4446 | 16.2753 |  | 17.1207 | 2.9884e-05 | 2.8544 |  |  | 7.0447e-07 | 1.4686e-05 | -5.9124e-07 | 1 | 0 | 0 | 47.0205 | 10.4417 | 16.2709 | 0 | 17.1165 | 1.9421e-05 | 2.8537 | 0 | 0 | -3.4633e-06 | -9.8931e-06 | -2.5926e-06 | 0 | 0 | 1 | 1 | 1 | 1.6964 | 47.0205 | 10.4417 | 16.271 | 17.1165 | 2.8538 | 0.22207 | 0.34604 | 0.36402 | 0.060692 | 1.5583 | 1.6393 | 0.27331 | 1.052 | 0.17539 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6964 | 0.21793 | 0.33336 | 0.3381 | 0.054388 | 1.5296 | 1.5514 | 0.24956 | 1.0142 | 0.16315 | 0.16086 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014119 | -1.8131 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 104.1% | 1.7548 | 6.1889 | -6.3245 | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.68% | 4.2918 | 3.3048 | 8.1331 | SRM 3169 | 2SE | 4.9678 | 3.3328 | 9.7838
Lab 2 | May 28, 2020 | 1 | 219 | SRM 3169 |  |  | 0.013638 | 0.0029013 | 0.0043253 |  | 0.0041225 | 1.6433e-05 | 0.00065453 |  |  | 9.8241e-06 | 2.6452e-05 | 1.8241e-06 |  |  | 45.3116 | 10.0624 | 15.6804 |  | 16.4963 | 2.9126e-05 | 2.7505 |  |  | -1.0656e-06 | 1.6684e-05 | -1.6514e-06 | 1 | 0 | 0 | 45.2979 | 10.0595 | 15.6761 | 0 | 16.4921 | 1.2693e-05 | 2.7499 | 0 | 0 | -1.089e-05 | -9.7676e-06 | -3.4754e-06 | 0 | 0 | 1 | 1 | 1 | 1.7001 | 45.2979 | 10.0595 | 15.6761 | 16.4921 | 2.7499 | 0.22208 | 0.34607 | 0.36408 | 0.060707 | 1.5583 | 1.6395 | 0.27336 | 1.0521 | 0.17542 | 0.16674 | -1.5047 | -1.0611 | -1.0104 | -2.8017 | [10 empty] | 0 | 1.7001 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16315 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.014125 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100% | 1.4732 | 0.62772 | 5.1002 | [8 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 101.9% | 23.5368 | 0.63457 | 15.7721 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 222 | bottle 5 |  |  | 0.014633 | 0.0031195 | 0.0046358 |  | 0.0044514 | 1.4837e-05 | 0.00071438 |  |  | -2.5975e-06 | 1.8116e-05 | -2.7476e-06 |  |  | 47.8151 | 10.618 | 16.5457 |  | 17.4055 | 1.6985e-05 | 2.902 |  |  | 3.9209e-06 | 1.4342e-06 | 4.9411e-06 | 1 | 0 | 0 | 47.8004 | 10.6149 | 16.5411 | 0 | 17.4011 | 2.1486e-06 | 2.9013 | 0 | 0 | 6.5185e-06 | -1.6682e-05 | 7.6888e-06 | 0 | 0 | 1 | 1 | 1 | 1.6973 | 47.8004 | 10.6149 | 16.5411 | 17.4011 | 2.9013 | 0.22207 | 0.34604 | 0.36404 | 0.060697 | 1.5583 | 1.6393 | 0.27333 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6973 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014124 | -1.813 | -1.8272 | [14 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 105.2% | -6.8426 | 1.6624 | 10.8917 | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.49% | -4.9096 | -13.2281 | -1.9245 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 225 | SRM 3169 |  |  | 0.013372 | 0.0028503 | 0.0041985 |  | 0.0040174 | 1.6835e-05 | 0.00063425 |  |  | 4.1337e-06 | 8.1202e-06 | 3.7643e-07 |  |  | 45.5624 | 10.1181 | 15.7672 |  | 16.5877 | 2.3802e-05 | 2.7658 |  |  | 9.8e-06 | 2.1969e-05 | 7.9295e-07 | 1 | 0 | 0 | 45.549 | 10.1153 | 15.763 | 0 | 16.5837 | 6.967e-06 | 2.7652 | 0 | 0 | 5.6663e-06 | 1.3849e-05 | 4.1652e-07 | 0 | 0 | 1 | 1 | 1 | 1.7004 | 45.549 | 10.1153 | 15.763 | 16.5837 | 2.7652 | 0.22208 | 0.34607 | 0.36408 | 0.060708 | 1.5583 | 1.6395 | 0.27337 | 1.0521 | 0.17542 | 0.16674 | -1.5047 | -1.0611 | -1.0104 | -2.8017 | [10 empty] | 0 | 1.7003 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16315 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.014127 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.3% | -2.6095 | 0.70076 | -2.1874 | [8 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 100.3% | -2.005 | 1.0024 | -0.24635 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 228 | bottle 6 |  |  | 0.01744 | 0.0037513 | 0.00563 |  | 0.0054831 | 2.3183e-05 | 0.00088138 |  |  | 4.5163e-06 | 2.5558e-05 | 5.5647e-06 |  |  | 47.5097 | 10.5501 | 16.4398 |  | 17.2939 | 2.4276e-05 | 2.8833 |  |  | 2.2302e-06 | 1.1447e-05 | -2.6491e-07 | 1 | 0 | 0 | 47.4922 | 10.5464 | 16.4342 | 0 | 17.2884 | 1.0932e-06 | 2.8825 | 0 | 0 | -2.2861e-06 | -1.4111e-05 | -5.8296e-06 | 0 | 0 | 1 | 1 | 1 | 1.6966 | 47.4922 | 10.5464 | 16.4342 | 17.2884 | 2.8825 | 0.22207 | 0.34604 | 0.36403 | 0.060694 | 1.5583 | 1.6393 | 0.27332 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6967 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43915 | -1.388 | 0.014125 | -1.8131 | -1.8272 | [14 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 104.3% | -4.4856 | 3.7589 | 8.4772 | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 101.8% | -2.8513 | 8.5085 | -10.3395 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 231 | SRM 3169 |  |  | 0.013512 | 0.0028573 | 0.0042651 |  | 0.0040444 | 1.4743e-05 | 0.00064564 |  |  | -6.2632e-07 | 2.21e-05 | 8.2068e-06 |  |  | 45.5563 | 10.1167 | 15.7648 |  | 16.5849 | 2.8379e-05 | 2.7653 |  |  | -3.4225e-06 | 1.8603e-05 | -1.5807e-06 | 1 | 0 | 0 | 45.5428 | 10.1139 | 15.7605 | 0 | 16.5809 | 1.3636e-05 | 2.7647 | 0 | 0 | -2.7962e-06 | -3.4972e-06 | -9.7875e-06 | 0 | 0 | 1 | 1 | 1 | 1.6995 | 45.5428 | 10.1139 | 15.7605 | 16.5809 | 2.7647 | 0.22207 | 0.34606 | 0.36407 | 0.060705 | 1.5583 | 1.6394 | 0.27335 | 1.0521 | 0.17542 | 0.16674 | -1.5047 | -1.0611 | -1.0104 | -2.8017 | [10 empty] | 0 | 1.6995 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.01413 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | 3.3505 | -1.3965 | 2.9983 | [8 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.51% | -0.038047 | 12.9627 | -14.124 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 234 | bottle 1 |  |  | 0.014373 | 0.003065 | 0.0045415 |  | 0.0043498 | 1.0377e-05 | 0.00068529 |  |  | 6.2944e-06 | 9.49e-06 | 2.9403e-06 |  |  | 48.5804 | 10.788 | 16.8105 |  | 17.6842 | 3.8631e-05 | 2.9485 |  |  | 2.7402e-06 | 1.9185e-05 | 1.1576e-06 | 1 | 0 | 0 | 48.5661 | 10.7849 | 16.806 | 0 | 17.6798 | 2.8254e-05 | 2.9478 | 0 | 0 | -3.5542e-06 | 9.6953e-06 | -1.7828e-06 | 0 | 0 | 1 | 1 | 1 | 1.6973 | 48.5661 | 10.7849 | 16.8059 | 17.6798 | 2.9478 | 0.22207 | 0.34604 | 0.36404 | 0.060696 | 1.5583 | 1.6393 | 0.27332 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6973 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16315 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014129 | -1.8131 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | May 28, 2020 | 1 | 106.8% | -7.2807 | 1.1397 | 0.41744 | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | -5.4111 | 7.6793 | -19.1403 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 237 | SRM 3169 |  |  | 0.013746 | 0.0029152 | 0.0043203 |  | 0.0041238 | 2.261e-06 | 0.000656 |  |  | 7.3067e-06 | 1.928e-05 | 7.1315e-07 |  |  | 45.3775 | 10.0769 | 15.7026 |  | 16.5193 | 2.314e-05 | 2.7543 |  |  | 3.1305e-06 | 2.9205e-05 | 4.8535e-06 | 1 | 0 | 0 | 45.3637 | 10.074 | 15.6983 | 0 | 16.5151 | 2.0879e-05 | 2.7537 | 0 | 0 | -4.1762e-06 | 9.9252e-06 | 4.1404e-06 | 0 | 0 | 1 | 1 | 1 | 1.6988 | 45.3637 | 10.074 | 15.6983 | 16.5151 | 2.7536 | 0.22207 | 0.34605 | 0.36406 | 0.060701 | 1.5583 | 1.6394 | 0.27334 | 1.052 | 0.17541 | 0.16673 | -1.5048 | -1.0612 | -1.0104 | -2.8018 | [10 empty] | 0 | 1.6988 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16315 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.01413 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.17% | 5.4825 | 3.2857 | -17.9386 | [8 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.6% | 8.7377 | 2.8308 | 0.8286 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 240 | bottle 2 |  |  | 0.013884 | 0.0029418 | 0.0043669 |  | 0.0041884 | 1.517e-05 | 0.00066242 |  |  | 3.4057e-06 | 1.3395e-05 | 7.6773e-07 |  |  | 47.7449 | 10.6024 | 16.5213 |  | 17.3798 | 2.5553e-05 | 2.8977 |  |  | -3.5798e-06 | 1.2963e-05 | -1.296e-06 | 1 | 0 | 0 | 47.731 | 10.5995 | 16.5169 | 0 | 17.3756 | 1.0384e-05 | 2.897 | 0 | 0 | -6.9855e-06 | -4.3154e-07 | -2.0637e-06 | 0 | 0 | 1 | 1 | 1 | 1.697 | 47.731 | 10.5995 | 16.5169 | 17.3756 | 2.897 | 0.22207 | 0.34604 | 0.36403 | 0.060695 | 1.5583 | 1.6393 | 0.27332 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.697 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16315 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42502 | 0.43915 | -1.388 | 0.014128 | -1.8131 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | May 28, 2020 | 1 | 104.6% | 9.4189 | 5.6377 | -15.3036 | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.8% | -6.6617 | 1.6107 | -44.6196 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 243 | SRM 3169 |  |  | 0.013016 | 0.0027583 | 0.0041093 |  | 0.0039084 | 2.2684e-05 | 0.00062856 |  |  | 7.402e-06 | 1.5248e-05 | -2.8705e-06 |  |  | 45.9567 | 10.205 | 15.902 |  | 16.7281 | 4.4165e-05 | 2.789 |  |  | -1.8686e-06 | 1.0835e-05 | -5.3721e-07 | 1 | 0 | 0 | 45.9436 | 10.2023 | 15.8979 | 0 | 16.7242 | 2.1481e-05 | 2.7884 | 0 | 0 | -9.2706e-06 | -4.4131e-06 | 2.3333e-06 | 0 | 0 | 1 | 1 | 1 | 1.6959 | 45.9436 | 10.2023 | 15.8979 | 16.7242 | 2.7884 | 0.22206 | 0.34603 | 0.36402 | 0.060692 | 1.5583 | 1.6393 | 0.27331 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0106 | -2.8019 | [10 empty] | 0 | 1.6959 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014137 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.4% | -6.7098 | -3.1289 | 11.3563 | [8 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.4% | -22.3547 | -3.2295 | -25.9543 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 246 | bottle 3 |  |  | 0.01304 | 0.0027625 | 0.0040968 |  | 0.0039159 | 1.7739e-05 | 0.00061706 |  |  | -2.5513e-06 | 1.7464e-05 | 4.1456e-06 |  |  | 48.7319 | 10.8212 | 16.8621 |  | 17.738 | 2.3235e-05 | 2.9574 |  |  | 4.4932e-06 | 2.8457e-05 | -4.7964e-06 | 1 | 0 | 0 | 48.7188 | 10.8184 | 16.858 | 0 | 17.734 | 5.4964e-06 | 2.9568 | 0 | 0 | 7.0445e-06 | 1.0993e-05 | -8.942e-06 | 0 | 0 | 1 | 1 | 1 | 1.6955 | 48.7188 | 10.8184 | 16.858 | 17.734 | 2.9567 | 0.22206 | 0.34603 | 0.36401 | 0.06069 | 1.5583 | 1.6392 | 0.27331 | 1.052 | 0.17539 | 0.16673 | -1.5048 | -1.0612 | -1.0106 | -2.802 | [10 empty] | 0 | 1.6955 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014137 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | May 28, 2020 | 1 | 105.8% | -3.5671 | 0.36716 | -11.7904 | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.3% | 6.7257 | -5.7961 | -51.8264 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 249 | SRM 3169 |  |  | 0.013333 | 0.0028421 | 0.0041917 |  | 0.0039958 | 4.0023e-06 | 0.0006301 |  |  | 7.264e-07 | 2.1021e-05 | 4.1446e-06 |  |  | 46.1823 | 10.2553 | 15.9805 |  | 16.8113 | 1.9842e-05 | 2.803 |  |  | 1.6938e-06 | 2.0185e-05 | -2.0004e-06 | 1 | 0 | 0 | 46.1689 | 10.2525 | 15.9763 | 0 | 16.8073 | 1.584e-05 | 2.8024 | 0 | 0 | 9.6739e-07 | -8.364e-07 | -6.145e-06 | 0 | 0 | 1 | 1 | 1 | 1.6974 | 46.1689 | 10.2525 | 15.9763 | 16.8073 | 2.8024 | 0.22206 | 0.34604 | 0.36404 | 0.060699 | 1.5583 | 1.6393 | 0.27334 | 1.052 | 0.17541 | 0.16674 | -1.5048 | -1.0612 | -1.0105 | -2.8018 | [10 empty] | 0 | 1.6974 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014138 | -1.813 | -1.8271 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | -6.9724 | -4.1765 | 23.0857 | [8 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.3% | 2.8183 | 5.2652 | 5.0518 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 252 | bottle 4 |  |  | 0.015119 | 0.0032353 | 0.0048184 |  | 0.0046451 | 2.2344e-06 | 0.00074635 |  |  | 2.3887e-06 | 3.1373e-05 | 4.3402e-06 |  |  | 48.1378 | 10.6892 | 16.6561 |  | 17.5207 | 3.21e-05 | 2.921 |  |  | -1.0199e-06 | 7.0181e-06 | 2.7204e-06 | 1 | 0 | 0 | 48.1227 | 10.686 | 16.6513 | 0 | 17.516 | 2.9866e-05 | 2.9203 | 0 | 0 | -3.4086e-06 | -2.4355e-05 | -1.6198e-06 | 0 | 0 | 1 | 1 | 1 | 1.6941 | 48.1227 | 10.686 | 16.6513 | 17.516 | 2.9203 | 0.22206 | 0.34602 | 0.36399 | 0.060685 | 1.5582 | 1.6392 | 0.27329 | 1.0519 | 0.17538 | 0.16672 | -1.5048 | -1.0613 | -1.0106 | -2.8021 | [10 empty] | 0 | 1.6942 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014132 | -1.813 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | May 28, 2020 | 1 | 104.2% | -2.2537 | 1.2503 | 8.6686 | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 99.23% | 2.4161 | -7.9842 | -33.4991 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 255 | SRM 3169 |  |  | 0.014551 | 0.0030973 | 0.0046073 |  | 0.0044421 | 1.3238e-05 | 0.00070878 |  |  | 1.1539e-06 | 2.7174e-06 | 2.9208e-07 |  |  | 46.2405 | 10.2684 | 16.0008 |  | 16.8325 | 3.7837e-05 | 2.8065 |  |  | -7.2122e-08 | 1.7676e-05 | 1.8187e-06 | 1 | 0 | 0 | 46.2259 | 10.2653 | 15.9962 | 0 | 16.8281 | 2.4599e-05 | 2.8058 | 0 | 0 | -1.226e-06 | 1.4959e-05 | 1.5266e-06 | 0 | 0 | 1 | 1 | 1 | 1.6975 | 46.2259 | 10.2653 | 15.9962 | 16.8281 | 2.8058 | 0.22207 | 0.34604 | 0.36404 | 0.060697 | 1.5583 | 1.6393 | 0.27333 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6975 | 0.21793 | 0.33336 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.01413 | -1.8131 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 99.82% | 10.9992 | 7.2515 | -27.266 | [8 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 99.14% | -1.241 | 6.6038 | -1.8746 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 258 | bottle 5 |  |  | 0.014947 | 0.0031809 | 0.0047487 |  | 0.0045638 | 2.3845e-06 | 0.00072757 |  |  | -9.988e-07 | 1.3188e-05 | 1.657e-07 |  |  | 48.6473 | 10.8023 | 16.8324 |  | 17.706 | 2.6872e-05 | 2.952 |  |  | 5.8409e-06 | 1.1027e-05 | 2.7271e-06 | 1 | 0 | 0 | 48.6323 | 10.7992 | 16.8276 | 0 | 17.7015 | 2.4487e-05 | 2.9512 | 0 | 0 | 6.8397e-06 | -2.1601e-06 | 2.5614e-06 | 0 | 0 | 1 | 1 | 1 | 1.6941 | 48.6323 | 10.7992 | 16.8276 | 17.7015 | 2.9512 | 0.22206 | 0.34602 | 0.36399 | 0.060685 | 1.5582 | 1.6392 | 0.27328 | 1.0519 | 0.17538 | 0.16672 | -1.5048 | -1.0613 | -1.0106 | -2.8021 | [10 empty] | 0 | 1.6941 | 0.21793 | 0.33336 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014133 | -1.813 | -1.8272 | [14 empty] | bottle 5 | Lab 2 | May 28, 2020 | 1 | 105% | -2.9715 | 0.25743 | 11.9319 | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100% | 18.7534 | 6.5304 | -45.8332 | SRM 3169
Lab 2 | May 28, 2020 | 1 | 261 | SRM 3169 |  |  | 0.014629 | 0.0031253 | 0.0046268 |  | 0.0044477 | 8.53e-06 | 0.00071444 |  |  | 5.8277e-06 | 1.4229e-05 | 2.1222e-06 |  |  | 46.4629 | 10.3176 | 16.0774 |  | 16.913 | 3.2505e-05 | 2.8199 |  |  | 4.1986e-06 | 1.6734e-05 | -8.3118e-07 | 1 | 0 | 0 | 46.4483 | 10.3145 | 16.0728 | 0 | 16.9085 | 2.3975e-05 | 2.8192 | 0 | 0 | -1.6291e-06 | 2.5049e-06 | -2.9534e-06 | 0 | 0 | 1 | 1 | 1 | 1.6968 | 46.4483 | 10.3145 | 16.0728 | 16.9085 | 2.8192 | 0.22206 | 0.34604 | 0.36403 | 0.060695 | 1.5583 | 1.6393 | 0.27332 | 1.052 | 0.1754 | 0.16673 | -1.5048 | -1.0612 | -1.0105 | -2.8019 | [10 empty] | 0 | 1.6968 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43916 | -1.388 | 0.014136 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | May 28, 2020 | 1 | 100.2% | -3.1984 | -1.0389 | 0.4553
Lab 2 | May 28, 2020 | 1 | 264 | bottle 6 |  |  | 0.013766 | 0.0029293 | 0.0043374 |  | 0.0041438 | 1.3986e-05 | 0.0006531 |  |  | 4.344e-06 | 1.259e-05 | -2.9262e-06 |  |  | 48.5088 | 10.7714 | 16.784 |  | 17.6548 | 4.0942e-05 | 2.9434 |  |  | 2.1843e-06 | 6.8118e-06 | -4.1288e-06 | 1 | 0 | 0 | 48.495 | 10.7685 | 16.7797 | 0 | 17.6506 | 2.6956e-05 | 2.9427 | 0 | 0 | -2.1596e-06 | -5.7778e-06 | -1.2026e-06 | 0 | 0 | 1 | 1 | 1 | 1.693 | 48.495 | 10.7685 | 16.7797 | 17.6506 | 2.9427 | 0.22205 | 0.34601 | 0.36397 | 0.060681 | 1.5582 | 1.6391 | 0.27327 | 1.0519 | 0.17537 | 0.16672 | -1.5048 | -1.0613 | -1.0107 | -2.8021 | [10 empty] | 0 | 1.693 | 0.21793 | 0.33336 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0985 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014133 | -1.813 | -1.8271 | [14 empty] | bottle 6 | Lab 2 | May 28, 2020 | 1 | 104.4% | -0.2295 | 5.4971 | 7.9089 | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.9% | 4.4623 | 0.22932 | 12.3283 | SRM 3169 | Avg | -2.9774 | 0.87742 | 0.72423
Lab 2 | May 28, 2020 | 1 | 267 | SRM 3169 |  |  | 0.013686 | 0.0029151 | 0.0043167 |  | 0.0041137 | 7.9588e-06 | 0.00066018 |  |  | 6.4535e-06 | 6.2305e-06 | 7.9768e-06 |  |  | 46.4952 | 10.3245 | 16.0878 |  | 16.9232 | 2.4537e-05 | 2.8215 |  |  | 5.9745e-06 | 1.7124e-05 | 3.5304e-06 | 1 | 0 | 0 | 46.4815 | 10.3216 | 16.0835 | 0 | 16.9191 | 1.6578e-05 | 2.8209 | 0 | 0 | -4.7902e-07 | 1.0894e-05 | -4.4464e-06 | 0 | 0 | 1 | 1 | 1 | 1.6947 | 46.4815 | 10.3216 | 16.0835 | 16.9191 | 2.8209 | 0.22206 | 0.34602 | 0.364 | 0.060688 | 1.5582 | 1.6392 | 0.2733 | 1.052 | 0.17539 | 0.16673 | -1.5048 | -1.0613 | -1.0106 | -2.802 | [10 empty] | 0 | 1.6947 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42503 | 0.43917 | -1.388 | 0.014141 | -1.813 | -1.8271 | [22 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.7% | -8.2372 | -2.1133 | -43.6591 | SRM 3169 | 2SD | 10.405 | 6.5866 | 51.3135
Lab 2 | June 3, 2020 | 2 | 68 | SRM 3169 |  |  | 0.010393 | 0.0022113 | 0.0032714 |  | 0.0031032 | 1.6519e-05 | 0.00050336 |  |  | 5.4354e-06 | 6.1762e-06 | 1.5294e-06 |  |  | 41.2112 | 9.1486 | 14.251 |  | 14.9823 | 2.6797e-05 | 2.4964 |  |  | -8.5739e-07 | 1.0761e-05 | 5.3731e-06 | 1 | 0 | 0 | 41.2008 | 9.1464 | 14.2478 | 0 | 14.9792 | 1.0278e-05 | 2.4959 | 0 | 0 | -6.2928e-06 | 4.5851e-06 | 3.8437e-06 | 0 | 0 | 1 | 1 | 1 | 1.6676 | 41.2008 | 9.1464 | 14.2478 | 14.9792 | 2.4959 | 0.22199 | 0.34581 | 0.36357 | 0.06058 | 1.5578 | 1.6377 | 0.27289 | 1.0513 | 0.17518 | 0.16663 | -1.5051 | -1.0619 | -1.0118 | -2.8038 | [9 empty] | L2 Jun | 0 | 1.6676 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014141 | -1.813 | -1.8272 | [22 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -7.3172 | 1.2003 | 12.4733 | SRM 3169 | n | 5 | 5 | 5
Lab 2 | June 3, 2020 | 2 | 71 | SRM 3169 |  |  | 0.01077 | 0.0022981 | 0.0033915 |  | 0.0032251 | 1.844e-05 | 0.00052105 |  |  | 1.0982e-05 | 1.3573e-05 | 1.7949e-06 |  |  | 41.2662 | 9.1607 | 14.2696 |  | 15.0014 | 2.6976e-05 | 2.4996 |  |  | 7.8046e-06 | 9.6842e-06 | -3.4589e-07 | 1 | 0 | 0 | 41.2554 | 9.1584 | 14.2662 | 0 | 14.9982 | 8.536e-06 | 2.499 | 0 | 0 | -3.1775e-06 | -3.8892e-06 | -2.1408e-06 | 0 | 0 | 1 | 1 | 1 | 1.6662 | 41.2554 | 9.1584 | 14.2662 | 14.9982 | 2.499 | 0.22199 | 0.3458 | 0.36354 | 0.060575 | 1.5577 | 1.6376 | 0.27287 | 1.0513 | 0.17517 | 0.16662 | -1.5051 | -1.0619 | -1.0119 | -2.8039 | [10 empty] | 0 | 1.6662 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014144 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 100.2% | 5.0631 | -3.0038 | -0.75686 | [8 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.4% | -3.3261 | 6.3048 | 20.4419 | SRM 3169 | 2SE | 4.6533 | 2.9456 | 22.9481
Lab 2 | June 3, 2020 | 2 | 74 | SRM 3169 |  |  | 0.010915 | 0.0023246 | 0.0034155 |  | 0.003257 | -4.3237e-06 | 0.00051104 |  |  | -1.3739e-06 | 2.1859e-05 | -3.0479e-07 |  |  | 41.1853 | 9.1426 | 14.2414 |  | 14.9715 | 2.1761e-05 | 2.4945 |  |  | 1.0999e-06 | 4.2695e-06 | 6.1206e-07 | 1 | 0 | 0 | 41.1744 | 9.1402 | 14.2379 | 0 | 14.9682 | 2.6085e-05 | 2.494 | 0 | 0 | 2.4738e-06 | -1.759e-05 | 9.1685e-07 | 0 | 0 | 1 | 1 | 1 | 1.6655 | 41.1744 | 9.1402 | 14.238 | 14.9682 | 2.4941 | 0.22199 | 0.3458 | 0.36353 | 0.060573 | 1.5577 | 1.6376 | 0.27286 | 1.0513 | 0.17517 | 0.16662 | -1.5051 | -1.0619 | -1.0119 | -2.8039 | [10 empty] | 0 | 1.6655 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014141 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.21% | -5.955 | 6.2986 | 15.9334 | [8 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -0.46873 | -1.234 | 2.0367 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 77 | Bottle 1 |  |  | 0.011145 | 0.0023656 | 0.0035143 |  | 0.0033322 | 2.0451e-06 | 0.0005335 |  |  | -1.5848e-06 | 1.3938e-06 | 5.2852e-06 |  |  | 43.8075 | 9.7241 | 15.1465 |  | 15.9214 | 2.7098e-05 | 2.6525 |  |  | -6.4384e-09 | 1.1232e-05 | -3.9101e-06 | 1 | 0 | 0 | 43.7963 | 9.7218 | 15.1429 | 0 | 15.9181 | 2.5053e-05 | 2.652 | 0 | 0 | 1.5783e-06 | 9.8382e-06 | -9.1953e-06 | 0 | 0 | 1 | 1 | 1 | 1.6607 | 43.7963 | 9.7218 | 15.1429 | 15.918 | 2.652 | 0.22198 | 0.34576 | 0.36346 | 0.060553 | 1.5576 | 1.6374 | 0.27279 | 1.0512 | 0.17513 | 0.1666 | -1.5052 | -1.062 | -1.0121 | -2.8042 | [10 empty] | 0 | 1.6607 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014147 | -1.813 | -1.8272 | [14 empty] | Bottle 1 | Lab 2 | June 3, 2020 | 2 | 105.6% | -0.6633 | -0.18269 | -11.5675
Lab 2 | June 3, 2020 | 2 | 80 | Bottle 2 |  |  | 0.011728 | 0.0024956 | 0.003696 |  | 0.0035367 | 9.3863e-06 | 0.00057091 |  |  | 6.738e-06 | 2.0093e-05 | -4.6513e-06 |  |  | 43.6731 | 9.6942 | 15.0994 |  | 15.8714 | 3.327e-05 | 2.6441 |  |  | 2.7444e-06 | -3.2302e-06 | 7.3748e-07 | 1 | 0 | 0 | 43.6613 | 9.6917 | 15.0957 | 0 | 15.8678 | 2.3884e-05 | 2.6435 | 0 | 0 | -3.9935e-06 | -2.3323e-05 | 5.3888e-06 | 0 | 0 | 1 | 1 | 1 | 1.659 | 43.6613 | 9.6917 | 15.0957 | 15.8679 | 2.6436 | 0.22197 | 0.34575 | 0.36343 | 0.060548 | 1.5576 | 1.6373 | 0.27277 | 1.0511 | 0.17512 | 0.1666 | -1.5052 | -1.0621 | -1.0122 | -2.8043 | [10 empty] | 0 | 1.659 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014145 | -1.813 | -1.8272 | [14 empty] | Bottle 2 | Lab 2 | June 3, 2020 | 2 | 105.3% | 5.9912 | 0.96502 | 7.093 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.48% | 0.74303 | -11.9217 | -10.3052 | SRM 3169 | Avg | -0.64855 | -0.60179 | -2.3274
Lab 2 | June 3, 2020 | 2 | 83 | Bottle 3 |  |  | 0.012461 | 0.0026583 | 0.0039669 |  | 0.0037883 | 1.1236e-05 | 0.00060319 |  |  | 3.5492e-06 | 1.7224e-05 | -4.5236e-06 |  |  | 43.2931 | 9.6095 | 14.9671 |  | 15.7314 | 2.0886e-05 | 2.6207 |  |  | 5.8395e-06 | 2.0677e-05 | 8.7369e-07 | 1 | 0 | 0 | 43.2806 | 9.6068 | 14.9632 | 0 | 15.7276 | 9.65e-06 | 2.6201 | 0 | 0 | 2.2903e-06 | 3.4523e-06 | 5.3973e-06 | 0 | 0 | 1 | 1 | 1 | 1.6563 | 43.2806 | 9.6068 | 14.9632 | 15.7276 | 2.6201 | 0.22197 | 0.34572 | 0.36339 | 0.060537 | 1.5576 | 1.6371 | 0.27273 | 1.0511 | 0.1751 | 0.16659 | -1.5052 | -1.0621 | -1.0123 | -2.8045 | [10 empty] | 0 | 1.6563 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014148 | -1.813 | -1.8272 | [14 empty] | Bottle 3 | Lab 2 | June 3, 2020 | 2 | 104.4% | -2.1335 | -1.5104 | 8.0981 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.59% | -1.6184 | 6.051 | -29.2726 | SRM 3169 | 2SD | 11.4287 | 9.9576 | 36.02
Lab 2 | June 3, 2020 | 2 | 86 | Bottle 4 |  |  | 0.012655 | 0.0026905 | 0.004025 |  | 0.0038359 | -6.5132e-07 | 0.00061124 |  |  | 5.1148e-06 | 1.0454e-05 | 1.2603e-06 |  |  | 43.8566 | 9.7344 | 15.1614 |  | 15.9349 | 2.3816e-05 | 2.6545 |  |  | -1.9875e-06 | 1.6376e-05 | 5.7456e-06 | 1 | 0 | 0 | 43.8439 | 9.7317 | 15.1573 | 0 | 15.9311 | 2.4467e-05 | 2.6539 | 0 | 0 | -7.1023e-06 | 5.9222e-06 | 4.4853e-06 | 0 | 0 | 1 | 1 | 1 | 1.6545 | 43.8439 | 9.7317 | 15.1573 | 15.9311 | 2.6539 | 0.22196 | 0.34571 | 0.36336 | 0.06053 | 1.5575 | 1.637 | 0.2727 | 1.051 | 0.17509 | 0.16658 | -1.5052 | -1.0622 | -1.0124 | -2.8046 | [10 empty] | 0 | 1.6545 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014147 | -1.813 | -1.8272 | [14 empty] | Bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.8% | 0.046197 | -0.80021 | 3.7165 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.61% | 3.0792 | -1.2563 | 20.6223 | SRM 3169 | n | 18 | 18 | 18
Lab 2 | June 3, 2020 | 2 | 89 | Bottle 5 |  |  | 0.01105 | 0.0023378 | 0.003503 |  | 0.003332 | 8.7907e-06 | 0.00053126 |  |  | 3.7441e-07 | 9.0642e-06 | 1.5929e-06 |  |  | 43.2136 | 9.5903 | 14.9348 |  | 15.6926 | 3.0793e-05 | 2.6135 |  |  | 3.6922e-06 | 1.3506e-05 | 1.7505e-06 | 1 | 0 | 0 | 43.2026 | 9.5879 | 14.9313 | 0 | 15.6892 | 2.2003e-05 | 2.6129 | 0 | 0 | 3.3178e-06 | 4.4416e-06 | 1.5756e-07 | 0 | 0 | 1 | 1 | 1 | 1.6416 | 43.2026 | 9.5879 | 14.9312 | 15.6892 | 2.6129 | 0.22193 | 0.34561 | 0.36316 | 0.060481 | 1.5573 | 1.6364 | 0.27252 | 1.0508 | 0.175 | 0.16654 | -1.5054 | -1.0624 | -1.0129 | -2.8054 | [10 empty] | 0 | 1.6416 | 0.21793 | 0.33335 | 0.3381 | 0.054392 | 1.5296 | 1.5514 | 0.24958 | 1.0143 | 0.16317 | 0.16087 | -1.5236 | -1.0986 | -2.9115 | 0.425 | 0.43916 | -1.388 | 0.014156 | -1.813 | -1.8271 | [14 empty] | Bottle 5 | Lab 2 | June 3, 2020 | 2 | 104.2% | -4.9136 | -9.3449 | 34.5917 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.31% | 12.2993 | 2.0736 | 6.4481 | SRM 3169 | 2SE | 2.6938 | 2.347 | 8.49
Lab 2 | June 3, 2020 | 2 | 92 | Bottle 6 |  |  | 0.012773 | 0.0027142 | 0.0040339 |  | 0.0038452 | 1.7578e-05 | 0.00062182 |  |  | 1.0669e-06 | 1.5033e-05 | 4.5621e-06 |  |  | 43.9749 | 9.7606 | 15.2022 |  | 15.9778 | 1.6337e-05 | 2.6616 |  |  | 2.664e-06 | 2.5178e-06 | 4.6132e-06 | 1 | 0 | 0 | 43.9621 | 9.7579 | 15.1981 | 0 | 15.974 | -1.2418e-06 | 2.661 | 0 | 0 | 1.5971e-06 | -1.2515e-05 | 5.105e-08 | 0 | 0 | 1 | 1 | 1 | 1.6544 | 43.9621 | 9.7579 | 15.1981 | 15.974 | 2.661 | 0.22196 | 0.34571 | 0.36336 | 0.060529 | 1.5575 | 1.637 | 0.2727 | 1.051 | 0.17509 | 0.16658 | -1.5053 | -1.0622 | -1.0124 | -2.8046 | [10 empty] | 0 | 1.6544 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014148 | -1.813 | -1.8272 | [14 empty] | Bottle 6 | Lab 2 | June 3, 2020 | 2 | 106% | -5.0128 | -2.1319 | 0.054605 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.82% | -5.4721 | -1.4654 | 6.9443 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 95 | SRM 3169 |  |  | 0.013534 | 0.0028895 | 0.0043007 |  | 0.0040998 | 5.4705e-06 | 0.00064419 |  |  | 3.3895e-07 | 1.4958e-07 | 7.2092e-06 |  |  | 41.7588 | 9.2693 | 14.4378 |  | 15.1763 | 3.3557e-05 | 2.5284 |  |  | -7.6304e-07 | 2.4754e-05 | 3.0646e-06 | 1 | 0 | 0 | 41.7453 | 9.2665 | 14.4335 | 0 | 15.1722 | 2.8086e-05 | 2.5277 | 0 | 0 | -1.102e-06 | 2.4604e-05 | -4.1447e-06 | 0 | 0 | 1 | 1 | 1 | 1.6601 | 41.7453 | 9.2665 | 14.4335 | 15.1722 | 2.5277 | 0.22198 | 0.34575 | 0.36345 | 0.06055 | 1.5576 | 1.6373 | 0.27278 | 1.0512 | 0.17513 | 0.1666 | -1.5052 | -1.062 | -1.0121 | -2.8043 | [10 empty] | 0 | 1.66 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0143 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014151 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 100.5% | -0.16222 | -6.1953 | -11.1069 | [8 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.71% | -4.3392 | 0.08556 | -13.1055 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 98 | bottle 1 |  |  | 0.013371 | 0.0028444 | 0.0042502 |  | 0.0040628 | 2.7498e-06 | 0.00064753 |  |  | -2.3477e-06 | 8.8517e-06 | 1.4407e-06 |  |  | 44.7681 | 9.9366 | 15.4762 |  | 16.2656 | 3.2875e-05 | 2.7095 |  |  | 4.4989e-06 | 1.3808e-05 | 2.4148e-06 | 1 | 0 | 0 | 44.7547 | 9.9338 | 15.4719 | 0 | 16.2615 | 3.0125e-05 | 2.7089 | 0 | 0 | 6.8466e-06 | 4.9559e-06 | 9.741e-07 | 0 | 0 | 1 | 1 | 1 | 1.6538 | 44.7547 | 9.9338 | 15.4719 | 16.2615 | 2.7088 | 0.22196 | 0.34571 | 0.36335 | 0.060526 | 1.5575 | 1.637 | 0.27269 | 1.051 | 0.17508 | 0.16658 | -1.5053 | -1.0622 | -1.0124 | -2.8047 | [10 empty] | 0 | 1.6538 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014148 | -1.813 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 107% | -0.97925 | 1.764 | 2.3811 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.62% | -1.4742 | 3.4491 | 5.2311 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 101 | bottle 2 |  |  | 0.012803 | 0.0027103 | 0.0040386 |  | 0.0038398 | 1.324e-05 | 0.00060047 |  |  | -3.262e-07 | 1.4788e-05 | -3.3824e-06 |  |  | 44.5099 | 9.8793 | 15.3867 |  | 16.1712 | 2.9658e-05 | 2.6937 |  |  | 6.7972e-06 | -3.321e-06 | -1.3747e-06 | 1 | 0 | 0 | 44.4971 | 9.8765 | 15.3826 | 0 | 16.1673 | 1.6418e-05 | 2.6931 | 0 | 0 | 7.1234e-06 | -1.8109e-05 | 2.0077e-06 | 0 | 0 | 1 | 1 | 1 | 1.653 | 44.4971 | 9.8765 | 15.3827 | 16.1673 | 2.6932 | 0.22196 | 0.3457 | 0.36333 | 0.060525 | 1.5575 | 1.6369 | 0.27268 | 1.051 | 0.17508 | 0.16658 | -1.5053 | -1.0622 | -1.0124 | -2.8047 | [10 empty] | 0 | 1.653 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014145 | -1.813 | -1.8271 | [14 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 106.4% | -0.4044 | 4.8708 | 27.7052 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 98.75% | -2.5364 | -0.89683 | -14.4114 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 104 | bottle 3 |  |  | 0.013699 | 0.0029308 | 0.0043571 |  | 0.0041688 | 1.1234e-05 | 0.00064232 |  |  | 1.0621e-06 | 1.0579e-06 | 1.055e-06 |  |  | 43.9922 | 9.7642 | 15.2074 |  | 15.9825 | 1.9691e-05 | 2.6623 |  |  | 6.682e-07 | 1.8606e-05 | 2.1615e-06 | 1 | 0 | 0 | 43.9785 | 9.7613 | 15.203 | 0 | 15.9784 | 8.4571e-06 | 2.6616 | 0 | 0 | -3.939e-07 | 1.7548e-05 | 1.1064e-06 | 0 | 0 | 1 | 1 | 1 | 1.6522 | 43.9785 | 9.7613 | 15.203 | 15.9783 | 2.6616 | 0.22196 | 0.34569 | 0.36332 | 0.06052 | 1.5575 | 1.6369 | 0.27267 | 1.051 | 0.17507 | 0.16658 | -1.5053 | -1.0622 | -1.0125 | -2.8048 | [10 empty] | 0 | 1.6521 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0143 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014152 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | June 3, 2020 | 2 | 105.2% | -7.7928 | -1.855 | 11.9504 | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.21% | -6.2846 | 3.3272 | -26.0889 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 107 | bottle 4 |  |  | 0.012678 | 0.0027149 | 0.0040516 |  | 0.00388 | 1.8359e-05 | 0.0006275 |  |  | 2.5881e-06 | 1.0658e-05 | 2.9363e-06 |  |  | 43.8744 | 9.7368 | 15.163 |  | 15.9322 | 4.0677e-05 | 2.6533 |  |  | -4.2445e-06 | 1.1631e-05 | 4.1548e-06 | 1 | 0 | 0 | 43.8617 | 9.7341 | 15.1589 | 0 | 15.9283 | 2.2318e-05 | 2.6527 | 0 | 0 | -6.8326e-06 | 9.7331e-07 | 1.2185e-06 | 0 | 0 | 1 | 1 | 1 | 1.6412 | 43.8617 | 9.7341 | 15.1589 | 15.9283 | 2.6527 | 0.22193 | 0.34561 | 0.36315 | 0.060479 | 1.5573 | 1.6363 | 0.27252 | 1.0508 | 0.17499 | 0.16654 | -1.5054 | -1.0625 | -1.0129 | -2.8055 | [10 empty] | 0 | 1.6412 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0143 | 0.16317 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014154 | -1.813 | -1.8271 | [14 empty] | bottle 4 | Lab 2 | June 3, 2020 | 2 | 104.9% | -11.582 | -4.5879 | 31.6407 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 98.12% | -1.8073 | -4.1721 | 12.9664 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 110 | bottle 5 |  |  | 0.013915 | 0.0029699 | 0.004419 |  | 0.0042557 | 1.1719e-05 | 0.00067399 |  |  | -1.262e-10 | 1.4203e-05 | 4.7176e-06 |  |  | 44.4868 | 9.8733 | 15.3763 |  | 16.158 | 2.8458e-05 | 2.6912 |  |  | 2.2661e-06 | 1.6641e-05 | 2.828e-06 | 1 | 0 | 0 | 44.4728 | 9.8703 | 15.3719 | 0 | 16.1538 | 1.674e-05 | 2.6905 | 0 | 0 | 2.2662e-06 | 2.4377e-06 | -1.8896e-06 | 0 | 0 | 1 | 1 | 1 | 1.6462 | 44.4728 | 9.8703 | 15.3719 | 16.1538 | 2.6905 | 0.22194 | 0.34565 | 0.36323 | 0.060498 | 1.5574 | 1.6366 | 0.27259 | 1.0509 | 0.17503 | 0.16656 | -1.5053 | -1.0623 | -1.0127 | -2.8051 | [10 empty] | 0 | 1.6462 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0143 | 0.16317 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014151 | -1.813 | -1.8271 | [14 empty] | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.4% | -12.1061 | -1.5531 | 31.6137 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 98.51% | -7.169 | 1.3515 | -13.1292 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 113 | bottle 6 |  |  | 0.013142 | 0.0027878 | 0.0041274 |  | 0.0039239 | 2.009e-05 | 0.00062001 |  |  | 1.339e-06 | 1.4222e-05 | -3.5748e-07 |  |  | 44.4724 | 9.8706 | 15.373 |  | 16.156 | 2.7349e-05 | 2.6911 |  |  | 5.8266e-06 | 2.1999e-05 | -1.8287e-06 | 1 | 0 | 0 | 44.4593 | 9.8678 | 15.3689 | 0 | 16.1521 | 7.259e-06 | 2.6905 | 0 | 0 | 4.4877e-06 | 7.7769e-06 | -1.4712e-06 | 0 | 0 | 1 | 1 | 1 | 1.6508 | 44.4593 | 9.8678 | 15.3688 | 16.1521 | 2.6905 | 0.22195 | 0.34568 | 0.3633 | 0.060515 | 1.5575 | 1.6368 | 0.27265 | 1.051 | 0.17506 | 0.16657 | -1.5053 | -1.0622 | -1.0125 | -2.8049 | [10 empty] | 0 | 1.6508 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014144 | -1.813 | -1.8272 | [14 empty] | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.3% | -7.5427 | 5.5312 | 9.4619 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.54% | -3.283 | 2.2247 | -3.8054 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 116 | SRM 3169 |  |  | 0.016951 | 0.0036427 | 0.0054456 |  | 0.0053085 | 1.4302e-05 | 0.00084903 |  |  | 2.7637e-06 | 9.9157e-06 | -1.1754e-06 |  |  | 41.8989 | 9.3003 | 14.4857 |  | 15.2261 | 2.8735e-05 | 2.5366 |  |  | 4.1979e-06 | 1.9596e-05 | 3.8143e-06 | 1 | 0 | 0 | 41.8819 | 9.2967 | 14.4803 | 0 | 15.2208 | 1.4432e-05 | 2.5357 | 0 | 0 | 1.4342e-06 | 9.6799e-06 | 4.9897e-06 | 0 | 0 | 1 | 1 | 1 | 1.6584 | 41.8819 | 9.2967 | 14.4803 | 15.2208 | 2.5357 | 0.22197 | 0.34574 | 0.36342 | 0.060544 | 1.5576 | 1.6372 | 0.27276 | 1.0511 | 0.17512 | 0.1666 | -1.5052 | -1.0621 | -1.0122 | -2.8044 | [10 empty] | 0 | 1.6584 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014149 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.71% | 2.9689 | -1.005 | 2.7341 | [8 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.5% | 7.0327 | 7.7044 | -10.4999 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 119 | bottle 1 |  |  | 0.014347 | 0.0030437 | 0.0045372 |  | 0.0043368 | 7.7526e-06 | 0.00069159 |  |  | 7.4127e-06 | 1.99e-05 | 1.0164e-05 |  |  | 45.3222 | 10.0596 | 15.6676 |  | 16.4665 | 2.4046e-05 | 2.7429 |  |  | -3.2705e-06 | 8.5662e-06 | 6.6544e-06 | 1 | 0 | 0 | 45.3079 | 10.0566 | 15.663 | 0 | 16.4622 | 1.6293e-05 | 2.7422 | 0 | 0 | -1.0683e-05 | -1.1334e-05 | -3.5092e-06 | 0 | 0 | 1 | 1 | 1 | 1.6533 | 45.3079 | 10.0566 | 15.6631 | 16.4622 | 2.7422 | 0.22196 | 0.3457 | 0.36334 | 0.060525 | 1.5575 | 1.637 | 0.27268 | 1.051 | 0.17508 | 0.16658 | -1.5053 | -1.0622 | -1.0124 | -2.8047 | [10 empty] | 0 | 1.6533 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014145 | -1.813 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 107.7% | -0.5958 | 1.2993 | 3.0076 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.87% | -7.4011 | -1.873 | 25.0781 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 122 | bottle 2 |  |  | 0.013858 | 0.0029409 | 0.0043628 |  | 0.0041388 | 1.3195e-05 | 0.00064278 |  |  | 5.5605e-07 | 6.7852e-06 | -2.898e-06 |  |  | 44.851 | 9.9551 | 15.5051 |  | 16.2961 | 2.3637e-05 | 2.7146 |  |  | -1.5784e-06 | 2.7161e-05 | 2.331e-06 | 1 | 0 | 0 | 44.8371 | 9.9522 | 15.5007 | 0 | 16.2919 | 1.0442e-05 | 2.7139 | 0 | 0 | -2.1344e-06 | 2.0376e-05 | 5.229e-06 | 0 | 0 | 1 | 1 | 1 | 1.6545 | 44.8371 | 9.9522 | 15.5007 | 16.2919 | 2.7139 | 0.22196 | 0.34571 | 0.36336 | 0.060528 | 1.5575 | 1.637 | 0.2727 | 1.051 | 0.17508 | 0.16658 | -1.5052 | -1.0622 | -1.0124 | -2.8046 | [10 empty] | 0 | 1.6544 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014145 | -1.813 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 106.6% | -0.0079351 | 1.4716 | -9.0703 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.99% | 2.0996 | 2.1737 | 34.0865 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 125 | bottle 3 |  |  | 0.013749 | 0.0029115 | 0.0043248 |  | 0.0041028 | 6.9423e-06 | 0.00063803 |  |  | 1.2368e-06 | -1.2678e-06 | 1.951e-06 |  |  | 44.4776 | 9.8719 | 15.3751 |  | 16.1585 | 3.5597e-05 | 2.6916 |  |  | 2.0364e-06 | 8.2401e-06 | 2.7578e-06 | 1 | 0 | 0 | 44.4639 | 9.869 | 15.3707 | 0 | 16.1544 | 2.8655e-05 | 2.6909 | 0 | 0 | 7.9966e-07 | 9.5079e-06 | 8.0685e-07 | 0 | 0 | 1 | 1 | 1 | 1.6517 | 44.4639 | 9.869 | 15.3707 | 16.1544 | 2.6909 | 0.22196 | 0.34569 | 0.36331 | 0.060519 | 1.5575 | 1.6369 | 0.27266 | 1.051 | 0.17507 | 0.16657 | -1.5053 | -1.0622 | -1.0125 | -2.8048 | [10 empty] | 0 | 1.6517 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014146 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | June 3, 2020 | 2 | 105.7% | -5.6981 | 0.32831 | 13.2187 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 96.94% | -3.8649 | -5.3327 | -27.8798 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 128 | bottle 4 |  |  | 0.014079 | 0.002991 | 0.0044406 |  | 0.0042319 | 9.9228e-06 | 0.00065679 |  |  | 1.8333e-06 | 1.0388e-05 | -6.706e-07 |  |  | 44.5259 | 9.8826 | 15.3917 |  | 16.1758 | 3.2027e-05 | 2.6944 |  |  | 2.3353e-06 | 1.2897e-05 | 2.4449e-06 | 1 | 0 | 0 | 44.5119 | 9.8796 | 15.3872 | 0 | 16.1716 | 2.2104e-05 | 2.6937 | 0 | 0 | 5.0193e-07 | 2.5095e-06 | 3.1155e-06 | 0 | 0 | 1 | 1 | 1 | 1.6514 | 44.5119 | 9.8796 | 15.3872 | 16.1716 | 2.6937 | 0.22195 | 0.34569 | 0.36331 | 0.060516 | 1.5575 | 1.6369 | 0.27265 | 1.051 | 0.17506 | 0.16657 | -1.5053 | -1.0622 | -1.0125 | -2.8048 | [10 empty] | 0 | 1.6514 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014144 | -1.813 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.8% | -8.7607 | 2.0741 | -7.65 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.42% | -2.093 | -2.9486 | -3.1949 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 131 | bottle 5 |  |  | 0.01632 | 0.0034837 | 0.0052145 |  | 0.0050106 | 3.6986e-06 | 0.0007901 |  |  | 4.2815e-06 | 6.4511e-06 | 2.6209e-06 |  |  | 44.8437 | 9.9528 | 15.5006 |  | 16.2894 | 2.2586e-05 | 2.7132 |  |  | 4.762e-07 | 1.3828e-05 | -5.5568e-06 | 1 | 0 | 0 | 44.8274 | 9.9494 | 15.4954 | 0 | 16.2844 | 1.8888e-05 | 2.7124 | 0 | 0 | -3.8053e-06 | 7.3772e-06 | -8.1777e-06 | 0 | 0 | 1 | 1 | 1 | 1.6489 | 44.8274 | 9.9494 | 15.4954 | 16.2844 | 2.7124 | 0.22195 | 0.34567 | 0.36327 | 0.060508 | 1.5574 | 1.6367 | 0.27262 | 1.0509 | 0.17505 | 0.16656 | -1.5053 | -1.0623 | -1.0126 | -2.805 | [10 empty] | 0 | 1.6488 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014144 | -1.813 | -1.8272 | [14 empty] | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.6% | -5.7717 | 2.0264 | 14.2531 | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.09% | 10.4157 | -9.4061 | -1.5772 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 134 | bottle 6 |  |  | 0.014682 | 0.0031322 | 0.0046359 |  | 0.004422 | 1.7239e-05 | 0.00068858 |  |  | 8.6045e-06 | 1.0954e-05 | 3.5462e-07 |  |  | 44.8184 | 9.9472 | 15.4918 |  | 16.28 | 3.1179e-05 | 2.7116 |  |  | 5.35e-07 | 2.9061e-06 | 2.9897e-06 | 1 | 0 | 0 | 44.8038 | 9.9441 | 15.4872 | 0 | 16.2756 | 1.394e-05 | 2.7109 | 0 | 0 | -8.0695e-06 | -8.048e-06 | 2.6351e-06 | 0 | 0 | 1 | 1 | 1 | 1.6485 | 44.8038 | 9.9441 | 15.4872 | 16.2756 | 2.7109 | 0.22195 | 0.34567 | 0.36326 | 0.060507 | 1.5574 | 1.6367 | 0.27262 | 1.0509 | 0.17504 | 0.16656 | -1.5053 | -1.0623 | -1.0126 | -2.805 | [10 empty] | 0 | 1.6485 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014142 | -1.813 | -1.8271 | [14 empty] | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.5% | -3.8999 | 4.5634 | 19.0303
Lab 2 | June 3, 2020 | 2 | 137 | SRM 3169 |  |  | 0.014433 | 0.00306 | 0.0045498 |  | 0.0043343 | -8.7878e-06 | 0.00067503 |  |  | -1.7588e-06 | 8.2998e-07 | 3.2535e-06 |  |  | 42.2753 | 9.3834 | 14.6147 |  | 15.3603 | 3.4166e-05 | 2.5587 |  |  | 4.7035e-06 | 7.4151e-06 | 8.3536e-07 | 1 | 0 | 0 | 42.2608 | 9.3804 | 14.6101 | 0 | 15.3559 | 4.2953e-05 | 2.558 | 0 | 0 | 6.4622e-06 | 6.5851e-06 | -2.4182e-06 | 0 | 0 | 1 | 1 | 1 | 1.6546 | 42.2608 | 9.3804 | 14.6101 | 15.3559 | 2.558 | 0.22196 | 0.34571 | 0.36336 | 0.060529 | 1.5575 | 1.637 | 0.2727 | 1.051 | 0.17509 | 0.16658 | -1.5052 | -1.0621 | -1.0124 | -2.8046 | [10 empty] | 0 | 1.6546 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014144 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 100.1% | 2.2062 | 5.2455 | -7.906 | [8 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 99.41% | -11.2534 | 0.062756 | 18.2876 | SRM 3169 | Avg | -0.019457 | 0.8487 | 10.6151
Lab 2 | June 3, 2020 | 2 | 140 | bottle 1 |  |  | 0.014688 | 0.003117 | 0.0046343 |  | 0.0044099 | 5.0299e-06 | 0.00069649 |  |  | 3.8368e-06 | 1.1325e-05 | 1.4719e-06 |  |  | 46.0632 | 10.2231 | 15.9211 |  | 16.7302 | 2.4141e-05 | 2.7864 |  |  | 2.7668e-06 | 4.3791e-06 | -3.5677e-06 | 1 | 0 | 0 | 46.0485 | 10.22 | 15.9165 | 0 | 16.7258 | 1.9111e-05 | 2.7857 | 0 | 0 | -1.07e-06 | -6.9457e-06 | -5.0396e-06 | 0 | 0 | 1 | 1 | 1 | 1.6459 | 46.0485 | 10.22 | 15.9165 | 16.7258 | 2.7858 | 0.22194 | 0.34565 | 0.36322 | 0.060496 | 1.5574 | 1.6366 | 0.27258 | 1.0509 | 0.17502 | 0.16655 | -1.5053 | -1.0623 | -1.0127 | -2.8052 | [10 empty] | 0 | 1.6459 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014144 | -1.813 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 108.6% | -5.4689 | 2.4144 | 6.3336 | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 100.9% | -3.4693 | 1.2326 | 14.4317 | SRM 3169 | 2SD | 20.8945 | 15.9083 | 21.0421
Lab 2 | June 3, 2020 | 2 | 143 | bottle 2 |  |  | 0.014246 | 0.0030403 | 0.0044953 |  | 0.0042676 | 1.0425e-05 | 0.00067462 |  |  | -4.2672e-06 | 1.1275e-05 | 4.5543e-06 |  |  | 45.832 | 10.1718 | 15.841 |  | 16.6459 | 3.1283e-05 | 2.7724 |  |  | -3.3843e-06 | 1.6247e-05 | 5.6568e-06 | 1 | 0 | 0 | 45.8178 | 10.1687 | 15.8365 | 0 | 16.6416 | 2.0859e-05 | 2.7717 | 0 | 0 | 8.8292e-07 | 4.9718e-06 | 1.1025e-06 | 0 | 0 | 1 | 1 | 1 | 1.6453 | 45.8178 | 10.1687 | 15.8365 | 16.6416 | 2.7717 | 0.22194 | 0.34564 | 0.36321 | 0.060494 | 1.5574 | 1.6365 | 0.27257 | 1.0508 | 0.17502 | 0.16655 | -1.5054 | -1.0624 | -1.0128 | -2.8052 | [10 empty] | 0 | 1.6453 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014145 | -1.813 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 108.1% | -5.1477 | 2.176 | 11.1908 | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 101.9% | -3.3751 | -2.6575 | 19.8375 | SRM 3169 | n | 16 | 16 | 16
Lab 2 | June 3, 2020 | 2 | 146 | bottle 3 |  |  | 0.013947 | 0.0029605 | 0.0043854 |  | 0.0041633 | 1.6132e-05 | 0.00064602 |  |  | 3.6231e-06 | 1.4153e-05 | 2.0873e-06 |  |  | 45.2217 | 10.0363 | 15.6299 |  | 16.4238 | 3.8212e-05 | 2.7354 |  |  | -2.1195e-06 | 1.964e-05 | -4.043e-06 | 1 | 0 | 0 | 45.2078 | 10.0333 | 15.6255 | 0 | 16.4197 | 2.208e-05 | 2.7347 | 0 | 0 | -5.7426e-06 | 5.4876e-06 | -6.1303e-06 | 0 | 0 | 1 | 1 | 1 | 1.6447 | 45.2078 | 10.0333 | 15.6255 | 16.4197 | 2.7347 | 0.22194 | 0.34564 | 0.3632 | 0.060492 | 1.5574 | 1.6365 | 0.27257 | 1.0508 | 0.17502 | 0.16655 | -1.5054 | -1.0624 | -1.0128 | -2.8052 | [10 empty] | 0 | 1.6447 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014147 | -1.813 | -1.8271 | [14 empty] | bottle 3 | Lab 2 | June 3, 2020 | 2 | 106.6% | -4.0346 | 0.14098 | 20.8489 | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 99.34% | -7.9832 | 0.6368 | 10.7505 | SRM 3169 | 2SE | 5.2236 | 3.9771 | 5.2605
Lab 2 | June 3, 2020 | 2 | 149 | bottle 4 |  |  | 0.013906 | 0.0029519 | 0.0043661 |  | 0.0041563 | 9.832e-06 | 0.00067569 |  |  | 4.1697e-07 | 4.3131e-06 | 8.2309e-06 |  |  | 44.9231 | 9.9704 | 15.5275 |  | 16.3171 | 4.1479e-05 | 2.7177 |  |  | 4.0645e-06 | 1.0513e-05 | -7.4976e-07 | 1 | 0 | 0 | 44.9092 | 9.9674 | 15.5231 | 0 | 16.3129 | 3.1647e-05 | 2.717 | 0 | 0 | 3.6476e-06 | 6.2e-06 | -8.9806e-06 | 0 | 0 | 1 | 1 | 1 | 1.6472 | 44.9092 | 9.9674 | 15.5231 | 16.3129 | 2.717 | 0.22195 | 0.34566 | 0.36324 | 0.060501 | 1.5574 | 1.6366 | 0.27259 | 1.0509 | 0.17503 | 0.16656 | -1.5053 | -1.0623 | -1.0127 | -2.8051 | [10 empty] | 0 | 1.6472 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014144 | -1.813 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.9% | 4.9061 | 2.795 | -1.2178 | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 100.8% | 11.5945 | 3.8951 | 11.4528 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 152 | bottle 5 |  |  | 0.015138 | 0.0032339 | 0.0048118 |  | 0.0045888 | 6.8669e-06 | 0.00072083 |  |  | 6.8175e-07 | 1.8305e-05 | 1.6856e-07 |  |  | 45.2037 | 10.0321 | 15.6231 |  | 16.4163 | 3.1911e-05 | 2.734 |  |  | 4.9802e-06 | 1.5438e-05 | 2.136e-06 | 1 | 0 | 0 | 45.1885 | 10.0289 | 15.6183 | 0 | 16.4117 | 2.5044e-05 | 2.7333 | 0 | 0 | 4.2985e-06 | -2.8666e-06 | 1.9675e-06 | 0 | 0 | 1 | 1 | 1 | 1.6434 | 45.1885 | 10.0289 | 15.6183 | 16.4117 | 2.7333 | 0.22193 | 0.34563 | 0.36318 | 0.060487 | 1.5573 | 1.6364 | 0.27255 | 1.0508 | 0.17501 | 0.16655 | -1.5054 | -1.0624 | -1.0128 | -2.8053 | [10 empty] | 0 | 1.6434 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014148 | -1.813 | -1.8271 | [14 empty] | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.6% | -5.7039 | -1.4478 | 18.6152 | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 101.4% | -9.3041 | -5.1812 | 9.4346 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 155 | bottle 6 |  |  | 0.014724 | 0.0031249 | 0.0046343 |  | 0.0044229 | 5.913e-06 | 0.00070046 |  |  | 3.2872e-06 | 7.9671e-06 | 6.0756e-06 |  |  | 45.1026 | 10.01 | 15.5893 |  | 16.3817 | 1.9388e-05 | 2.7284 |  |  | 4.1035e-06 | 1.2231e-05 | 7.7543e-08 | 1 | 0 | 0 | 45.0879 | 10.0069 | 15.5846 | 0 | 16.3773 | 1.3475e-05 | 2.7277 | 0 | 0 | 8.1637e-07 | 4.264e-06 | -5.998e-06 | 0 | 0 | 1 | 1 | 1 | 1.6464 | 45.0879 | 10.0069 | 15.5846 | 16.3773 | 2.7277 | 0.22194 | 0.34565 | 0.36323 | 0.060497 | 1.5574 | 1.6366 | 0.27258 | 1.0509 | 0.17502 | 0.16655 | -1.5053 | -1.0623 | -1.0127 | -2.8052 | [10 empty] | 0 | 1.6464 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014144 | -1.813 | -1.8272 | [14 empty] | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.4% | -2.2142 | 2.4611 | -5.1259 | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 99.12% | 3.2016 | -4.6299 | 4.6052 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 158 | SRM 3169 |  |  | 0.017794 | 0.0038115 | 0.0056975 |  | 0.0055156 | 1.7005e-05 | 0.00086238 |  |  | 1.4149e-06 | 5.7902e-06 | 1.0074e-05 |  |  | 42.546 | 9.4432 | 14.7073 |  | 15.4568 | 2.8098e-05 | 2.5747 |  |  | -1.4119e-07 | 1.8723e-05 | 1.1093e-06 | 1 | 0 | 0 | 42.5282 | 9.4394 | 14.7016 | 0 | 15.4513 | 1.1094e-05 | 2.5738 | 0 | 0 | -1.5561e-06 | 1.2933e-05 | -8.9642e-06 | 0 | 0 | 1 | 1 | 1 | 1.652 | 42.5282 | 9.4394 | 14.7016 | 15.4513 | 2.5738 | 0.22196 | 0.34569 | 0.36332 | 0.06052 | 1.5575 | 1.6369 | 0.27267 | 1.051 | 0.17507 | 0.16658 | -1.5053 | -1.0622 | -1.0125 | -2.8048 | [10 empty] | 0 | 1.652 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0143 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.01415 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.99% | -3.6295 | -4.2218 | 7.1394 | [8 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 100.3% | 10.9911 | 1.8885 | 22.7935 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 161 | bottle 1 |  |  | 0.015061 | 0.0031974 | 0.0047348 |  | 0.0045024 | 2.115e-06 | 0.00070964 |  |  | 6.7929e-06 | 1.3398e-05 | -6.2842e-06 |  |  | 46.4309 | 10.3051 | 16.0491 |  | 16.8655 | 4.3875e-05 | 2.8091 |  |  | 9.7809e-08 | 8.3614e-06 | 3.8127e-06 | 1 | 0 | 0 | 46.4158 | 10.3019 | 16.0444 | 0 | 16.861 | 4.176e-05 | 2.8084 | 0 | 0 | -6.6951e-06 | -5.0361e-06 | 1.0097e-05 | 0 | 0 | 1 | 1 | 1 | 1.6483 | 46.4158 | 10.3019 | 16.0444 | 16.861 | 2.8084 | 0.22195 | 0.34567 | 0.36326 | 0.060505 | 1.5574 | 1.6367 | 0.27261 | 1.0509 | 0.17504 | 0.16656 | -1.5053 | -1.0623 | -1.0126 | -2.805 | [10 empty] | 0 | 1.6483 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014141 | -1.813 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 108.8% | 6.1721 | 8.0101 | -3.7042 | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 101.9% | 10.5067 | 7.44 | 10.4908 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 164 | bottle 2 |  |  | 0.015069 | 0.00321 | 0.0047284 |  | 0.0044945 | 1.1814e-05 | 0.0006929 |  |  | 2.7061e-06 | 1.8201e-05 | 1.6737e-08 |  |  | 46.0909 | 10.2298 | 15.932 |  | 16.743 | 2.5293e-05 | 2.7887 |  |  | 3.6287e-06 | 1.5985e-05 | 3.1939e-06 | 1 | 0 | 0 | 46.0759 | 10.2265 | 15.9273 | 0 | 16.7385 | 1.348e-05 | 2.788 | 0 | 0 | 9.226e-07 | -2.2155e-06 | 3.1772e-06 | 0 | 0 | 1 | 1 | 1 | 1.6496 | 46.0759 | 10.2265 | 15.9273 | 16.7385 | 2.788 | 0.22195 | 0.34567 | 0.36328 | 0.06051 | 1.5574 | 1.6368 | 0.27263 | 1.0509 | 0.17505 | 0.16657 | -1.5053 | -1.0623 | -1.0126 | -2.8049 | [10 empty] | 0 | 1.6496 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014143 | -1.813 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 108% | 0.26107 | 5.5659 | -11.7135 | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 7.9236 | 16.2301 | 21.3491 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 167 | bottle 3 |  |  | 0.015131 | 0.0032199 | 0.004749 |  | 0.0045102 | 4.9963e-06 | 0.0006839 |  |  | 1.5801e-06 | 2.9078e-06 | 3.8947e-06 |  |  | 45.5044 | 10.0995 | 15.7289 |  | 16.5293 | 3.9042e-05 | 2.7531 |  |  | 5.3567e-06 | 1.7175e-05 | -4.5391e-06 | 1 | 0 | 0 | 45.4893 | 10.0963 | 15.7242 | 0 | 16.5248 | 3.4046e-05 | 2.7524 | 0 | 0 | 3.7765e-06 | 1.4267e-05 | -8.4338e-06 | 0 | 0 | 1 | 1 | 1 | 1.6487 | 45.4893 | 10.0963 | 15.7242 | 16.5248 | 2.7524 | 0.22195 | 0.34567 | 0.36327 | 0.060506 | 1.5574 | 1.6367 | 0.27261 | 1.0509 | 0.17504 | 0.16656 | -1.5053 | -1.0623 | -1.0126 | -2.805 | [10 empty] | 0 | 1.6487 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014146 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | June 3, 2020 | 2 | 106.6% | 4.7211 | 2.9817 | -16.5573 | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.6% | -7.2536 | -20.1022 | 20.5925 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 170 | bottle 4 |  |  | 0.016427 | 0.0035048 | 0.0052096 |  | 0.0049806 | 1.3778e-06 | 0.00079762 |  |  | 1.0106e-06 | 1.0814e-05 | 5.4304e-07 |  |  | 45.2073 | 10.0333 | 15.6258 |  | 16.4203 | 3.2026e-05 | 2.7349 |  |  | -1.1619e-06 | 2.0407e-05 | 2.7187e-06 | 1 | 0 | 0 | 45.1909 | 10.0298 | 15.6205 | 0 | 16.4153 | 3.0648e-05 | 2.7341 | 0 | 0 | -2.1725e-06 | 9.5924e-06 | 2.1757e-06 | 0 | 0 | 1 | 1 | 1 | 1.6473 | 45.1909 | 10.0298 | 15.6205 | 16.4153 | 2.7341 | 0.22194 | 0.34566 | 0.36324 | 0.0605 | 1.5574 | 1.6366 | 0.27259 | 1.0509 | 0.17503 | 0.16656 | -1.5053 | -1.0623 | -1.0127 | -2.8051 | [10 empty] | 0 | 1.6472 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014145 | -1.813 | -1.8272 | [14 empty] | bottle 4 | Lab 2 | June 3, 2020 | 2 | 105.9% | -2.664 | 3.8111 | -19.3774 | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.9% | -19.3296 | 5.027 | -1.6762 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 173 | bottle 5 |  |  | 0.016564 | 0.0035347 | 0.0052517 |  | 0.0050334 | 1.7403e-05 | 0.00078663 |  |  | 3.1515e-06 | 1.5647e-05 | -1.2696e-06 |  |  | 45.4907 | 10.0962 | 15.7234 |  | 16.5226 | 3.9642e-05 | 2.7519 |  |  | -4.7385e-08 | 1.3909e-05 | -1.003e-06 | 1 | 0 | 0 | 45.4741 | 10.0927 | 15.7182 | 0 | 16.5176 | 2.224e-05 | 2.7511 | 0 | 0 | -3.1989e-06 | -1.7384e-06 | 2.6664e-07 | 0 | 0 | 1 | 1 | 1 | 1.6464 | 45.4741 | 10.0927 | 15.7182 | 16.5176 | 2.7511 | 0.22194 | 0.34565 | 0.36323 | 0.060498 | 1.5574 | 1.6366 | 0.27258 | 1.0509 | 0.17502 | 0.16655 | -1.5053 | -1.0623 | -1.0127 | -2.8052 | [10 empty] | 0 | 1.6464 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014142 | -1.813 | -1.8272 | [14 empty] | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.6% | 2.3497 | 6.3966 | -10.9165 | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | 18.2122 | -3.5481 | 19.6356 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 176 | bottle 6 |  |  | 0.015356 | 0.0032603 | 0.004834 |  | 0.0045813 | 3.8076e-06 | 0.00070683 |  |  | 2.5729e-06 | 6.9279e-06 | 5.6771e-07 |  |  | 45.5936 | 10.1187 | 15.7578 |  | 16.5577 | 2.1377e-05 | 2.7575 |  |  | -1.6663e-07 | 5.364e-06 | 2.9002e-07 | 1 | 0 | 0 | 45.5782 | 10.1154 | 15.753 | 0 | 16.5531 | 1.7569e-05 | 2.7568 | 0 | 0 | -2.7395e-06 | -1.5639e-06 | -2.777e-07 | 0 | 0 | 1 | 1 | 1 | 1.6432 | 45.5782 | 10.1154 | 15.753 | 16.5531 | 2.7568 | 0.22194 | 0.34563 | 0.36318 | 0.060486 | 1.5573 | 1.6364 | 0.27254 | 1.0508 | 0.175 | 0.16654 | -1.5054 | -1.0624 | -1.0129 | -2.8054 | [10 empty] | 0 | 1.6432 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014146 | -1.813 | -1.8272 | [14 empty] | bottle 6 | Lab 2 | June 3, 2020 | 2 | 106.8% | 3.7989 | 2.0662 | -1.2906 | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.9% | -0.12937 | 1.6245 | 8.1355 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 179 | SRM 3169 |  |  | 0.015088 | 0.0032071 | 0.0047377 |  | 0.0044928 | 4.036e-06 | 0.00068961 |  |  | 3.1519e-06 | 5.1291e-06 | 2.9281e-06 |  |  | 42.8212 | 9.504 | 14.8017 |  | 15.5551 | 4.206e-05 | 2.5909 |  |  | 2.4484e-06 | 8.9771e-06 | -1.8841e-07 | 1 | 0 | 0 | 42.8061 | 9.5008 | 14.7969 | 0 | 15.5506 | 3.8024e-05 | 2.5902 | 0 | 0 | -7.0355e-07 | 3.848e-06 | -3.1165e-06 | 0 | 0 | 1 | 1 | 1 | 1.6496 | 42.8061 | 9.5008 | 14.7969 | 15.5506 | 2.5902 | 0.22195 | 0.34567 | 0.36328 | 0.060511 | 1.5574 | 1.6368 | 0.27263 | 1.0509 | 0.17505 | 0.16657 | -1.5053 | -1.0623 | -1.0126 | -2.8049 | [10 empty] | 0 | 1.6496 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014147 | -1.813 | -1.8272 | [6 empty] | SRM 3169 | Lab 2 | June 3, 2020 | 2 | 99.79% | 2.7351 | 1.8993 | -4.5777 | [8 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | 8.0448 | 10.9871 | -10.6413 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 182 | bottle 1 |  |  | 0.023959 | 0.0051569 | 0.0078017 |  | 0.0076998 | -1.2031e-06 | 0.0012391 |  |  | 3.4735e-07 | 1.3992e-05 | 1.908e-06 |  |  | 47.1667 | 10.4681 | 16.3023 |  | 17.1307 | 2.2537e-05 | 2.8531 |  |  | 2.6172e-06 | 8.8051e-06 | -4.027e-06 | 1 | 0 | 0 | 47.1428 | 10.4629 | 16.2945 | 0 | 17.123 | 2.374e-05 | 2.8519 | 0 | 0 | 2.2698e-06 | -5.1871e-06 | -5.935e-06 | 0 | 0 | 1 | 1 | 1 | 1.6454 | 47.1428 | 10.4629 | 16.2945 | 17.123 | 2.8519 | 0.22194 | 0.34564 | 0.36322 | 0.060494 | 1.5574 | 1.6365 | 0.27257 | 1.0508 | 0.17502 | 0.16655 | -1.5053 | -1.0624 | -1.0128 | -2.8052 | [10 empty] | 0 | 1.6454 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014148 | -1.813 | -1.8272 | [14 empty] | bottle 1 | Lab 2 | June 3, 2020 | 2 | 109.5% | 4.0714 | 0.32291 | -12.3461 | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.9% | -8.688 | 0.67367 | -9.6369 | SRM 3169
Lab 2 | June 3, 2020 | 2 | 185 | bottle 2 |  |  | 0.017481 | 0.0037227 | 0.005554 |  | 0.0053409 | -7.2724e-07 | 0.00084628 |  |  | -2.5413e-06 | 3.281e-06 | 8.2379e-07 |  |  | 46.9561 | 10.4212 | 16.2295 |  | 17.054 | 4.4988e-05 | 2.8403 |  |  | 2.6505e-07 | 1.6783e-05 | -3.8022e-06 | 1 | 0 | 0 | 46.9386 | 10.4175 | 16.224 | 0 | 17.0486 | 4.5715e-05 | 2.8395 | 0 | 0 | 2.8064e-06 | 1.3502e-05 | -4.626e-06 | 0 | 0 | 1 | 1 | 1 | 1.6452 | 46.9386 | 10.4175 | 16.224 | 17.0486 | 2.8394 | 0.22194 | 0.34564 | 0.36321 | 0.060492 | 1.5574 | 1.6365 | 0.27256 | 1.0508 | 0.17501 | 0.16655 | -1.5054 | -1.0624 | -1.0128 | -2.8052 | [10 empty] | 0 | 1.6452 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43915 | -1.388 | 0.014141 | -1.813 | -1.8272 | [14 empty] | bottle 2 | Lab 2 | June 3, 2020 | 2 | 109.1% | -0.41889 | 6.9819 | -19.9101
Lab 2 | June 3, 2020 | 2 | 188 | bottle 3 |  |  | 0.01579 | 0.0033702 | 0.0049783 |  | 0.0047184 | 4.4926e-06 | 0.00073213 |  |  | 1.2448e-06 | 2.2512e-06 | 6.6473e-08 |  |  | 46.3225 | 10.2806 | 16.0102 |  | 16.8234 | 2.294e-05 | 2.8018 |  |  | 2.0213e-06 | 1.6113e-05 | 6.2984e-06 | 1 | 0 | 0 | 46.3067 | 10.2772 | 16.0052 | 0 | 16.8187 | 1.8448e-05 | 2.8011 | 0 | 0 | 7.7646e-07 | 1.3862e-05 | 6.2319e-06 | 0 | 0 | 1 | 1 | 1 | 1.6445 | 46.3067 | 10.2772 | 16.0052 | 16.8186 | 2.8011 | 0.22194 | 0.34563 | 0.3632 | 0.06049 | 1.5573 | 1.6365 | 0.27255 | 1.0508 | 0.17501 | 0.16655 | -1.5054 | -1.0624 | -1.0128 | -2.8053 | [10 empty] | 0 | 1.6445 | 0.21793 | 0.33335 | 0.3381 | 0.054389 | 1.5296 | 1.5514 | 0.24957 | 1.0143 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.425 | 0.43915 | -1.388 | 0.014151 | -1.813 | -1.8272 | [14 empty] | bottle 3 | Lab 2 | June 3, 2020 | 2 | 107.6% | 0.89306 | -2.8911 | -24.8284 | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.45% | 10.2236 | 0.31302 | 22.4311 | SRM 3169 | Avg | 2.6242 | -0.77758 | 0.98632
Lab 2 | June 3, 2020 | 2 | 191 | bottle 4 |  |  | 0.016694 | 0.0035495 | 0.0052894 |  | 0.0050602 | 2.9418e-05 | 0.00079459 |  |  | 2.0321e-06 | 1.5206e-05 | -2.5244e-07 |  |  | 45.0926 | 10.006 | 15.5803 |  | 16.367 | 3.0508e-05 | 2.7252 |  |  | 4.809e-06 | 8.3437e-06 | 1.8267e-06 | 1 | 0 | 0 | 45.0759 | 10.0024 | 15.575 | 0 | 16.3619 | 1.0904e-06 | 2.7244 | 0 | 0 | 2.7769e-06 | -6.8628e-06 | 2.0791e-06 | 0 | 0 | 1 | 1 | 1 | 1.6309 | 45.0759 | 10.0024 | 15.575 | 16.3619 | 2.7244 | 0.2219 | 0.34553 | 0.36299 | 0.06044 | 1.5571 | 1.6358 | 0.27237 | 1.0505 | 0.17492 | 0.16651 | -1.5055 | -1.0627 | -1.0134 | -2.8061 | [10 empty] | 0 | 1.6309 | 0.21793 | 0.33335 | 0.3381 | 0.054392 | 1.5296 | 1.5514 | 0.24958 | 1.0143 | 0.16317 | 0.16088 | -1.5236 | -1.0986 | -2.9115 | 0.42501 | 0.43916 | -1.388 | 0.014157 | -1.813 | -1.8271 | [14 empty] | bottle 4 | Lab 2 | June 3, 2020 | 2 | 104.7% | -9.4461 | -9.0355 | 27.1817 | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.2% | 0.50305 | 5.3597 | -11.2572 | SRM 3169 | 2SD | 9.5464 | 7.8571 | 43.4203
Lab 2 | June 3, 2020 | 2 | 194 | bottle 5 |  |  | 0.016144 | 0.0034502 | 0.0051157 |  | 0.0049028 | 1.6515e-05 | 0.00077549 |  |  | 4.754e-06 | 4.0911e-06 | -9.4996e-07 |  |  | 46.0017 | 10.2085 | 15.8967 |  | 16.7015 | 2.101e-05 | 2.7812 |  |  | 1.3249e-06 | 9.2004e-06 | 2.2249e-06 | 1 | 0 | 0 | 45.9855 | 10.205 | 15.8916 | 0 | 16.6966 | 4.4946e-06 | 2.7804 | 0 | 0 | -3.4291e-06 | 5.1093e-06 | 3.1749e-06 | 0 | 0 | 1 | 1 | 1 | 1.6372 | 45.9855 | 10.205 | 15.8916 | 16.6966 | 2.7804 | 0.22192 | 0.34558 | 0.36308 | 0.060463 | 1.5572 | 1.6361 | 0.27246 | 1.0507 | 0.17496 | 0.16653 | -1.5054 | -1.0625 | -1.0131 | -2.8057 | [10 empty] | 0 | 1.6372 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16317 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014149 | -1.813 | -1.8271 | [14 empty] | bottle 5 | Lab 2 | June 3, 2020 | 2 | 106.9% | -2.3783 | -0.67613 | 9.6052 | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.23% | -0.53785 | -3.4229 | -13.3367 | SRM 3169 | n | 5 | 5 | 5
Lab 2 | June 3, 2020 | 2 | 197 | bottle 6 |  |  | 0.018252 | 0.0038984 | 0.0058332 |  | 0.0056331 | 2.5528e-06 | 0.00088549 |  |  | 3.4784e-06 | 9.026e-06 | -9.9147e-07 |  |  | 46.0437 | 10.2179 | 15.9115 |  | 16.7174 | 2.1594e-05 | 2.7838 |  |  | 2.3836e-06 | 2.5954e-05 | 2.487e-06 | 1 | 0 | 0 | 46.0254 | 10.214 | 15.9057 | 0 | 16.7117 | 1.9041e-05 | 2.783 | 0 | 0 | -1.0949e-06 | 1.6928e-05 | 3.4784e-06 | 0 | 0 | 1 | 1 | 1 | 1.638 | 46.0254 | 10.214 | 15.9056 | 16.7117 | 2.7829 | 0.22192 | 0.34558 | 0.3631 | 0.060465 | 1.5572 | 1.6362 | 0.27246 | 1.0507 | 0.17496 | 0.16653 | -1.5054 | -1.0625 | -1.0131 | -2.8057 | [10 empty] | 0 | 1.638 | 0.21793 | 0.33335 | 0.3381 | 0.05439 | 1.5296 | 1.5514 | 0.24957 | 1.0143 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014151 | -1.813 | -1.8272 | [14 empty] | bottle 6 | Lab 2 | June 3, 2020 | 2 | 107% | -4.2375 | -2.752 | -11.7731 | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 98.98% | 4.3124 | -1.4287 | -19.4449 | SRM 3169 | 2SE | 4.2693 | 3.5138 | 19.4181
Lab 2 | June 3, 2020 | 2 | 200 | SRM 3169 |  |  | 0.015637 | 0.0033083 | 0.0049299 |  | 0.0046917 | 1.5026e-05 | 0.00074222 |  |  | 4.1336e-06 | 1.3969e-05 | 6.2851e-06 |  |  | 43.2773 | 9.6044 | 14.9567 |  | 15.7153 | 4.2099e-05 | 2.6172 |  |  | 2.586e-06 | 1.6729e-05 | 5.5278e-06 | 1 | 0 | 0 | 43.2617 | 9.601 | 14.9518 | 0 | 15.7107 | 2.7073e-05 | 2.6165 | 0 | 0 | -1.5476e-06 | 2.7599e-06 | -7.5734e-07 | 0 | 0 | 1 | 1 | 1 | 1.6416 | 43.2617 | 9.601 | 14.9517 | 15.7106 | 2.6165 | 0.22193 | 0.34561 | 0.36315 | 0.06048 | 1.5573 | 1.6363 | 0.27252 | 1.0508 | 0.17499 | 0.16654 | -1.5054 | -1.0624 | -1.0129 | -2.8054 | [10 empty] | 0 | 1.6416 | 0.21793 | 0.33335 | 0.3381 | 0.054391 | 1.5296 | 1.5514 | 0.24958 | 1.0142 | 0.16316 | 0.16087 | -1.5236 | -1.0986 | -2.9116 | 0.42501 | 0.43916 | -1.388 | 0.014148 | -1.813 | -1.8271 | [22 empty] | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.91% | -1.3804 | -4.709 | 26.5393 | SRM 3169
[56 empty] | 0 | 0 | 0
Lab 3 | Feb 4, 2020 | 1 | 2 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: -1.7442e-06; SAMPLE 49Ti40Ar: -2.6382e-06; SAMPLE 90Zr: 29.9579; SAMPLE 91Zr: 6.6635; SAMPLE 92Zr: 10.3972; SAMPLE 94Zr: 10.9688; SAMPLE 96Zr: 1.8336 | 1 | -0.00015145 | -2.3128e-05 | 29.9573 | 6.6633 | 10.3968 | 0 | 10.9684 | 0 | 1.8333 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8292 | 29.9573 | 6.6633 | 10.3968 | 10.9684 | 1.8333 | 0.22243 | 0.34705 | 0.36613 | 0.061197 | 1.5603 | 1.6461 | 0.27513 | 1.055 | 0.17633 | 0.16714 | -1.5032 | -1.0583 | -1.0048 | -2.7937 | [9 empty] | L3 FebS1 | 0 | 1.8292 | 0.21797 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.014112 | -1.8134 | -1.8275 | [22 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.8% | -9.5918 | -2.4354 | -9.3319 | SRM 3169 | Avg | -1.0744 | -1.2983 | -0.69774
Lab 3 | Feb 4, 2020 | 1 | 3 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.5747e-05; SAMPLE 49Ti40Ar: 2.9456e-05; SAMPLE 90Zr: 29.8408; SAMPLE 91Zr: 6.6375; SAMPLE 92Zr: 10.3568; SAMPLE 94Zr: 10.9265; SAMPLE 96Zr: 1.8265 | 1 | -9.3962e-05 | 8.9664e-06 | 29.8401 | 6.6373 | 10.3564 | 0 | 10.9261 | 0 | 1.8263 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8305 | 29.8401 | 6.6373 | 10.3564 | 10.9261 | 1.8263 | 0.22243 | 0.34706 | 0.36615 | 0.061202 | 1.5603 | 1.6462 | 0.27515 | 1.055 | 0.17634 | 0.16715 | -1.5031 | -1.0582 | -1.0047 | -2.7936 | [10 empty] | 0 | 1.8305 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014114 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.3% | -3.5107 | -3.6879 | -7.8244 | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.53% | 0.53586 | -1.0297 | -3.147 | SRM 3169 | 2SD | 9.7058 | 7.9162 | 24.2156
Lab 3 | Feb 4, 2020 | 1 | 4 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.9778e-05; SAMPLE 49Ti40Ar: 1.1789e-05; SAMPLE 90Zr: 29.57; SAMPLE 91Zr: 6.5774; SAMPLE 92Zr: 10.2632; SAMPLE 94Zr: 10.828; SAMPLE 96Zr: 1.8101 | 1 | -8.993e-05 | -8.7005e-06 | 29.5694 | 6.5772 | 10.2628 | 0 | 10.8276 | 0 | 1.8099 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8319 | 29.5694 | 6.5772 | 10.2628 | 10.8276 | 1.8099 | 0.22243 | 0.34708 | 0.36618 | 0.061208 | 1.5604 | 1.6462 | 0.27517 | 1.055 | 0.17635 | 0.16716 | -1.5031 | -1.0582 | -1.0046 | -2.7935 | [10 empty] | 0 | 1.8319 | 0.21797 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.014108 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.38% | -0.65514 | 6.726 | 14.227 | [8 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.12% | -5.9114 | 1.3511 | 9.6823 | SRM 3169 | n | 18 | 18 | 18
Lab 3 | Feb 4, 2020 | 1 | 5 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: -9.6121e-06; SAMPLE 49Ti40Ar: -8.374e-06; SAMPLE 90Zr: 29.4063; SAMPLE 91Zr: 6.541; SAMPLE 92Zr: 10.2063; SAMPLE 94Zr: 10.7684; SAMPLE 96Zr: 1.8002 | 1 | -0.00015932 | -2.8864e-05 | 29.4056 | 6.5409 | 10.206 | 0 | 10.768 | 0 | 1.7999 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8326 | 29.4056 | 6.5409 | 10.206 | 10.768 | 1.7999 | 0.22244 | 0.34708 | 0.36619 | 0.06121 | 1.5603 | 1.6463 | 0.27518 | 1.0551 | 0.17636 | 0.16715 | -1.5031 | -1.0582 | -1.0046 | -2.7934 | [10 empty] | 0 | 1.8326 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014125 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.48% | 0.74303 | -11.9217 | -10.3052 | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.26% | 1.5974 | -5.337 | -3.1668 | SRM 3169 | 2SE | 2.2877 | 1.8659 | 5.7077
Lab 3 | Feb 4, 2020 | 1 | 6 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.5161e-05; SAMPLE 49Ti40Ar: 1.1109e-05; SAMPLE 90Zr: 29.5494; SAMPLE 91Zr: 6.5729; SAMPLE 92Zr: 10.2563; SAMPLE 94Zr: 10.8211; SAMPLE 96Zr: 1.809 | 1 | -0.00010455 | -9.3811e-06 | 29.5488 | 6.5728 | 10.2559 | 0 | 10.8207 | 0 | 1.8088 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8333 | 29.5488 | 6.5728 | 10.2559 | 10.8207 | 1.8088 | 0.22244 | 0.34708 | 0.3662 | 0.061212 | 1.5604 | 1.6463 | 0.27519 | 1.0551 | 0.17636 | 0.16716 | -1.5031 | -1.0582 | -1.0046 | -2.7934 | [10 empty] | 0 | 1.8333 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.014117 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100% | -7.2394 | -6.366 | -10.732 | [8 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99% | -3.1739 | -2.5887 | -4.5144 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 7 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.1549e-05; SAMPLE 49Ti40Ar: 9.8536e-06; SAMPLE 90Zr: 29.1802; SAMPLE 91Zr: 6.4909; SAMPLE 92Zr: 10.1284; SAMPLE 94Zr: 10.6866; SAMPLE 96Zr: 1.7865 | 1 | -0.00013816 | -1.0636e-05 | 29.1796 | 6.4907 | 10.1281 | 0 | 10.6862 | 0 | 1.7863 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8347 | 29.1796 | 6.4907 | 10.1281 | 10.6862 | 1.7863 | 0.22244 | 0.34709 | 0.36622 | 0.061217 | 1.5604 | 1.6464 | 0.27521 | 1.0551 | 0.17637 | 0.16716 | -1.5031 | -1.0582 | -1.0045 | -2.7933 | [10 empty] | 0 | 1.8347 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014118 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.8% | -9.5918 | -2.4354 | -9.3319 | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.25% | 5.049 | 0.22908 | 5.1839 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 8 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 7.9464e-06; SAMPLE 49Ti40Ar: 9.3543e-06; SAMPLE 90Zr: 29.5193; SAMPLE 91Zr: 6.5664; SAMPLE 92Zr: 10.2463; SAMPLE 94Zr: 10.8111; SAMPLE 96Zr: 1.8074 | 1 | -0.00014176 | -1.1135e-05 | 29.5186 | 6.5663 | 10.246 | 0 | 10.8107 | 0 | 1.8072 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8354 | 29.5186 | 6.5663 | 10.246 | 10.8107 | 1.8072 | 0.22245 | 0.3471 | 0.36623 | 0.061221 | 1.5604 | 1.6464 | 0.27522 | 1.0551 | 0.17638 | 0.16716 | -1.5031 | -1.0581 | -1.0045 | -2.7933 | [10 empty] | 0 | 1.8354 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24944 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43898 | -1.3885 | 0.014114 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.2% | 12.3514 | 6.7904 | -1.7828 | [8 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.82% | -4.5589 | -6.7181 | 7.2977 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 9 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.973e-05; SAMPLE 49Ti40Ar: 3.0602e-05; SAMPLE 90Zr: 29.4293; SAMPLE 91Zr: 6.5464; SAMPLE 92Zr: 10.2153; SAMPLE 94Zr: 10.7785; SAMPLE 96Zr: 1.802 | 1 | -8.9979e-05 | 1.0113e-05 | 29.4287 | 6.5462 | 10.2149 | 0 | 10.7782 | 0 | 1.8018 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8363 | 29.4287 | 6.5462 | 10.2149 | 10.7782 | 1.8018 | 0.22244 | 0.34711 | 0.36625 | 0.061225 | 1.5604 | 1.6465 | 0.27524 | 1.0551 | 0.17638 | 0.16717 | -1.5031 | -1.0581 | -1.0044 | -2.7932 | [10 empty] | 0 | 1.8363 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014116 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 99.9% | -14.2033 | 3.0566 | 0.99206 | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.45% | 1.0114 | -1.3722 | 11.3754 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 10 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 9.2913e-05; SAMPLE 49Ti40Ar: -2.7499e-06; SAMPLE 90Zr: 29.3979; SAMPLE 91Zr: 6.5395; SAMPLE 92Zr: 10.2045; SAMPLE 94Zr: 10.7675; SAMPLE 96Zr: 1.8002 | 1 | -5.6795e-05 | -2.324e-05 | 29.3973 | 6.5394 | 10.2041 | 0 | 10.7671 | 0 | 1.7999 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8372 | 29.3973 | 6.5394 | 10.2041 | 10.7671 | 1.7999 | 0.22245 | 0.34711 | 0.36626 | 0.061228 | 1.5604 | 1.6465 | 0.27525 | 1.0552 | 0.17639 | 0.16717 | -1.5031 | -1.0581 | -1.0044 | -2.7931 | [10 empty] | 0 | 1.8372 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014124 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.98% | 0.30494 | -6.7606 | 0.091153 | [8 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.75% | 3.4629 | 5.0603 | -20.8672 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 11 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.6695e-05; SAMPLE 49Ti40Ar: 7.1226e-06; SAMPLE 90Zr: 29.2239; SAMPLE 91Zr: 6.5009; SAMPLE 92Zr: 10.1445; SAMPLE 94Zr: 10.7045; SAMPLE 96Zr: 1.7897 | 1 | -0.00013301 | -1.3367e-05 | 29.2233 | 6.5007 | 10.1442 | 0 | 10.7041 | 0 | 1.7894 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8389 | 29.2233 | 6.5007 | 10.1442 | 10.7041 | 1.7894 | 0.22245 | 0.34713 | 0.36629 | 0.061233 | 1.5605 | 1.6466 | 0.27527 | 1.0552 | 0.1764 | 0.16717 | -1.5031 | -1.0581 | -1.0043 | -2.7931 | [10 empty] | 0 | 1.8389 | 0.21797 | 0.33336 | 0.3381 | 0.05437 | 1.5294 | 1.5512 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014116 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.59% | -1.6184 | 6.051 | -29.2726 | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.39% | -1.6124 | -7.9603 | -18.4919 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 12 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 3.1971e-05; SAMPLE 49Ti40Ar: -1.3955e-05; SAMPLE 90Zr: 29.2898; SAMPLE 91Zr: 6.5156; SAMPLE 92Zr: 10.1675; SAMPLE 94Zr: 10.729; SAMPLE 96Zr: 1.7939 | 1 | -0.00011774 | -3.4445e-05 | 29.2892 | 6.5154 | 10.1672 | 0 | 10.7287 | 0 | 1.7936 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8397 | 29.2892 | 6.5154 | 10.1672 | 10.7287 | 1.7936 | 0.22245 | 0.34713 | 0.3663 | 0.061238 | 1.5605 | 1.6467 | 0.27529 | 1.0552 | 0.17641 | 0.16718 | -1.503 | -1.0581 | -1.0043 | -2.793 | [10 empty] | 0 | 1.8397 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.4% | -9.2645 | -1.396 | 8.6175 | [8 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.4% | 4.716 | 0.16623 | -7.1154 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 13 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 3.8787e-05; SAMPLE 49Ti40Ar: 5.2531e-06; SAMPLE 90Zr: 28.9805; SAMPLE 91Zr: 6.4469; SAMPLE 92Zr: 10.0605; SAMPLE 94Zr: 10.6164; SAMPLE 96Zr: 1.7751 | 1 | -0.00011092 | -1.5237e-05 | 28.9799 | 6.4468 | 10.0602 | 0 | 10.6161 | 0 | 1.7748 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8412 | 28.9799 | 6.4468 | 10.0602 | 10.6161 | 1.7748 | 0.22246 | 0.34714 | 0.36633 | 0.061243 | 1.5605 | 1.6467 | 0.27531 | 1.0553 | 0.17642 | 0.16718 | -1.503 | -1.058 | -1.0042 | -2.7929 | [10 empty] | 0 | 1.8412 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.53% | 0.53586 | -1.0297 | -3.147 | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.31% | 4.1424 | 0.71719 | -2.7621 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 14 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 2.6428e-05; SAMPLE 49Ti40Ar: 2.964e-06; SAMPLE 90Zr: 29.537; SAMPLE 91Zr: 6.5708; SAMPLE 92Zr: 10.254; SAMPLE 94Zr: 10.8208; SAMPLE 96Zr: 1.8093 | 1 | -0.00012328 | -1.7526e-05 | 29.5364 | 6.5707 | 10.2537 | 0 | 10.8204 | 0 | 1.809 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8423 | 29.5364 | 6.5707 | 10.2537 | 10.8204 | 1.809 | 0.22246 | 0.34715 | 0.36634 | 0.061248 | 1.5605 | 1.6468 | 0.27532 | 1.0553 | 0.17643 | 0.16719 | -1.503 | -1.058 | -1.0042 | -2.7928 | [10 empty] | 0 | 1.8423 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.014114 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 101.3% | 6.4608 | 4.0978 | -2.3109 | [8 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.02% | -1.9139 | 3.3235 | 0.32492 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 15 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.3158e-05; SAMPLE 49Ti40Ar: 1.5157e-06; SAMPLE 90Zr: 29.422; SAMPLE 91Zr: 6.5453; SAMPLE 92Zr: 10.2143; SAMPLE 94Zr: 10.7793; SAMPLE 96Zr: 1.8024 | 1 | -0.00010655 | -1.8974e-05 | 29.4214 | 6.5452 | 10.2139 | 0 | 10.7789 | 0 | 1.8022 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8435 | 29.4214 | 6.5452 | 10.2139 | 10.7789 | 1.8022 | 0.22246 | 0.34716 | 0.36636 | 0.061253 | 1.5605 | 1.6468 | 0.27534 | 1.0553 | 0.17644 | 0.16719 | -1.503 | -1.058 | -1.0041 | -2.7927 | [10 empty] | 0 | 1.8435 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014124 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.5% | 6.7219 | -8.3217 | 18.0906 | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 98.09% | 4.7477 | 3.4747 | 14.8564 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 16 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 3.4699e-05; SAMPLE 49Ti40Ar: -3.5842e-05; SAMPLE 90Zr: 29.0296; SAMPLE 91Zr: 6.4581; SAMPLE 92Zr: 10.0784; SAMPLE 94Zr: 10.6361; SAMPLE 96Zr: 1.7785 | 1 | -0.00011501 | -5.6332e-05 | 29.0289 | 6.4579 | 10.0781 | 0 | 10.6357 | 0 | 1.7782 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8449 | 29.0289 | 6.4579 | 10.0781 | 10.6357 | 1.7782 | 0.22246 | 0.34717 | 0.36638 | 0.061257 | 1.5606 | 1.6469 | 0.27536 | 1.0553 | 0.17645 | 0.16719 | -1.503 | -1.0579 | -1.0041 | -2.7927 | [10 empty] | 0 | 1.8449 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014116 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 98.87% | -5.1388 | 1.6008 | -17.5875 | [8 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.01% | -8.93 | -7.3544 | 8.4222 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 17 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.389e-05; SAMPLE 49Ti40Ar: 1.6684e-05; SAMPLE 90Zr: 28.9952; SAMPLE 91Zr: 6.4506; SAMPLE 92Zr: 10.0668; SAMPLE 94Zr: 10.6243; SAMPLE 96Zr: 1.7766 | 1 | -9.5819e-05 | -3.8052e-06 | 28.9946 | 6.4504 | 10.0665 | 0 | 10.6239 | 0 | 1.7764 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8465 | 28.9946 | 6.4504 | 10.0665 | 10.6239 | 1.7764 | 0.22247 | 0.34718 | 0.36641 | 0.061265 | 1.5606 | 1.647 | 0.27539 | 1.0554 | 0.17646 | 0.1672 | -1.503 | -1.0579 | -1.004 | -2.7925 | [10 empty] | 0 | 1.8465 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.61% | 3.0792 | -1.2563 | 20.6223 | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.56% | -6.4549 | -5.649 | -9.9788 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 18 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.3781e-05; SAMPLE 49Ti40Ar: -1.0064e-05; SAMPLE 90Zr: 29.1885; SAMPLE 91Zr: 6.4936; SAMPLE 92Zr: 10.1341; SAMPLE 94Zr: 10.6955; SAMPLE 96Zr: 1.7886 | 1 | -0.00010593 | -3.0554e-05 | 29.1878 | 6.4935 | 10.1337 | 0 | 10.6952 | 0 | 1.7883 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8475 | 29.1878 | 6.4935 | 10.1337 | 10.6952 | 1.7883 | 0.22247 | 0.34719 | 0.36643 | 0.061269 | 1.5606 | 1.6471 | 0.2754 | 1.0554 | 0.17647 | 0.16721 | -1.503 | -1.0579 | -1.004 | -2.7925 | [10 empty] | 0 | 1.8475 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100% | 3.9178 | -2.2499 | 17.5847 | [8 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.56% | -5.0593 | 0.11489 | -15.1197 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 19 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 2.5378e-06; SAMPLE 49Ti40Ar: -2.1471e-05; SAMPLE 90Zr: 29.0032; SAMPLE 91Zr: 6.4524; SAMPLE 92Zr: 10.07; SAMPLE 94Zr: 10.6282; SAMPLE 96Zr: 1.7773 | 1 | -0.00014717 | -4.1961e-05 | 29.0026 | 6.4523 | 10.0697 | 0 | 10.6278 | 0 | 1.7771 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8486 | 29.0026 | 6.4523 | 10.0697 | 10.6278 | 1.7771 | 0.22247 | 0.3472 | 0.36644 | 0.061274 | 1.5606 | 1.6471 | 0.27542 | 1.0554 | 0.17648 | 0.16721 | -1.5029 | -1.0579 | -1.0039 | -2.7924 | [10 empty] | 0 | 1.8486 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.12% | -5.9114 | 1.3511 | 9.6823 | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.62% | 2.604 | 2.6373 | 24.793 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 20 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 9.7695e-06; SAMPLE 49Ti40Ar: 2.0633e-05; SAMPLE 90Zr: 29.3327; SAMPLE 91Zr: 6.5259; SAMPLE 92Zr: 10.1847; SAMPLE 94Zr: 10.7495; SAMPLE 96Zr: 1.7977 | 1 | -0.00013994 | 1.4303e-07 | 29.3321 | 6.5257 | 10.1844 | 0 | 10.7491 | 0 | 1.7974 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8499 | 29.3321 | 6.5257 | 10.1844 | 10.7491 | 1.7974 | 0.22248 | 0.34721 | 0.36646 | 0.061278 | 1.5606 | 1.6472 | 0.27543 | 1.0555 | 0.17649 | 0.16721 | -1.5029 | -1.0578 | -1.0039 | -2.7923 | [10 empty] | 0 | 1.8499 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.8% | -1.772 | 1.3651 | 1.7535
Lab 3 | Feb 4, 2020 | 1 | 21 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 3.8799e-05; SAMPLE 49Ti40Ar: -1.1781e-05; SAMPLE 90Zr: 29.3264; SAMPLE 91Zr: 6.5246; SAMPLE 92Zr: 10.1829; SAMPLE 94Zr: 10.7479; SAMPLE 96Zr: 1.7974 | 1 | -0.00011091 | -3.2271e-05 | 29.3257 | 6.5245 | 10.1825 | 0 | 10.7476 | 0 | 1.7971 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8515 | 29.3257 | 6.5245 | 10.1825 | 10.7476 | 1.7971 | 0.22248 | 0.34722 | 0.36649 | 0.061282 | 1.5607 | 1.6473 | 0.27545 | 1.0555 | 0.17649 | 0.16721 | -1.5029 | -1.0578 | -1.0038 | -2.7923 | [10 empty] | 0 | 1.8515 | 0.21797 | 0.33336 | 0.3381 | 0.05437 | 1.5294 | 1.5511 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014123 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.5% | 9.6758 | 1.4344 | -23.5506 | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.05% | -6.6903 | -2.7174 | 10.7629 | SRM 3169 | Avg | -5.407 | 2.0403 | 1.884
Lab 3 | Feb 4, 2020 | 1 | 22 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 9.6741e-05; SAMPLE 49Ti40Ar: -1.3223e-05; SAMPLE 90Zr: 29.0379; SAMPLE 91Zr: 6.4605; SAMPLE 92Zr: 10.0829; SAMPLE 94Zr: 10.6427; SAMPLE 96Zr: 1.7799 | 1 | -5.2967e-05 | -3.3712e-05 | 29.0373 | 6.4603 | 10.0826 | 0 | 10.6424 | 0 | 1.7796 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8526 | 29.0373 | 6.4603 | 10.0826 | 10.6424 | 1.7796 | 0.22248 | 0.34723 | 0.36651 | 0.061288 | 1.5607 | 1.6473 | 0.27547 | 1.0555 | 0.1765 | 0.16722 | -1.5029 | -1.0578 | -1.0037 | -2.7922 | [10 empty] | 0 | 1.8526 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.439 | -1.3885 | 0.014126 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.25% | 1.6555 | -6.0196 | -10.116 | [8 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.66% | -1.961 | -12.6449 | 25.1645 | SRM 3169 | 2SD | 19.0946 | 17.9771 | 50.3823
Lab 3 | Feb 4, 2020 | 1 | 23 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.0397e-05; SAMPLE 49Ti40Ar: -2.0392e-05; SAMPLE 90Zr: 28.9111; SAMPLE 91Zr: 6.4324; SAMPLE 92Zr: 10.0393; SAMPLE 94Zr: 10.5969; SAMPLE 96Zr: 1.7723 | 1 | -0.00013931 | -4.0882e-05 | 28.9105 | 6.4323 | 10.0389 | 0 | 10.5966 | 0 | 1.772 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.854 | 28.9105 | 6.4323 | 10.0389 | 10.5966 | 1.772 | 0.22249 | 0.34724 | 0.36653 | 0.061294 | 1.5607 | 1.6474 | 0.27549 | 1.0555 | 0.17652 | 0.16723 | -1.5029 | -1.0577 | -1.0037 | -2.7921 | [10 empty] | 0 | 1.854 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.31% | 12.2993 | 2.0736 | 6.4481 | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.63% | -4.8627 | 7.8351 | 25.0743 | SRM 3169 | n | 15 | 15 | 15
Lab 3 | Feb 4, 2020 | 1 | 24 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.7319e-05; SAMPLE 49Ti40Ar: 5.9361e-06; SAMPLE 90Zr: 29.1839; SAMPLE 91Zr: 6.4931; SAMPLE 92Zr: 10.1343; SAMPLE 94Zr: 10.6974; SAMPLE 96Zr: 1.7891 | 1 | -0.00010239 | -1.4554e-05 | 29.1833 | 6.493 | 10.1339 | 0 | 10.697 | 0 | 1.7889 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8551 | 29.1833 | 6.493 | 10.1339 | 10.697 | 1.7889 | 0.22249 | 0.34725 | 0.36655 | 0.061298 | 1.5607 | 1.6475 | 0.27551 | 1.0556 | 0.17652 | 0.16723 | -1.5029 | -1.0577 | -1.0036 | -2.792 | [10 empty] | 0 | 1.8551 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014117 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.8% | -4.3905 | 4.1697 | -5.0917 | [8 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.86% | -9.752 | 12.6985 | -22.1612 | SRM 3169 | 2SE | 4.9302 | 4.6417 | 13.0087
Lab 3 | Feb 4, 2020 | 1 | 25 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 2.7108e-05; SAMPLE 49Ti40Ar: -7.6926e-06; SAMPLE 90Zr: 28.5315; SAMPLE 91Zr: 6.3481; SAMPLE 92Zr: 9.9079; SAMPLE 94Zr: 10.4588; SAMPLE 96Zr: 1.7493 | 1 | -0.0001226 | -2.8182e-05 | 28.5309 | 6.3479 | 9.9076 | 0 | 10.4585 | 0 | 1.749 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8563 | 28.5309 | 6.3479 | 9.9076 | 10.4585 | 1.749 | 0.22249 | 0.34726 | 0.36657 | 0.061303 | 1.5608 | 1.6475 | 0.27553 | 1.0556 | 0.17653 | 0.16724 | -1.5029 | -1.0577 | -1.0036 | -2.7919 | [10 empty] | 0 | 1.8563 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.439 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.26% | 1.5974 | -5.337 | -3.1668 | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.92% | -14.6631 | -9.8627 | 3.4383 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 26 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.5405e-05; SAMPLE 49Ti40Ar: -5.8882e-07; SAMPLE 90Zr: 28.8908; SAMPLE 91Zr: 6.4281; SAMPLE 92Zr: 10.0329; SAMPLE 94Zr: 10.5909; SAMPLE 96Zr: 1.7714 | 1 | -0.0001343 | -2.1078e-05 | 28.8902 | 6.4279 | 10.0326 | 0 | 10.5905 | 0 | 1.7712 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.857 | 28.8902 | 6.4279 | 10.0326 | 10.5905 | 1.7712 | 0.2225 | 0.34727 | 0.36658 | 0.061307 | 1.5608 | 1.6476 | 0.27554 | 1.0556 | 0.17654 | 0.16724 | -1.5028 | -1.0577 | -1.0035 | -2.7919 | [10 empty] | 0 | 1.857 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.014117 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.39% | 5.3501 | 4.4688 | 20.7006 | [8 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.19% | -18.6254 | -1.0449 | -10.4323 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 27 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.931e-05; SAMPLE 49Ti40Ar: -2.7365e-05; SAMPLE 90Zr: 28.9276; SAMPLE 91Zr: 6.4363; SAMPLE 92Zr: 10.0458; SAMPLE 94Zr: 10.6047; SAMPLE 96Zr: 1.7737 | 1 | -0.0001004 | -4.7855e-05 | 28.927 | 6.4361 | 10.0455 | 0 | 10.6044 | 0 | 1.7735 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8579 | 28.927 | 6.4361 | 10.0455 | 10.6044 | 1.7735 | 0.2225 | 0.34727 | 0.36659 | 0.061309 | 1.5608 | 1.6476 | 0.27555 | 1.0556 | 0.17654 | 0.16724 | -1.5028 | -1.0577 | -1.0035 | -2.7918 | [10 empty] | 0 | 1.8579 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100% | -6.4198 | 0.43341 | -4.8855 | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.01% | -23.3863 | 10.4738 | 7.1893 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 28 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 8.0608e-05; SAMPLE 49Ti40Ar: -1.2971e-05; SAMPLE 90Zr: 28.9524; SAMPLE 91Zr: 6.4419; SAMPLE 92Zr: 10.0547; SAMPLE 94Zr: 10.6144; SAMPLE 96Zr: 1.7754 | 1 | -6.91e-05 | -3.3461e-05 | 28.9517 | 6.4418 | 10.0543 | 0 | 10.6141 | 0 | 1.7751 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8592 | 28.9517 | 6.4418 | 10.0543 | 10.6141 | 1.7751 | 0.2225 | 0.34728 | 0.36661 | 0.061314 | 1.5608 | 1.6477 | 0.27557 | 1.0557 | 0.17655 | 0.16724 | -1.5028 | -1.0576 | -1.0035 | -2.7918 | [10 empty] | 0 | 1.8592 | 0.21797 | 0.33336 | 0.3381 | 0.05437 | 1.5294 | 1.5512 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.439 | -1.3885 | 0.014126 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.2% | -2.3192 | -8.0462 | -16.8958 | [8 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.22% | 8.2655 | 9.9206 | -1.7241 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 29 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 9.0006e-07; SAMPLE 49Ti40Ar: 2.0671e-05; SAMPLE 90Zr: 28.8829; SAMPLE 91Zr: 6.4265; SAMPLE 92Zr: 10.0309; SAMPLE 94Zr: 10.5896; SAMPLE 96Zr: 1.7713 | 1 | -0.00014881 | 1.8095e-07 | 28.8823 | 6.4264 | 10.0305 | 0 | 10.5892 | 0 | 1.771 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8605 | 28.8823 | 6.4264 | 10.0305 | 10.5892 | 1.771 | 0.2225 | 0.34729 | 0.36663 | 0.06132 | 1.5608 | 1.6478 | 0.27559 | 1.0557 | 0.17657 | 0.16725 | -1.5028 | -1.0576 | -1.0034 | -2.7917 | [10 empty] | 0 | 1.8605 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014123 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.82% | -5.4721 | -1.4654 | 6.9443 | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.6% | 7.9493 | 10.581 | 39.476 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 30 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 3.8727e-05; SAMPLE 49Ti40Ar: -7.1896e-06; SAMPLE 90Zr: 28.9179; SAMPLE 91Zr: 6.4344; SAMPLE 92Zr: 10.0434; SAMPLE 94Zr: 10.603; SAMPLE 96Zr: 1.7736 | 1 | -0.00011098 | -2.7679e-05 | 28.9173 | 6.4343 | 10.043 | 0 | 10.6026 | 0 | 1.7733 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8618 | 28.9173 | 6.4343 | 10.043 | 10.6026 | 1.7733 | 0.22251 | 0.3473 | 0.36665 | 0.061325 | 1.5609 | 1.6478 | 0.27561 | 1.0557 | 0.17657 | 0.16725 | -1.5028 | -1.0576 | -1.0033 | -2.7916 | [10 empty] | 0 | 1.8618 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014118 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.84% | -2.7075 | 1.5819 | 1.1183 | [8 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | -15.4762 | -0.22744 | -2.3322 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 31 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: -7.3941e-06; SAMPLE 49Ti40Ar: -3.5091e-07; SAMPLE 90Zr: 28.6568; SAMPLE 91Zr: 6.3764; SAMPLE 92Zr: 9.9529; SAMPLE 94Zr: 10.5077; SAMPLE 96Zr: 1.7577 | 1 | -0.0001571 | -2.0841e-05 | 28.6562 | 6.3762 | 9.9525 | 0 | 10.5073 | 0 | 1.7574 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8627 | 28.6562 | 6.3762 | 9.9525 | 10.5073 | 1.7574 | 0.22251 | 0.34731 | 0.36667 | 0.061328 | 1.5609 | 1.6479 | 0.27562 | 1.0557 | 0.17658 | 0.16726 | -1.5028 | -1.0575 | -1.0033 | -2.7915 | [10 empty] | 0 | 1.8627 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99% | -3.1739 | -2.5887 | -4.5144 | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.7% | -4.3567 | 3.9379 | 22.0271 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 32 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 3.0476e-07; SAMPLE 49Ti40Ar: -8.7129e-07; SAMPLE 90Zr: 28.9759; SAMPLE 91Zr: 6.4475; SAMPLE 92Zr: 10.0641; SAMPLE 94Zr: 10.6254; SAMPLE 96Zr: 1.7774 | 1 | -0.0001494 | -2.1361e-05 | 28.9752 | 6.4474 | 10.0637 | 0 | 10.625 | 0 | 1.7772 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8643 | 28.9752 | 6.4474 | 10.0637 | 10.625 | 1.7772 | 0.22251 | 0.34732 | 0.36669 | 0.061335 | 1.5609 | 1.648 | 0.27564 | 1.0558 | 0.17659 | 0.16726 | -1.5028 | -1.0575 | -1.0032 | -2.7914 | [10 empty] | 0 | 1.8643 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.014114 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.4% | 1.5758 | 6.3546 | -8.3733 | [8 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 101% | -1.4651 | -8.9293 | -15.6156 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 33 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.3988e-05; SAMPLE 49Ti40Ar: -1.2216e-05; SAMPLE 90Zr: 29.0238; SAMPLE 91Zr: 6.4583; SAMPLE 92Zr: 10.081; SAMPLE 94Zr: 10.6436; SAMPLE 96Zr: 1.7806 | 1 | -9.5721e-05 | -3.2706e-05 | 29.0232 | 6.4582 | 10.0807 | 0 | 10.6433 | 0 | 1.7803 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8657 | 29.0232 | 6.4582 | 10.0807 | 10.6433 | 1.7803 | 0.22252 | 0.34733 | 0.36672 | 0.061341 | 1.5609 | 1.648 | 0.27567 | 1.0558 | 0.17661 | 0.16727 | -1.5027 | -1.0575 | -1.0032 | -2.7913 | [10 empty] | 0 | 1.8657 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.4% | 2.397 | -0.27255 | -0.3493 | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | -3.9123 | -7.9507 | 31.9688 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 34 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 6.1344e-05; SAMPLE 49Ti40Ar: -3.7026e-05; SAMPLE 90Zr: 28.8263; SAMPLE 91Zr: 6.4145; SAMPLE 92Zr: 10.0127; SAMPLE 94Zr: 10.5718; SAMPLE 96Zr: 1.7686 | 1 | -8.8365e-05 | -5.7516e-05 | 28.8256 | 6.4143 | 10.0123 | 0 | 10.5715 | 0 | 1.7684 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8671 | 28.8256 | 6.4143 | 10.0123 | 10.5715 | 1.7684 | 0.22252 | 0.34734 | 0.36674 | 0.061347 | 1.5609 | 1.6481 | 0.27569 | 1.0558 | 0.17662 | 0.16728 | -1.5027 | -1.0574 | -1.0031 | -2.7912 | [10 empty] | 0 | 1.8671 | 0.21797 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.43% | 4.2847 | -5.4986 | 19.4169 | [8 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | 4.3298 | 4.0267 | -43.2574 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 35 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 2.3969e-05; SAMPLE 49Ti40Ar: -4.4662e-06; SAMPLE 90Zr: 28.8321; SAMPLE 91Zr: 6.4158; SAMPLE 92Zr: 10.015; SAMPLE 94Zr: 10.5746; SAMPLE 96Zr: 1.7691 | 1 | -0.00012574 | -2.4956e-05 | 28.8315 | 6.4157 | 10.0146 | 0 | 10.5742 | 0 | 1.7688 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8683 | 28.8315 | 6.4157 | 10.0146 | 10.5742 | 1.7688 | 0.22252 | 0.34735 | 0.36676 | 0.061351 | 1.561 | 1.6482 | 0.27571 | 1.0559 | 0.17663 | 0.16728 | -1.5027 | -1.0574 | -1.0031 | -2.7911 | [10 empty] | 0 | 1.8683 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.71% | -4.3392 | 0.08556 | -13.1055 | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 101.2% | 3.5013 | 14.5076 | -41.3192 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 36 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.6949e-05; SAMPLE 49Ti40Ar: -6.4572e-06; SAMPLE 90Zr: 29.0055; SAMPLE 91Zr: 6.4545; SAMPLE 92Zr: 10.0755; SAMPLE 94Zr: 10.6388; SAMPLE 96Zr: 1.7799 | 1 | -9.2759e-05 | -2.6947e-05 | 29.0048 | 6.4543 | 10.0752 | 0 | 10.6384 | 0 | 1.7796 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8697 | 29.0048 | 6.4543 | 10.0752 | 10.6384 | 1.7796 | 0.22253 | 0.34736 | 0.36678 | 0.061357 | 1.561 | 1.6483 | 0.27573 | 1.0559 | 0.17664 | 0.16728 | -1.5027 | -1.0574 | -1.003 | -2.7911 | [10 empty] | 0 | 1.8697 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.3% | -2.8833 | 1.7275 | 1.7652
Lab 3 | Feb 4, 2020 | 1 | 37 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: -2.7071e-06; SAMPLE 49Ti40Ar: 2.9757e-06; SAMPLE 90Zr: 28.7786; SAMPLE 91Zr: 6.4041; SAMPLE 92Zr: 9.997; SAMPLE 94Zr: 10.5562; SAMPLE 96Zr: 1.7661 | 1 | -0.00015242 | -1.7514e-05 | 28.778 | 6.404 | 9.9967 | 0 | 10.5558 | 0 | 1.7659 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8711 | 28.778 | 6.404 | 9.9967 | 10.5558 | 1.7659 | 0.22253 | 0.34737 | 0.3668 | 0.061362 | 1.561 | 1.6483 | 0.27575 | 1.0559 | 0.17665 | 0.16729 | -1.5027 | -1.0574 | -1.0029 | -2.791 | [10 empty] | 0 | 1.8711 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.25% | 5.049 | 0.22908 | 5.1839 | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 98.77% | 14.7698 | -3.8197 | 9.2462 | SRM 3169 | Avg | 0.58319 | -2.3577 | -3.7591
Lab 3 | Feb 4, 2020 | 1 | 38 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 9.549e-06; SAMPLE 49Ti40Ar: 2.0068e-05; SAMPLE 90Zr: 28.9851; SAMPLE 91Zr: 6.4501; SAMPLE 92Zr: 10.069; SAMPLE 94Zr: 10.6326; SAMPLE 96Zr: 1.7789 | 1 | -0.00014016 | -4.2153e-07 | 28.9845 | 6.45 | 10.0687 | 0 | 10.6322 | 0 | 1.7787 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8724 | 28.9845 | 6.45 | 10.0687 | 10.6322 | 1.7787 | 0.22253 | 0.34738 | 0.36682 | 0.061366 | 1.561 | 1.6484 | 0.27576 | 1.056 | 0.17665 | 0.16729 | -1.5027 | -1.0573 | -1.0029 | -2.7909 | [10 empty] | 0 | 1.8724 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.1% | -3.2531 | 1.1435 | -5.9106 | [8 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 101.4% | -2.3395 | -4.4268 | -20.5948 | SRM 3169 | 2SD | 14.2808 | 14.0589 | 33.4863
Lab 3 | Feb 4, 2020 | 1 | 39 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.6245e-05; SAMPLE 49Ti40Ar: -1.2846e-05; SAMPLE 90Zr: 29.0706; SAMPLE 91Zr: 6.4693; SAMPLE 92Zr: 10.0991; SAMPLE 94Zr: 10.6648; SAMPLE 96Zr: 1.7844 | 1 | -9.3464e-05 | -3.3336e-05 | 29.07 | 6.4692 | 10.0988 | 0 | 10.6645 | 0 | 1.7842 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8744 | 29.07 | 6.4692 | 10.0988 | 10.6645 | 1.7842 | 0.22254 | 0.3474 | 0.36685 | 0.061374 | 1.5611 | 1.6485 | 0.27579 | 1.056 | 0.17667 | 0.1673 | -1.5027 | -1.0573 | -1.0028 | -2.7908 | [10 empty] | 0 | 1.8744 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.439 | -1.3885 | 0.014125 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 100.4% | -1.1831 | -2.8879 | 8.6975 | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.96% | 8.2102 | 7.1171 | 7.2745 | SRM 3169 | n | 9 | 9 | 9
Lab 3 | Feb 4, 2020 | 1 | 40 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.9379e-05; SAMPLE 49Ti40Ar: -6.3975e-06; SAMPLE 90Zr: 28.9008; SAMPLE 91Zr: 6.4317; SAMPLE 92Zr: 10.0405; SAMPLE 94Zr: 10.6032; SAMPLE 96Zr: 1.7741 | 1 | -9.033e-05 | -2.6887e-05 | 28.9002 | 6.4315 | 10.0401 | 0 | 10.6028 | 0 | 1.7739 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8758 | 28.9002 | 6.4315 | 10.0401 | 10.6028 | 1.7739 | 0.22254 | 0.34741 | 0.36688 | 0.061379 | 1.5611 | 1.6486 | 0.27581 | 1.056 | 0.17668 | 0.1673 | -1.5026 | -1.0573 | -1.0027 | -2.7907 | [10 empty] | 0 | 1.8758 | 0.21797 | 0.33336 | 0.3381 | 0.05437 | 1.5294 | 1.5511 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014124 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100% | 2.3987 | -3.8894 | -5.4993 | [8 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.17% | 2.0473 | 5.6897 | 8.0355 | SRM 3169 | 2SE | 4.7603 | 4.6863 | 11.1621
Lab 3 | Feb 4, 2020 | 1 | 41 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 3.7878e-05; SAMPLE 49Ti40Ar: 4.8264e-06; SAMPLE 90Zr: 28.7423; SAMPLE 91Zr: 6.3965; SAMPLE 92Zr: 9.9858; SAMPLE 94Zr: 10.5459; SAMPLE 96Zr: 1.7646 | 1 | -0.00011183 | -1.5663e-05 | 28.7416 | 6.3963 | 9.9855 | 0 | 10.5455 | 0 | 1.7644 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8776 | 28.7416 | 6.3963 | 9.9855 | 10.5455 | 1.7644 | 0.22255 | 0.34742 | 0.36691 | 0.061387 | 1.5611 | 1.6487 | 0.27584 | 1.0561 | 0.17669 | 0.16731 | -1.5026 | -1.0572 | -1.0026 | -2.7906 | [10 empty] | 0 | 1.8776 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.014118 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.62% | -1.4742 | 3.4491 | 5.2311 | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 101.6% | -5.0379 | -0.87205 | 9.4707 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 42 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 6.0121e-05; SAMPLE 49Ti40Ar: 5.036e-06; SAMPLE 90Zr: 28.8042; SAMPLE 91Zr: 6.4104; SAMPLE 92Zr: 10.0076; SAMPLE 94Zr: 10.5691; SAMPLE 96Zr: 1.7685 | 1 | -8.9588e-05 | -1.5454e-05 | 28.8036 | 6.4102 | 10.0072 | 0 | 10.5687 | 0 | 1.7683 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8787 | 28.8036 | 6.4102 | 10.0072 | 10.5687 | 1.7683 | 0.22255 | 0.34743 | 0.36692 | 0.061391 | 1.5611 | 1.6487 | 0.27585 | 1.0561 | 0.1767 | 0.16731 | -1.5026 | -1.0572 | -1.0026 | -2.7905 | [10 empty] | 0 | 1.8787 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5511 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.66% | 0.80296 | -2.4608 | -12.4965 | [8 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.81% | -2.978 | -4.8659 | 7.0344 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 43 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 2.6945e-05; SAMPLE 49Ti40Ar: -2.8348e-06; SAMPLE 90Zr: 28.513; SAMPLE 91Zr: 6.3456; SAMPLE 92Zr: 9.9068; SAMPLE 94Zr: 10.463; SAMPLE 96Zr: 1.7509 | 1 | -0.00012276 | -2.3324e-05 | 28.5123 | 6.3455 | 9.9064 | 0 | 10.4626 | 0 | 1.7506 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8804 | 28.5123 | 6.3455 | 9.9064 | 10.4626 | 1.7506 | 0.22255 | 0.34744 | 0.36695 | 0.061399 | 1.5612 | 1.6488 | 0.27589 | 1.0561 | 0.17672 | 0.16732 | -1.5026 | -1.0572 | -1.0025 | -2.7904 | [10 empty] | 0 | 1.8804 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 98.82% | -4.5589 | -6.7181 | 7.2977 | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 98.8% | -1.51 | -12.3758 | -3.3747 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 44 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 9.4977e-06; SAMPLE 49Ti40Ar: 6.1786e-06; SAMPLE 90Zr: 28.9046; SAMPLE 91Zr: 6.4329; SAMPLE 92Zr: 10.0432; SAMPLE 94Zr: 10.6073; SAMPLE 96Zr: 1.7751 | 1 | -0.00014021 | -1.4311e-05 | 28.904 | 6.4328 | 10.0429 | 0 | 10.6069 | 0 | 1.7748 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8817 | 28.904 | 6.4328 | 10.0429 | 10.6069 | 1.7748 | 0.22256 | 0.34746 | 0.36697 | 0.061405 | 1.5612 | 1.6489 | 0.27591 | 1.0562 | 0.17673 | 0.16733 | -1.5026 | -1.0571 | -1.0025 | -2.7903 | [10 empty] | 0 | 1.8817 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43899 | -1.3885 | 0.01411 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.5% | -0.71756 | 11.0962 | 15.2357 | [8 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 101.8% | 0.93657 | 3.9495 | -13.3058 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 45 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.8306e-05; SAMPLE 49Ti40Ar: 1.2169e-05; SAMPLE 90Zr: 28.8154; SAMPLE 91Zr: 6.4131; SAMPLE 92Zr: 10.0125; SAMPLE 94Zr: 10.5755; SAMPLE 96Zr: 1.7698 | 1 | -0.0001014 | -8.3208e-06 | 28.8148 | 6.413 | 10.0122 | 0 | 10.5751 | 0 | 1.7696 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8836 | 28.8148 | 6.413 | 10.0122 | 10.5751 | 1.7696 | 0.22256 | 0.34747 | 0.367 | 0.061413 | 1.5612 | 1.649 | 0.27594 | 1.0562 | 0.17674 | 0.16734 | -1.5026 | -1.0571 | -1.0024 | -2.7901 | [10 empty] | 0 | 1.8836 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014125 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 99.99% | -10.6318 | -8.4965 | 3.8123 | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.84% | -8.8498 | -11.6158 | -37.6179 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 46 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 0.00010892; SAMPLE 49Ti40Ar: 2.6175e-05; SAMPLE 90Zr: 28.7344; SAMPLE 91Zr: 6.3953; SAMPLE 92Zr: 9.9847; SAMPLE 94Zr: 10.5464; SAMPLE 96Zr: 1.765 | 1 | -4.0785e-05 | 5.6854e-06 | 28.7338 | 6.3951 | 9.9844 | 0 | 10.546 | 0 | 1.7648 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.885 | 28.7338 | 6.3951 | 9.9844 | 10.546 | 1.7648 | 0.22257 | 0.34748 | 0.36703 | 0.061418 | 1.5612 | 1.6491 | 0.27596 | 1.0563 | 0.17675 | 0.16734 | -1.5025 | -1.0571 | -1.0023 | -2.7901 | [10 empty] | 0 | 1.885 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014123 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.98% | 5.3237 | -7.3137 | 0.004329
Lab 3 | Feb 4, 2020 | 1 | 47 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 5.2475e-06; SAMPLE 49Ti40Ar: -2.3151e-06; SAMPLE 90Zr: 28.2979; SAMPLE 91Zr: 6.2983; SAMPLE 92Zr: 9.8335; SAMPLE 94Zr: 10.3872; SAMPLE 96Zr: 1.7384 | 1 | -0.00014446 | -2.2805e-05 | 28.2973 | 6.2981 | 9.8332 | 0 | 10.3869 | 0 | 1.7382 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8874 | 28.2973 | 6.2981 | 9.8332 | 10.3869 | 1.7382 | 0.22257 | 0.3475 | 0.36706 | 0.061426 | 1.5613 | 1.6492 | 0.27599 | 1.0563 | 0.17677 | 0.16735 | -1.5025 | -1.057 | -1.0022 | -2.7899 | [10 empty] | 0 | 1.8874 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014123 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 98.75% | -2.5364 | -0.89683 | -14.4114 | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -0.84304 | 1.1443 | -8.3872 | SRM 3169 | Avg | -0.011466 | 2.742 | 0.18722
Lab 3 | Feb 4, 2020 | 1 | 48 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.2095e-05; SAMPLE 49Ti40Ar: 2.1568e-05; SAMPLE 90Zr: 28.5761; SAMPLE 91Zr: 6.3602; SAMPLE 92Zr: 9.9305; SAMPLE 94Zr: 10.49; SAMPLE 96Zr: 1.7557 | 1 | -0.00010761 | 1.078e-06 | 28.5754 | 6.3601 | 9.9302 | 0 | 10.4896 | 0 | 1.7555 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8888 | 28.5754 | 6.3601 | 9.9302 | 10.4896 | 1.7555 | 0.22257 | 0.34751 | 0.36709 | 0.061433 | 1.5613 | 1.6493 | 0.27601 | 1.0563 | 0.17678 | 0.16735 | -1.5025 | -1.057 | -1.0022 | -2.7898 | [10 empty] | 0 | 1.8888 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.54% | -6.6654 | 0.60882 | 10.3551 | [8 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | 0.46612 | 4.4106 | -23.9164 | SRM 3169 | 2SD | 11.0916 | 22.8406 | 36.4775
Lab 3 | Feb 4, 2020 | 1 | 49 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 2.35e-05; SAMPLE 49Ti40Ar: -6.9569e-06; SAMPLE 90Zr: 28.4712; SAMPLE 91Zr: 6.3371; SAMPLE 92Zr: 9.8945; SAMPLE 94Zr: 10.4525; SAMPLE 96Zr: 1.7495 | 1 | -0.00012621 | -2.7447e-05 | 28.4706 | 6.3369 | 9.8942 | 0 | 10.4521 | 0 | 1.7493 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8909 | 28.4706 | 6.3369 | 9.8942 | 10.4521 | 1.7493 | 0.22258 | 0.34752 | 0.36712 | 0.061441 | 1.5614 | 1.6494 | 0.27604 | 1.0564 | 0.1768 | 0.16736 | -1.5025 | -1.0569 | -1.0021 | -2.7897 | [10 empty] | 0 | 1.8909 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.45% | 1.0114 | -1.3722 | 11.3754 | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | 5.0183 | -6.6714 | -28.0897 | SRM 3169 | n | 14 | 14 | 14
Lab 3 | Feb 4, 2020 | 1 | 50 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.0891e-05; SAMPLE 49Ti40Ar: 1.8044e-06; SAMPLE 90Zr: 28.6818; SAMPLE 91Zr: 6.384; SAMPLE 92Zr: 9.968; SAMPLE 94Zr: 10.5303; SAMPLE 96Zr: 1.7625 | 1 | -0.00013882 | -1.8685e-05 | 28.6812 | 6.3839 | 9.9676 | 0 | 10.5299 | 0 | 1.7623 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.892 | 28.6812 | 6.3839 | 9.9676 | 10.5299 | 1.7623 | 0.22258 | 0.34753 | 0.36714 | 0.061444 | 1.5614 | 1.6495 | 0.27605 | 1.0564 | 0.1768 | 0.16736 | -1.5025 | -1.0569 | -1.002 | -2.7896 | [10 empty] | 0 | 1.892 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.7% | 1.2097 | 2.2562 | -22.1904 | [8 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.2% | -5.1419 | 5.6344 | 27.1017 | SRM 3169 | 2SE | 2.9643 | 6.1044 | 9.749
Lab 3 | Feb 4, 2020 | 1 | 51 | iRM4 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.7197e-05; SAMPLE 49Ti40Ar: -1.8641e-05; SAMPLE 90Zr: 28.5113; SAMPLE 91Zr: 6.3462; SAMPLE 92Zr: 9.9092; SAMPLE 94Zr: 10.4687; SAMPLE 96Zr: 1.7523 | 1 | -0.00013251 | -3.913e-05 | 28.5107 | 6.3461 | 9.9088 | 0 | 10.4683 | 0 | 1.752 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8941 | 28.5107 | 6.3461 | 9.9088 | 10.4683 | 1.752 | 0.22259 | 0.34755 | 0.36717 | 0.061452 | 1.5614 | 1.6496 | 0.27608 | 1.0565 | 0.17682 | 0.16737 | -1.5024 | -1.0569 | -1.0019 | -2.7895 | [10 empty] | 0 | 1.8941 | 0.21797 | 0.33336 | 0.3381 | 0.05437 | 1.5294 | 1.5511 | 0.24944 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Feb 4, 2020 | 1 | 99.92% | 2.1509 | 1.022 | -24.568 | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.5% | -1.6577 | 10.2111 | -4.4936 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 52 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 6.1834e-05; SAMPLE 49Ti40Ar: 4.7969e-06; SAMPLE 90Zr: 28.3871; SAMPLE 91Zr: 6.3187; SAMPLE 92Zr: 9.8663; SAMPLE 94Zr: 10.4237; SAMPLE 96Zr: 1.7449 | 1 | -8.7874e-05 | -1.5693e-05 | 28.3864 | 6.3185 | 9.8659 | 0 | 10.4234 | 0 | 1.7446 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8957 | 28.3864 | 6.3185 | 9.8659 | 10.4234 | 1.7446 | 0.22259 | 0.34756 | 0.3672 | 0.06146 | 1.5614 | 1.6496 | 0.27611 | 1.0565 | 0.17683 | 0.16738 | -1.5024 | -1.0568 | -1.0019 | -2.7894 | [10 empty] | 0 | 1.8957 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.439 | -1.3885 | 0.014124 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.09% | -2.493 | -2.5694 | -1.5906 | [8 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.4% | 7.7493 | 14.335 | 7.543 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 53 | iRM5 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 4.232e-05; SAMPLE 49Ti40Ar: 3.1742e-06; SAMPLE 90Zr: 28.2757; SAMPLE 91Zr: 6.294; SAMPLE 92Zr: 9.828; SAMPLE 94Zr: 10.3837; SAMPLE 96Zr: 1.7382 | 1 | -0.00010739 | -1.7315e-05 | 28.2751 | 6.2939 | 9.8277 | 0 | 10.3833 | 0 | 1.738 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8975 | 28.2751 | 6.2939 | 9.8277 | 10.3833 | 1.738 | 0.22259 | 0.34757 | 0.36722 | 0.061466 | 1.5615 | 1.6497 | 0.27614 | 1.0565 | 0.17684 | 0.16738 | -1.5024 | -1.0568 | -1.0018 | -2.7893 | [10 empty] | 0 | 1.8975 | 0.21797 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Feb 4, 2020 | 1 | 99.21% | -6.2846 | 3.3272 | -26.0889 | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.6% | 2.4594 | 4.838 | 29.3741 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 54 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 2.4228e-05; SAMPLE 49Ti40Ar: -3.3601e-07; SAMPLE 90Zr: 28.6155; SAMPLE 91Zr: 6.3698; SAMPLE 92Zr: 9.9464; SAMPLE 94Zr: 10.5091; SAMPLE 96Zr: 1.7593 | 1 | -0.00012548 | -2.0826e-05 | 28.6148 | 6.3696 | 9.946 | 0 | 10.5087 | 0 | 1.7591 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.8989 | 28.6148 | 6.3696 | 9.946 | 10.5087 | 1.7591 | 0.2226 | 0.34758 | 0.36725 | 0.061474 | 1.5615 | 1.6498 | 0.27617 | 1.0566 | 0.17686 | 0.16739 | -1.5024 | -1.0567 | -1.0017 | -2.7891 | [10 empty] | 0 | 1.8989 | 0.21797 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014123 | -1.8133 | -1.8275 | [6 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.8% | 2.9709 | 4.1042 | 22.3398 | [8 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.2% | 3.7215 | 28.2253 | -3.2707 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 55 | iRM6 | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.5459e-05; SAMPLE 49Ti40Ar: -6.7787e-06; SAMPLE 90Zr: 28.4423; SAMPLE 91Zr: 6.3314; SAMPLE 92Zr: 9.8866; SAMPLE 94Zr: 10.4462; SAMPLE 96Zr: 1.7488 | 1 | -0.00013425 | -2.7268e-05 | 28.4417 | 6.3312 | 9.8862 | 0 | 10.4458 | 0 | 1.7486 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9004 | 28.4417 | 6.3312 | 9.8862 | 10.4458 | 1.7486 | 0.2226 | 0.3476 | 0.36727 | 0.061479 | 1.5615 | 1.6499 | 0.27618 | 1.0566 | 0.17687 | 0.16739 | -1.5024 | -1.0567 | -1.0017 | -2.7891 | [10 empty] | 0 | 1.9004 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42487 | 0.43899 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM6 | Lab 3 | Feb 4, 2020 | 1 | 99.75% | 3.4629 | 5.0603 | -20.8672 | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 3.9118 | -3.7416 | -9.8633 | SRM 3169
Lab 3 | Feb 4, 2020 | 1 | 56 | STD | OPZ 177Hf++: 0.00014971; OPZ 49Ti40Ar: 2.049e-05; OPZ 90Zr: 0.00063448; OPZ 91Zr: 0.00014659; OPZ 92Zr: 0.0003432; OPZ 94Zr: 0.00037492; OPZ 96Zr: 0.00024974; SAMPLE 177Hf++: 1.9521e-05; SAMPLE 49Ti40Ar: -1.3471e-06; SAMPLE 90Zr: 28.4146; SAMPLE 91Zr: 6.3253; SAMPLE 92Zr: 9.8772; SAMPLE 94Zr: 10.4368; SAMPLE 96Zr: 1.7473 | 1 | -0.00013019 | -2.1837e-05 | 28.4139 | 6.3252 | 9.8769 | 0 | 10.4364 | 0 | 1.7471 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9023 | 28.4139 | 6.3252 | 9.8769 | 10.4364 | 1.7471 | 0.22261 | 0.34761 | 0.3673 | 0.061487 | 1.5615 | 1.65 | 0.27621 | 1.0567 | 0.17689 | 0.1674 | -1.5023 | -1.0567 | -1.0016 | -2.7889 | [10 empty] | 0 | 1.9023 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5511 | 0.24945 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42486 | 0.43899 | -1.3885 | 0.014129 | -1.8134 | -1.8275 | [22 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.6% | 5.1132 | -10.9547 | 2.2026 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 2 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.1002e-05; SAMPLE 49Ti40Ar: 4.4936e-06; SAMPLE 90Zr: 24.4296; SAMPLE 91Zr: 5.4461; SAMPLE 92Zr: 8.5171; SAMPLE 94Zr: 9.0258; SAMPLE 96Zr: 1.5155 | 1 | -4.6275e-05 | 4.5459e-07 | 24.4296 | 5.4461 | 8.517 | 0 | 9.0256 | 0 | 1.5154 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0365 | 24.4296 | 5.4461 | 8.517 | 9.0256 | 1.5154 | 0.22293 | 0.34864 | 0.36945 | 0.062032 | 1.5639 | 1.6573 | 0.27826 | 1.0597 | 0.17793 | 0.1679 | -1.5009 | -1.0537 | -0.99573 | -2.7801 | [9 empty] | L3 FebS2 | 0 | 2.0365 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.014129 | -1.8132 | -1.8273 | [22 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.6% | -6.8493 | 14.0345 | -11.5376 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 3 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 6.2424e-05; SAMPLE 49Ti40Ar: 1.2373e-05; SAMPLE 90Zr: 23.5822; SAMPLE 91Zr: 5.2574; SAMPLE 92Zr: 8.222; SAMPLE 94Zr: 8.7135; SAMPLE 96Zr: 1.4631 | 1 | -3.4852e-05 | 8.3335e-06 | 23.5822 | 5.2574 | 8.2219 | 0 | 8.7133 | 0 | 1.463 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0385 | 23.5822 | 5.2574 | 8.2219 | 8.7133 | 1.463 | 0.22294 | 0.34865 | 0.36949 | 0.06204 | 1.5639 | 1.6573 | 0.27828 | 1.0598 | 0.17794 | 0.16791 | -1.5009 | -1.0537 | -0.99564 | -2.78 | [10 empty] | 0 | 2.0385 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42488 | 0.43902 | -1.3883 | 0.014135 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 96.97% | 10.9909 | -6.2046 | -2.6194 | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.5% | -1.3583 | -7.9294 | -7.6742 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 4 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 4.4737e-05; SAMPLE 49Ti40Ar: -1.0671e-05; SAMPLE 90Zr: 24.2077; SAMPLE 91Zr: 5.3969; SAMPLE 92Zr: 8.4405; SAMPLE 94Zr: 8.9453; SAMPLE 96Zr: 1.5021 | 1 | -5.2539e-05 | -1.471e-05 | 24.2077 | 5.3969 | 8.4404 | 0 | 8.9451 | 0 | 1.502 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0404 | 24.2077 | 5.3969 | 8.4404 | 8.9451 | 1.502 | 0.22294 | 0.34867 | 0.36952 | 0.062048 | 1.5639 | 1.6575 | 0.27832 | 1.0598 | 0.17796 | 0.16792 | -1.5008 | -1.0536 | -0.99556 | -2.7798 | [10 empty] | 0 | 2.0404 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.014129 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.5% | -4.8883 | 4.8246 | -8.6769 | [8 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.5% | 0.44756 | -13.2895 | 3.6109 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 5 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 6.0664e-05; SAMPLE 49Ti40Ar: -4.177e-06; SAMPLE 90Zr: 23.5317; SAMPLE 91Zr: 5.2463; SAMPLE 92Zr: 8.205; SAMPLE 94Zr: 8.6961; SAMPLE 96Zr: 1.4604 | 1 | -3.6612e-05 | -8.216e-06 | 23.5316 | 5.2463 | 8.2049 | 0 | 8.6959 | 0 | 1.4603 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0419 | 23.5316 | 5.2463 | 8.2049 | 8.6959 | 1.4603 | 0.22295 | 0.34867 | 0.36954 | 0.062056 | 1.5639 | 1.6575 | 0.27834 | 1.0598 | 0.17798 | 0.16793 | -1.5008 | -1.0536 | -0.99549 | -2.7797 | [10 empty] | 0 | 2.0419 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.42488 | 0.43902 | -1.3883 | 0.014138 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 98.12% | -1.8073 | -4.1721 | 12.9664 | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.4% | -13.1974 | -1.8582 | 30.0214 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 6 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: -1.0339e-05; SAMPLE 49Ti40Ar: -2.4332e-06; SAMPLE 90Zr: 23.7557; SAMPLE 91Zr: 5.2964; SAMPLE 92Zr: 8.2834; SAMPLE 94Zr: 8.7796; SAMPLE 96Zr: 1.4745 | 1 | -0.00010762 | -6.4723e-06 | 23.7557 | 5.2964 | 8.2833 | 0 | 8.7794 | 0 | 1.4743 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0438 | 23.7557 | 5.2964 | 8.2833 | 8.7794 | 1.4743 | 0.22295 | 0.34869 | 0.36957 | 0.062063 | 1.564 | 1.6576 | 0.27837 | 1.0599 | 0.17799 | 0.16793 | -1.5008 | -1.0536 | -0.99541 | -2.7796 | [10 empty] | 0 | 2.0438 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.42488 | 0.43902 | -1.3883 | 0.014138 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.03% | 7.8548 | -10.9401 | 21.6619 | [8 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 84.89% | -11.6225 | 6.7419 | -53.7544 | SRM 3169 | Rejected for intensity mismatch > 12%
Lab 3 | Feb 5, 2020 | 2 | 7 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: -1.9308e-06; SAMPLE 49Ti40Ar: 1.149e-05; SAMPLE 90Zr: 23.1428; SAMPLE 91Zr: 5.1598; SAMPLE 92Zr: 8.0699; SAMPLE 94Zr: 8.5536; SAMPLE 96Zr: 1.4365 | 1 | -9.9207e-05 | 7.4512e-06 | 23.1427 | 5.1598 | 8.0698 | 0 | 8.5534 | 0 | 1.4364 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0452 | 23.1427 | 5.1598 | 8.0698 | 8.5534 | 1.4364 | 0.22295 | 0.3487 | 0.36959 | 0.062067 | 1.564 | 1.6577 | 0.27838 | 1.0599 | 0.178 | 0.16793 | -1.5008 | -1.0535 | -0.99535 | -2.7795 | [10 empty] | 0 | 2.0452 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9118 | 0.42488 | 0.43902 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.39% | -1.6124 | -7.9603 | -18.4919
Lab 3 | Feb 5, 2020 | 2 | 8 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 2.5774e-05; SAMPLE 49Ti40Ar: 1.3055e-05; SAMPLE 90Zr: 23.769; SAMPLE 91Zr: 5.2995; SAMPLE 92Zr: 8.2886; SAMPLE 94Zr: 8.7855; SAMPLE 96Zr: 1.4755 | 1 | -7.1502e-05 | 9.0158e-06 | 23.7689 | 5.2994 | 8.2885 | 0 | 8.7854 | 0 | 1.4754 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0466 | 23.7689 | 5.2994 | 8.2885 | 8.7854 | 1.4754 | 0.22296 | 0.34871 | 0.36962 | 0.062072 | 1.564 | 1.6578 | 0.27841 | 1.0599 | 0.178 | 0.16794 | -1.5008 | -1.0535 | -0.99529 | -2.7795 | [10 empty] | 0 | 2.0466 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014126 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | -3.0579 | 14.0674 | -14.4267 | [8 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | 9.4989 | 0.2013 | -14.6993 | SRM 3169 | Avg | -0.13569 | 0.69467 | -4.7363
Lab 3 | Feb 5, 2020 | 2 | 9 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.5925e-05; SAMPLE 49Ti40Ar: 1.96e-05; SAMPLE 90Zr: 23.1686; SAMPLE 91Zr: 5.1657; SAMPLE 92Zr: 8.0795; SAMPLE 94Zr: 8.5643; SAMPLE 96Zr: 1.4384 | 1 | -4.1351e-05 | 1.5561e-05 | 23.1685 | 5.1656 | 8.0794 | 0 | 8.5641 | 0 | 1.4383 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0483 | 23.1685 | 5.1656 | 8.0794 | 8.5641 | 1.4383 | 0.22296 | 0.34872 | 0.36964 | 0.062079 | 1.5641 | 1.6579 | 0.27843 | 1.06 | 0.17802 | 0.16794 | -1.5008 | -1.0535 | -0.99522 | -2.7793 | [10 empty] | 0 | 2.0483 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.014132 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.73% | -11.4982 | 2.1533 | -2.4676 | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.4% | 6.8667 | -4.3113 | 8.0067 | SRM 3169 | 2SD | 15.6533 | 11.4916 | 48.1497
Lab 3 | Feb 5, 2020 | 2 | 10 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.2335e-05; SAMPLE 49Ti40Ar: -5.1236e-06; SAMPLE 90Zr: 23.6433; SAMPLE 91Zr: 5.2716; SAMPLE 92Zr: 8.2452; SAMPLE 94Zr: 8.7402; SAMPLE 96Zr: 1.468 | 1 | -6.4941e-05 | -9.1627e-06 | 23.6433 | 5.2716 | 8.2451 | 0 | 8.74 | 0 | 1.4679 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0495 | 23.6433 | 5.2716 | 8.2451 | 8.74 | 1.4679 | 0.22296 | 0.34873 | 0.36966 | 0.062084 | 1.5641 | 1.6579 | 0.27845 | 1.06 | 0.17803 | 0.16795 | -1.5007 | -1.0535 | -0.99517 | -2.7793 | [10 empty] | 0 | 2.0495 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42488 | 0.43902 | -1.3883 | 0.014142 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | 1.4513 | -10.9498 | 0.53963 | [8 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | 9.2559 | 3.9726 | 39.4407 | SRM 3169 | n | 9 | 9 | 9
Lab 3 | Feb 5, 2020 | 2 | 11 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 4.9279e-05; SAMPLE 49Ti40Ar: 2.0715e-05; SAMPLE 90Zr: 23.1476; SAMPLE 91Zr: 5.1612; SAMPLE 92Zr: 8.0726; SAMPLE 94Zr: 8.5574; SAMPLE 96Zr: 1.4373 | 1 | -4.7997e-05 | 1.6676e-05 | 23.1476 | 5.1611 | 8.0725 | 0 | 8.5573 | 0 | 1.4372 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0507 | 23.1476 | 5.1611 | 8.0725 | 8.5573 | 1.4372 | 0.22297 | 0.34874 | 0.36968 | 0.062089 | 1.5641 | 1.658 | 0.27847 | 1.06 | 0.17804 | 0.16795 | -1.5007 | -1.0534 | -0.99511 | -2.7792 | [10 empty] | 0 | 2.0507 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9118 | 0.42489 | 0.43903 | -1.3883 | 0.014138 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 98.51% | -7.169 | 1.3515 | -13.1292 | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -9.093 | -3.5471 | -3.1532 | SRM 3169 | 2SE | 5.2178 | 3.8305 | 16.0499
Lab 3 | Feb 5, 2020 | 2 | 12 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.3024e-05; SAMPLE 49Ti40Ar: 1.8153e-06; SAMPLE 90Zr: 23.353; SAMPLE 91Zr: 5.2071; SAMPLE 92Zr: 8.1446; SAMPLE 94Zr: 8.6341; SAMPLE 96Zr: 1.4503 | 1 | -6.4252e-05 | -2.2237e-06 | 23.3529 | 5.207 | 8.1445 | 0 | 8.6339 | 0 | 1.4501 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0526 | 23.3529 | 5.207 | 8.1445 | 8.6339 | 1.4501 | 0.22297 | 0.34876 | 0.36971 | 0.062097 | 1.5641 | 1.6581 | 0.2785 | 1.0601 | 0.17805 | 0.16796 | -1.5007 | -1.0534 | -0.99503 | -2.7791 | [10 empty] | 0 | 2.0526 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014136 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.87% | 0.8018 | 0.45953 | 1.9173 | [8 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 102.2% | 0.21775 | 2.5404 | -15.9936 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 13 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 9.8958e-06; SAMPLE 49Ti40Ar: 1.3993e-05; SAMPLE 90Zr: 22.6327; SAMPLE 91Zr: 5.0466; SAMPLE 92Zr: 7.8937; SAMPLE 94Zr: 8.3683; SAMPLE 96Zr: 1.4056 | 1 | -8.738e-05 | 9.954e-06 | 22.6326 | 5.0465 | 7.8936 | 0 | 8.3681 | 0 | 1.4055 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0541 | 22.6326 | 5.0465 | 7.8936 | 8.3681 | 1.4055 | 0.22298 | 0.34877 | 0.36974 | 0.062103 | 1.5642 | 1.6582 | 0.27852 | 1.0601 | 0.17806 | 0.16796 | -1.5007 | -1.0533 | -0.99496 | -2.779 | [10 empty] | 0 | 2.0541 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014133 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.4% | 4.716 | 0.16623 | -7.1154 | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -2.785 | -2.1439 | -44.2701 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 14 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.3609e-05; SAMPLE 49Ti40Ar: -2.1202e-06; SAMPLE 90Zr: 23.1226; SAMPLE 91Zr: 5.1558; SAMPLE 92Zr: 8.0647; SAMPLE 94Zr: 8.5499; SAMPLE 96Zr: 1.4362 | 1 | -6.3667e-05 | -6.1593e-06 | 23.1225 | 5.1558 | 8.0646 | 0 | 8.5497 | 0 | 1.4361 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0552 | 23.1225 | 5.1558 | 8.0646 | 8.5497 | 1.4361 | 0.22298 | 0.34878 | 0.36976 | 0.062108 | 1.5642 | 1.6583 | 0.27854 | 1.0601 | 0.17807 | 0.16797 | -1.5007 | -1.0533 | -0.99491 | -2.7789 | [10 empty] | 0 | 2.0552 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014131 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.91% | 0.99418 | 7.3862 | -25.4678 | [8 empty] | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.2% | -3.7305 | -5.5205 | -0.38315 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 15 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 7.2955e-05; SAMPLE 49Ti40Ar: 2.6766e-05; SAMPLE 90Zr: 22.4474; SAMPLE 91Zr: 5.0054; SAMPLE 92Zr: 7.8295; SAMPLE 94Zr: 8.3007; SAMPLE 96Zr: 1.3944 | 1 | -2.4322e-05 | 2.2727e-05 | 22.4474 | 5.0054 | 7.8294 | 0 | 8.3006 | 0 | 1.3943 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0566 | 22.4474 | 5.0054 | 7.8294 | 8.3006 | 1.3943 | 0.22298 | 0.34879 | 0.36978 | 0.062115 | 1.5642 | 1.6583 | 0.27857 | 1.0602 | 0.17809 | 0.16798 | -1.5007 | -1.0533 | -0.99485 | -2.7788 | [10 empty] | 0 | 2.0566 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014136 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.48% | 5.4921 | -0.36523 | 2.8337 | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -12.6969 | 13.2978 | 13.0571 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 16 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 4.4515e-05; SAMPLE 49Ti40Ar: 1.6246e-05; SAMPLE 90Zr: 22.9315; SAMPLE 91Zr: 5.1134; SAMPLE 92Zr: 7.9985; SAMPLE 94Zr: 8.4802; SAMPLE 96Zr: 1.4246 | 1 | -5.2761e-05 | 1.2207e-05 | 22.9315 | 5.1133 | 7.9984 | 0 | 8.48 | 0 | 1.4245 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0579 | 22.9315 | 5.1133 | 7.9984 | 8.48 | 1.4245 | 0.22298 | 0.3488 | 0.3698 | 0.062121 | 1.5642 | 1.6584 | 0.27859 | 1.0602 | 0.1781 | 0.16799 | -1.5007 | -1.0533 | -0.9948 | -2.7787 | [10 empty] | 0 | 2.0579 | 0.21796 | 0.33335 | 0.3381 | 0.054384 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3882 | 0.014141 | -1.8131 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | -6.4906 | -7.3738 | 54.2709 | [8 empty] | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.4% | 1.2449 | 1.7627 | -24.6316 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 17 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.2396e-05; SAMPLE 49Ti40Ar: -2.8548e-05; SAMPLE 90Zr: 22.2055; SAMPLE 91Zr: 4.9515; SAMPLE 92Zr: 7.7455; SAMPLE 94Zr: 8.2122; SAMPLE 96Zr: 1.3796 | 1 | -4.488e-05 | -3.2587e-05 | 22.2054 | 4.9515 | 7.7454 | 0 | 8.212 | 0 | 1.3795 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0594 | 22.2054 | 4.9515 | 7.7454 | 8.212 | 1.3795 | 0.22299 | 0.34881 | 0.36982 | 0.062125 | 1.5643 | 1.6585 | 0.2786 | 1.0602 | 0.17811 | 0.16799 | -1.5006 | -1.0532 | -0.99473 | -2.7786 | [10 empty] | 0 | 2.0594 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014136 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.54% | -3.283 | 2.2247 | -3.8054
Lab 3 | Feb 5, 2020 | 2 | 18 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.0977e-05; SAMPLE 49Ti40Ar: 1.7803e-05; SAMPLE 90Zr: 22.6013; SAMPLE 91Zr: 5.0399; SAMPLE 92Zr: 7.8838; SAMPLE 94Zr: 8.3591; SAMPLE 96Zr: 1.4043 | 1 | -4.6299e-05 | 1.3764e-05 | 22.6012 | 5.0399 | 7.8837 | 0 | 8.3589 | 0 | 1.4042 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0606 | 22.6012 | 5.0399 | 7.8837 | 8.3589 | 1.4042 | 0.22299 | 0.34882 | 0.36984 | 0.062129 | 1.5643 | 1.6586 | 0.27861 | 1.0603 | 0.17811 | 0.16799 | -1.5006 | -1.0532 | -0.99468 | -2.7785 | [10 empty] | 0 | 2.0606 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9118 | 0.42489 | 0.43902 | -1.3883 | 0.014136 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.34% | 6.8507 | 4.419 | -33.9777 | [8 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.4% | -1.894 | 13.443 | -2.5636 | SRM 3169 | Avg | -3.1857 | 1.839 | -7.5664
Lab 3 | Feb 5, 2020 | 2 | 19 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: -3.2368e-06; SAMPLE 49Ti40Ar: -4.9402e-06; SAMPLE 90Zr: 21.9783; SAMPLE 91Zr: 4.9011; SAMPLE 92Zr: 7.6668; SAMPLE 94Zr: 8.1292; SAMPLE 96Zr: 1.3657 | 1 | -0.00010051 | -8.9792e-06 | 21.9783 | 4.9011 | 7.6667 | 0 | 8.1291 | 0 | 1.3656 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0622 | 21.9783 | 4.9011 | 7.6667 | 8.1291 | 1.3656 | 0.223 | 0.34883 | 0.36987 | 0.062135 | 1.5643 | 1.6586 | 0.27864 | 1.0603 | 0.17812 | 0.16799 | -1.5006 | -1.0532 | -0.99461 | -2.7784 | [10 empty] | 0 | 2.0622 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014137 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.31% | 4.1424 | 0.71719 | -2.7621 | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.5% | 11.0165 | 6.3856 | -17.258 | SRM 3169 | 2SD | 14.5208 | 18.4814 | 44.6152
Lab 3 | Feb 5, 2020 | 2 | 20 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.6821e-05; SAMPLE 49Ti40Ar: 3.4976e-06; SAMPLE 90Zr: 22.5705; SAMPLE 91Zr: 5.0332; SAMPLE 92Zr: 7.8735; SAMPLE 94Zr: 8.3487; SAMPLE 96Zr: 1.4026 | 1 | -8.0455e-05 | -5.4144e-07 | 22.5704 | 5.0331 | 7.8734 | 0 | 8.3485 | 0 | 1.4025 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0634 | 22.5704 | 5.0331 | 7.8734 | 8.3485 | 1.4025 | 0.223 | 0.34884 | 0.36989 | 0.062141 | 1.5643 | 1.6587 | 0.27866 | 1.0603 | 0.17814 | 0.168 | -1.5006 | -1.0531 | -0.99456 | -2.7784 | [10 empty] | 0 | 2.0634 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.93% | -4.5354 | -7.4393 | -10.0301 | [8 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.6% | -9.7713 | 8.7516 | -6.3123 | SRM 3169 | n | 16 | 16 | 16
Lab 3 | Feb 5, 2020 | 2 | 21 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 4.6711e-05; SAMPLE 49Ti40Ar: 1.562e-05; SAMPLE 90Zr: 21.832; SAMPLE 91Zr: 4.8685; SAMPLE 92Zr: 7.6162; SAMPLE 94Zr: 8.0759; SAMPLE 96Zr: 1.3569 | 1 | -5.0565e-05 | 1.1581e-05 | 21.8319 | 4.8685 | 7.6161 | 0 | 8.0757 | 0 | 1.3568 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0645 | 21.8319 | 4.8685 | 7.6161 | 8.0757 | 1.3568 | 0.223 | 0.34885 | 0.3699 | 0.062146 | 1.5644 | 1.6588 | 0.27868 | 1.0604 | 0.17815 | 0.16801 | -1.5006 | -1.0531 | -0.99451 | -2.7783 | [10 empty] | 0 | 2.0645 | 0.21796 | 0.33336 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.014132 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 96.72% | -5.0257 | 2.0391 | 6.3257 | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.2% | -18.6456 | -6.5614 | -22.9547 | SRM 3169 | 2SE | 3.6302 | 4.6204 | 11.1538
Lab 3 | Feb 5, 2020 | 2 | 22 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.7086e-05; SAMPLE 49Ti40Ar: -9.4588e-06; SAMPLE 90Zr: 22.5735; SAMPLE 91Zr: 5.034; SAMPLE 92Zr: 7.8751; SAMPLE 94Zr: 8.3506; SAMPLE 96Zr: 1.4031 | 1 | -6.019e-05 | -1.3498e-05 | 22.5735 | 5.034 | 7.875 | 0 | 8.3504 | 0 | 1.403 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0656 | 22.5735 | 5.034 | 7.875 | 8.3504 | 1.403 | 0.223 | 0.34886 | 0.36992 | 0.062151 | 1.5644 | 1.6588 | 0.2787 | 1.0604 | 0.17816 | 0.16801 | -1.5006 | -1.0531 | -0.99446 | -2.7782 | [10 empty] | 0 | 2.0656 | 0.21796 | 0.33336 | 0.3381 | 0.054383 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.014129 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.6% | 3.8094 | 13.2507 | 28.9739 | [8 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.1% | -5.5718 | 6.7614 | -5.7334 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 23 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.4392e-05; SAMPLE 49Ti40Ar: 1.2386e-05; SAMPLE 90Zr: 21.8788; SAMPLE 91Zr: 4.8792; SAMPLE 92Zr: 7.6329; SAMPLE 94Zr: 8.0941; SAMPLE 96Zr: 1.36 | 1 | -6.2884e-05 | 8.3466e-06 | 21.8788 | 4.8791 | 7.6328 | 0 | 8.0939 | 0 | 1.3599 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0669 | 21.8788 | 4.8791 | 7.6328 | 8.0939 | 1.3599 | 0.22301 | 0.34887 | 0.36994 | 0.062155 | 1.5644 | 1.6589 | 0.27871 | 1.0604 | 0.17816 | 0.16801 | -1.5006 | -1.0531 | -0.9944 | -2.7781 | [10 empty] | 0 | 2.0669 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014129 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.5% | 7.0327 | 7.7044 | -10.4999 | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | -3.6032 | 16.4609 | -9.2216 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 24 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.2464e-05; SAMPLE 49Ti40Ar: 2.1044e-05; SAMPLE 90Zr: 22.3069; SAMPLE 91Zr: 4.9746; SAMPLE 92Zr: 7.7824; SAMPLE 94Zr: 8.2529; SAMPLE 96Zr: 1.3867 | 1 | -8.4812e-05 | 1.7005e-05 | 22.3068 | 4.9746 | 7.7823 | 0 | 8.2527 | 0 | 1.3866 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0682 | 22.3068 | 4.9746 | 7.7823 | 8.2527 | 1.3866 | 0.22301 | 0.34887 | 0.36996 | 0.06216 | 1.5644 | 1.659 | 0.27873 | 1.0605 | 0.17817 | 0.16802 | -1.5005 | -1.053 | -0.99435 | -2.778 | [10 empty] | 0 | 2.0682 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014144 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.87% | -1.3111 | -9.0457 | -36.6994 | [8 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 100.8% | 0.73283 | 3.3364 | -30.8964 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 25 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 2.0174e-05; SAMPLE 49Ti40Ar: 1.0777e-05; SAMPLE 90Zr: 21.5405; SAMPLE 91Zr: 4.8038; SAMPLE 92Zr: 7.5152; SAMPLE 94Zr: 7.9698; SAMPLE 96Zr: 1.3392 | 1 | -7.7102e-05 | 6.7382e-06 | 21.5404 | 4.8037 | 7.5151 | 0 | 7.9696 | 0 | 1.3391 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0694 | 21.5404 | 4.8037 | 7.5151 | 7.9696 | 1.3391 | 0.22301 | 0.34888 | 0.36998 | 0.062166 | 1.5644 | 1.659 | 0.27876 | 1.0605 | 0.17818 | 0.16802 | -1.5005 | -1.053 | -0.9943 | -2.7779 | [10 empty] | 0 | 2.0694 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014139 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.02% | -1.9139 | 3.3235 | 0.32492 | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 100.7% | -5.4567 | -0.41162 | -10.5982 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 26 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.6711e-05; SAMPLE 49Ti40Ar: 2.6146e-05; SAMPLE 90Zr: 22.0968; SAMPLE 91Zr: 4.9279; SAMPLE 92Zr: 7.7095; SAMPLE 94Zr: 8.1761; SAMPLE 96Zr: 1.3739 | 1 | -6.0565e-05 | 2.2107e-05 | 22.0968 | 4.9279 | 7.7094 | 0 | 8.1759 | 0 | 1.3738 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0707 | 22.0968 | 4.9279 | 7.7094 | 8.1759 | 1.3738 | 0.22301 | 0.34889 | 0.37 | 0.062172 | 1.5644 | 1.6591 | 0.27878 | 1.0605 | 0.1782 | 0.16803 | -1.5005 | -1.053 | -0.99424 | -2.7778 | [10 empty] | 0 | 2.0707 | 0.21796 | 0.33335 | 0.3381 | 0.054383 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3882 | 0.014141 | -1.8131 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.32% | -3.6179 | -1.2147 | 39.2921 | [8 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.3% | -13.2272 | 4.6327 | -21.8299 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 27 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 2.5591e-05; SAMPLE 49Ti40Ar: 1.0753e-05; SAMPLE 90Zr: 21.4881; SAMPLE 91Zr: 4.7923; SAMPLE 92Zr: 7.4974; SAMPLE 94Zr: 7.9513; SAMPLE 96Zr: 1.3361 | 1 | -7.1686e-05 | 6.7136e-06 | 21.488 | 4.7922 | 7.4973 | 0 | 7.9511 | 0 | 1.336 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0719 | 21.488 | 4.7922 | 7.4973 | 7.9511 | 1.336 | 0.22302 | 0.3489 | 0.37002 | 0.062176 | 1.5645 | 1.6592 | 0.27879 | 1.0605 | 0.1782 | 0.16803 | -1.5005 | -1.053 | -0.99419 | -2.7778 | [10 empty] | 0 | 2.0719 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014138 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.04% | 3.5086 | 0.41351 | -6.9756 | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 99.98% | 6.0863 | -0.85961 | 46.419 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 28 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 2.1476e-05; SAMPLE 49Ti40Ar: 4.348e-06; SAMPLE 90Zr: 22.188; SAMPLE 91Zr: 4.9484; SAMPLE 92Zr: 7.7417; SAMPLE 94Zr: 8.2106; SAMPLE 96Zr: 1.3797 | 1 | -7.58e-05 | 3.0897e-07 | 22.1879 | 4.9484 | 7.7416 | 0 | 8.2104 | 0 | 1.3796 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0729 | 22.1879 | 4.9484 | 7.7416 | 8.2104 | 1.3796 | 0.22302 | 0.34891 | 0.37004 | 0.062179 | 1.5645 | 1.6592 | 0.2788 | 1.0606 | 0.17821 | 0.16803 | -1.5005 | -1.0529 | -0.99414 | -2.7777 | [10 empty] | 0 | 2.0729 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014136 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 101.3% | 3.7659 | 4.6149 | -13.5627 | [8 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100% | -0.89466 | -11.422 | -34.0219 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 29 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 2.8143e-05; SAMPLE 49Ti40Ar: 2.6611e-05; SAMPLE 90Zr: 21.4882; SAMPLE 91Zr: 4.7924; SAMPLE 92Zr: 7.4978; SAMPLE 94Zr: 7.9521; SAMPLE 96Zr: 1.3364 | 1 | -6.9133e-05 | 2.2572e-05 | 21.4882 | 4.7923 | 7.4977 | 0 | 7.952 | 0 | 1.3363 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0743 | 21.4882 | 4.7923 | 7.4977 | 7.952 | 1.3363 | 0.22302 | 0.34892 | 0.37006 | 0.062186 | 1.5645 | 1.6593 | 0.27883 | 1.0606 | 0.17822 | 0.16804 | -1.5005 | -1.0529 | -0.99408 | -2.7776 | [10 empty] | 0 | 2.0743 | 0.21796 | 0.33335 | 0.3381 | 0.054383 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.87% | -7.4011 | -1.873 | 25.0781 | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.1% | -6.1346 | -7.9869 | -14.8614 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 30 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.0871e-06; SAMPLE 49Ti40Ar: 2.0529e-05; SAMPLE 90Zr: 21.7259; SAMPLE 91Zr: 4.8455; SAMPLE 92Zr: 7.5809; SAMPLE 94Zr: 8.0406; SAMPLE 96Zr: 1.3512 | 1 | -9.2189e-05 | 1.6489e-05 | 21.7259 | 4.8454 | 7.5808 | 0 | 8.0404 | 0 | 1.3511 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0756 | 21.7259 | 4.8454 | 7.5808 | 8.0404 | 1.3511 | 0.22303 | 0.34893 | 0.37008 | 0.06219 | 1.5645 | 1.6594 | 0.27884 | 1.0606 | 0.17823 | 0.16804 | -1.5005 | -1.0529 | -0.99403 | -2.7776 | [10 empty] | 0 | 2.0756 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.04% | -4.2096 | -2.2919 | -11.1291 | [8 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 99.58% | 0.62147 | -10.3504 | -20.2272 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 31 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.553e-05; SAMPLE 49Ti40Ar: 1.5916e-05; SAMPLE 90Zr: 21.2929; SAMPLE 91Zr: 4.749; SAMPLE 92Zr: 7.43; SAMPLE 94Zr: 7.8807; SAMPLE 96Zr: 1.3244 | 1 | -4.1746e-05 | 1.1877e-05 | 21.2928 | 4.749 | 7.4299 | 0 | 7.8805 | 0 | 1.3243 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0767 | 21.2928 | 4.749 | 7.4299 | 7.8805 | 1.3243 | 0.22303 | 0.34894 | 0.3701 | 0.062195 | 1.5645 | 1.6594 | 0.27886 | 1.0606 | 0.17824 | 0.16805 | -1.5004 | -1.0529 | -0.99398 | -2.7775 | [10 empty] | 0 | 2.0767 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014137 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 98.09% | 4.7477 | 3.4747 | 14.8564 | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.1% | -7.3496 | -9.4811 | 16.3832 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 32 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.1658e-05; SAMPLE 49Ti40Ar: 1.1917e-05; SAMPLE 90Zr: 21.687; SAMPLE 91Zr: 4.8369; SAMPLE 92Zr: 7.5678; SAMPLE 94Zr: 8.027; SAMPLE 96Zr: 1.349 | 1 | -6.5618e-05 | 7.8781e-06 | 21.687 | 4.8369 | 7.5677 | 0 | 8.0268 | 0 | 1.3489 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0779 | 21.687 | 4.8369 | 7.5677 | 8.0268 | 1.3489 | 0.22303 | 0.34895 | 0.37012 | 0.0622 | 1.5646 | 1.6595 | 0.27888 | 1.0607 | 0.17825 | 0.16805 | -1.5004 | -1.0528 | -0.99393 | -2.7774 | [10 empty] | 0 | 2.0779 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.4% | 6.4009 | -0.93341 | 10.8244 | [8 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.2% | 2.2467 | 0.86291 | -21.6014 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 33 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.6361e-05; SAMPLE 49Ti40Ar: 2.4241e-05; SAMPLE 90Zr: 21.0884; SAMPLE 91Zr: 4.7034; SAMPLE 92Zr: 7.3591; SAMPLE 94Zr: 7.8058; SAMPLE 96Zr: 1.3119 | 1 | -8.0915e-05 | 2.0202e-05 | 21.0883 | 4.7034 | 7.359 | 0 | 7.8057 | 0 | 1.3118 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0792 | 21.0883 | 4.7034 | 7.359 | 7.8057 | 1.3118 | 0.22303 | 0.34896 | 0.37014 | 0.062207 | 1.5646 | 1.6596 | 0.27891 | 1.0607 | 0.17826 | 0.16806 | -1.5004 | -1.0528 | -0.99387 | -2.7773 | [10 empty] | 0 | 2.0792 | 0.21796 | 0.33335 | 0.3381 | 0.054383 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5235 | -1.0985 | -2.9117 | 0.4249 | 0.43904 | -1.3882 | 0.014134 | -1.8131 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.69% | -9.6749 | 4.3978 | 27.8727 | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.2% | 0.87435 | 15.8633 | 34.215 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 34 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.0714e-05; SAMPLE 49Ti40Ar: 9.2472e-06; SAMPLE 90Zr: 21.4852; SAMPLE 91Zr: 4.7921; SAMPLE 92Zr: 7.4978; SAMPLE 94Zr: 7.9533; SAMPLE 96Zr: 1.3367 | 1 | -6.6562e-05 | 5.2081e-06 | 21.4851 | 4.792 | 7.4977 | 0 | 7.9531 | 0 | 1.3366 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0807 | 21.4851 | 4.792 | 7.4977 | 7.9531 | 1.3366 | 0.22304 | 0.34897 | 0.37017 | 0.062211 | 1.5646 | 1.6596 | 0.27892 | 1.0607 | 0.17827 | 0.16806 | -1.5004 | -1.0528 | -0.9938 | -2.7772 | [10 empty] | 0 | 2.0807 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16314 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.014138 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.78% | -4.7843 | -2.645 | -6.0441
Lab 3 | Feb 5, 2020 | 2 | 35 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 4.0125e-05; SAMPLE 49Ti40Ar: -3.3331e-06; SAMPLE 90Zr: 21.0027; SAMPLE 91Zr: 4.6845; SAMPLE 92Zr: 7.3296; SAMPLE 94Zr: 7.775; SAMPLE 96Zr: 1.3068 | 1 | -5.7152e-05 | -7.3721e-06 | 21.0026 | 4.6845 | 7.3295 | 0 | 7.7748 | 0 | 1.3067 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0818 | 21.0026 | 4.6845 | 7.3295 | 7.7748 | 1.3067 | 0.22304 | 0.34898 | 0.37018 | 0.062217 | 1.5646 | 1.6597 | 0.27895 | 1.0608 | 0.17828 | 0.16807 | -1.5004 | -1.0527 | -0.99376 | -2.7771 | [10 empty] | 0 | 2.0818 | 0.21796 | 0.33336 | 0.3381 | 0.054384 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43903 | -1.3882 | 0.014132 | -1.8131 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.99% | 2.0996 | 2.1737 | 34.0865 | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 99.94% | -1.924 | 3.1764 | 15.1135 | SRM 3169 | Avg | 2.5994 | -0.82938 | 3.5411
Lab 3 | Feb 5, 2020 | 2 | 36 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 4.5419e-05; SAMPLE 49Ti40Ar: 1.8564e-05; SAMPLE 90Zr: 21.3802; SAMPLE 91Zr: 4.7688; SAMPLE 92Zr: 7.4617; SAMPLE 94Zr: 7.9153; SAMPLE 96Zr: 1.3304 | 1 | -5.1857e-05 | 1.4525e-05 | 21.3802 | 4.7688 | 7.4616 | 0 | 7.9151 | 0 | 1.3303 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0832 | 21.3802 | 4.7688 | 7.4616 | 7.9151 | 1.3303 | 0.22305 | 0.34899 | 0.37021 | 0.062221 | 1.5647 | 1.6598 | 0.27896 | 1.0608 | 0.17829 | 0.16807 | -1.5004 | -1.0527 | -0.99369 | -2.7771 | [10 empty] | 0 | 2.0832 | 0.21796 | 0.33336 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.01413 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | 2.6893 | 7.1043 | 10.1579 | [8 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 98.98% | 15.2714 | -0.65995 | 17.4744 | SRM 3169 | 2SD | 16.7823 | 11.8354 | 32.5665
Lab 3 | Feb 5, 2020 | 2 | 37 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: -3.767e-06; SAMPLE 49Ti40Ar: 1.3352e-05; SAMPLE 90Zr: 20.6284; SAMPLE 91Zr: 4.6011; SAMPLE 92Zr: 7.1994; SAMPLE 94Zr: 7.6373; SAMPLE 96Zr: 1.2837 | 1 | -0.00010104 | 9.3132e-06 | 20.6284 | 4.6011 | 7.1993 | 0 | 7.6372 | 0 | 1.2836 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0844 | 20.6284 | 4.6011 | 7.1993 | 7.6372 | 1.2836 | 0.22305 | 0.349 | 0.37023 | 0.062226 | 1.5647 | 1.6599 | 0.27898 | 1.0608 | 0.1783 | 0.16808 | -1.5004 | -1.0527 | -0.99364 | -2.777 | [10 empty] | 0 | 2.0844 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014141 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.01% | -8.93 | -7.3544 | 8.4222 | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 101.5% | 0.27042 | 2.5772 | -2.1929 | SRM 3169 | n | 9 | 9 | 9
Lab 3 | Feb 5, 2020 | 2 | 38 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.1781e-05; SAMPLE 49Ti40Ar: 1.8427e-05; SAMPLE 90Zr: 21.1477; SAMPLE 91Zr: 4.717; SAMPLE 92Zr: 7.3809; SAMPLE 94Zr: 7.83; SAMPLE 96Zr: 1.3161 | 1 | -8.5495e-05 | 1.4388e-05 | 21.1477 | 4.717 | 7.3808 | 0 | 7.8299 | 0 | 1.316 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0857 | 21.1477 | 4.717 | 7.3808 | 7.8299 | 1.316 | 0.22305 | 0.34901 | 0.37025 | 0.06223 | 1.5647 | 1.6599 | 0.279 | 1.0608 | 0.17831 | 0.16808 | -1.5004 | -1.0527 | -0.99359 | -2.7769 | [10 empty] | 0 | 2.0857 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.4249 | 0.43903 | -1.3883 | 0.014137 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.42% | -4.4766 | -4.2412 | -21.3077 | [8 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 99.75% | -6.1204 | 3.963 | 18.7942 | SRM 3169 | 2SE | 5.5941 | 3.9451 | 10.8555
Lab 3 | Feb 5, 2020 | 2 | 39 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.6041e-05; SAMPLE 49Ti40Ar: -4.1135e-06; SAMPLE 90Zr: 20.6338; SAMPLE 91Zr: 4.6025; SAMPLE 92Zr: 7.2017; SAMPLE 94Zr: 7.6402; SAMPLE 96Zr: 1.2843 | 1 | -4.1235e-05 | -8.1525e-06 | 20.6337 | 4.6025 | 7.2016 | 0 | 7.64 | 0 | 1.2842 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0871 | 20.6337 | 4.6025 | 7.2016 | 7.64 | 1.2842 | 0.22306 | 0.34902 | 0.37027 | 0.062236 | 1.5647 | 1.66 | 0.27902 | 1.0609 | 0.17832 | 0.16808 | -1.5003 | -1.0526 | -0.99353 | -2.7768 | [10 empty] | 0 | 2.0871 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014136 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.54% | -0.85292 | -0.21458 | -7.5691 | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 98.93% | 8.4657 | -1.6651 | -4.4758 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 40 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 9.3297e-07; SAMPLE 49Ti40Ar: -1.4422e-05; SAMPLE 90Zr: 21.1611; SAMPLE 91Zr: 4.7202; SAMPLE 92Zr: 7.3859; SAMPLE 94Zr: 7.8358; SAMPLE 96Zr: 1.3172 | 1 | -9.6343e-05 | -1.8461e-05 | 21.161 | 4.7202 | 7.3858 | 0 | 7.8357 | 0 | 1.3171 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0882 | 21.161 | 4.7202 | 7.3858 | 7.8357 | 1.3171 | 0.22306 | 0.34903 | 0.37029 | 0.062242 | 1.5647 | 1.66 | 0.27904 | 1.0609 | 0.17833 | 0.16809 | -1.5003 | -1.0526 | -0.99348 | -2.7767 | [10 empty] | 0 | 2.0882 | 0.21796 | 0.33335 | 0.3381 | 0.054383 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014135 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.9% | 6.2576 | 4.3813 | 26.6193 | [8 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 101.5% | 1.8244 | -6.2862 | 22.0284 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 41 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.5129e-05; SAMPLE 49Ti40Ar: 1.0537e-06; SAMPLE 90Zr: 20.3357; SAMPLE 91Zr: 4.5361; SAMPLE 92Zr: 7.098; SAMPLE 94Zr: 7.5306; SAMPLE 96Zr: 1.2659 | 1 | -8.2147e-05 | -2.9853e-06 | 20.3356 | 4.5361 | 7.0979 | 0 | 7.5305 | 0 | 1.2658 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0896 | 20.3356 | 4.5361 | 7.0979 | 7.5305 | 1.2658 | 0.22306 | 0.34904 | 0.37031 | 0.062245 | 1.5648 | 1.6601 | 0.27905 | 1.0609 | 0.17833 | 0.16809 | -1.5003 | -1.0526 | -0.99342 | -2.7767 | [10 empty] | 0 | 2.0896 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014144 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 96.94% | -3.8649 | -5.3327 | -27.8798 | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 99.89% | -4.4182 | -5.9165 | -24.9231 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 42 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.2742e-06; SAMPLE 49Ti40Ar: -2.3988e-05; SAMPLE 90Zr: 20.7942; SAMPLE 91Zr: 4.6385; SAMPLE 92Zr: 7.2583; SAMPLE 94Zr: 7.7009; SAMPLE 96Zr: 1.2946 | 1 | -9.6002e-05 | -2.8027e-05 | 20.7942 | 4.6384 | 7.2582 | 0 | 7.7007 | 0 | 1.2945 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0908 | 20.7942 | 4.6384 | 7.2582 | 7.7007 | 1.2945 | 0.22306 | 0.34905 | 0.37033 | 0.062251 | 1.5648 | 1.6602 | 0.27907 | 1.061 | 0.17835 | 0.1681 | -1.5003 | -1.0525 | -0.99336 | -2.7766 | [10 empty] | 0 | 2.0908 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42489 | 0.43903 | -1.3883 | 0.014142 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.22% | -1.6491 | -8.2986 | -27.9114 | [8 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 98.72% | 15.1341 | 7.9647 | -14.0923 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 43 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: -7.4911e-06; SAMPLE 49Ti40Ar: 2.0352e-05; SAMPLE 90Zr: 20.2684; SAMPLE 91Zr: 4.5212; SAMPLE 92Zr: 7.0749; SAMPLE 94Zr: 7.5064; SAMPLE 96Zr: 1.2619 | 1 | -0.00010477 | 1.6313e-05 | 20.2684 | 4.5211 | 7.0748 | 0 | 7.5062 | 0 | 1.2618 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0917 | 20.2684 | 4.5211 | 7.0748 | 7.5062 | 1.2618 | 0.22306 | 0.34905 | 0.37034 | 0.062255 | 1.5648 | 1.6603 | 0.27909 | 1.061 | 0.17835 | 0.1681 | -1.5003 | -1.0525 | -0.99333 | -2.7765 | [10 empty] | 0 | 2.0917 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014143 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.56% | -6.4549 | -5.649 | -9.9788 | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 101.8% | -5.1086 | -10.6181 | 4.1433 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 44 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 5.2644e-06; SAMPLE 49Ti40Ar: 3.1203e-05; SAMPLE 90Zr: 20.7547; SAMPLE 91Zr: 4.6297; SAMPLE 92Zr: 7.2449; SAMPLE 94Zr: 7.687; SAMPLE 96Zr: 1.2923 | 1 | -9.2012e-05 | 2.7164e-05 | 20.7546 | 4.6297 | 7.2448 | 0 | 7.6868 | 0 | 1.2922 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0931 | 20.7546 | 4.6297 | 7.2448 | 7.6868 | 1.2922 | 0.22307 | 0.34907 | 0.37036 | 0.062262 | 1.5648 | 1.6603 | 0.27912 | 1.061 | 0.17837 | 0.16811 | -1.5003 | -1.0525 | -0.99327 | -2.7764 | [10 empty] | 0 | 2.0931 | 0.21796 | 0.33336 | 0.3381 | 0.054383 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5234 | -1.0985 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014132 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.2% | 2.8763 | 11.291 | 39.1196
Lab 3 | Feb 5, 2020 | 2 | 45 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 2.5179e-05; SAMPLE 49Ti40Ar: 1.5706e-05; SAMPLE 90Zr: 20.2801; SAMPLE 91Zr: 4.5239; SAMPLE 92Zr: 7.0794; SAMPLE 94Zr: 7.5116; SAMPLE 96Zr: 1.2628 | 1 | -7.2097e-05 | 1.1667e-05 | 20.28 | 4.5238 | 7.0792 | 0 | 7.5114 | 0 | 1.2627 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0942 | 20.28 | 4.5238 | 7.0792 | 7.5114 | 1.2627 | 0.22307 | 0.34907 | 0.37038 | 0.062265 | 1.5649 | 1.6604 | 0.27913 | 1.061 | 0.17837 | 0.16811 | -1.5003 | -1.0525 | -0.99322 | -2.7764 | [10 empty] | 0 | 2.0942 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42491 | 0.43905 | -1.3883 | 0.014141 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 98.01% | -9.5685 | -2.2838 | -10.9981 | STD | Lab 3 | Feb 4, 2020 | 1 | 99.38% | -0.65514 | 6.726 | 14.227 | SRM 3169 | Avg | 0.023857 | 0.11918 | 0.54157
Lab 3 | Feb 5, 2020 | 2 | 46 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.4084e-05; SAMPLE 49Ti40Ar: 2.7144e-05; SAMPLE 90Zr: 20.6309; SAMPLE 91Zr: 4.6022; SAMPLE 92Zr: 7.202; SAMPLE 94Zr: 7.642; SAMPLE 96Zr: 1.2848 | 1 | -8.3192e-05 | 2.3105e-05 | 20.6309 | 4.6022 | 7.2019 | 0 | 7.6418 | 0 | 1.2847 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0956 | 20.6309 | 4.6022 | 7.2019 | 7.6418 | 1.2847 | 0.22307 | 0.34908 | 0.37041 | 0.06227 | 1.5649 | 1.6605 | 0.27914 | 1.0611 | 0.17838 | 0.16811 | -1.5003 | -1.0524 | -0.99315 | -2.7763 | [10 empty] | 0 | 2.0956 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014145 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | -6.7233 | -4.4336 | -50.0868 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100% | -7.2394 | -6.366 | -10.732 | SRM 3169 | 2SD | 9.1745 | 12.4834 | 43.7845
Lab 3 | Feb 5, 2020 | 2 | 47 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.4417e-05; SAMPLE 49Ti40Ar: -8.1193e-06; SAMPLE 90Zr: 19.9711; SAMPLE 91Zr: 4.4551; SAMPLE 92Zr: 6.9718; SAMPLE 94Zr: 7.3979; SAMPLE 96Zr: 1.2438 | 1 | -8.2859e-05 | -1.2158e-05 | 19.9711 | 4.4551 | 6.9717 | 0 | 7.3977 | 0 | 1.2437 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0966 | 19.9711 | 4.4551 | 6.9717 | 7.3977 | 1.2437 | 0.22308 | 0.34909 | 0.37042 | 0.062275 | 1.5649 | 1.6605 | 0.27916 | 1.0611 | 0.17839 | 0.16812 | -1.5002 | -1.0524 | -0.99311 | -2.7762 | [10 empty] | 0 | 2.0966 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0143 | 0.16314 | 0.16085 | -1.5235 | -1.0986 | -2.9117 | 0.42489 | 0.43904 | -1.3883 | 0.01415 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.42% | -2.093 | -2.9486 | -3.1949 | STD | Lab 3 | Feb 4, 2020 | 1 | 100.2% | 12.3514 | 6.7904 | -1.7828 | SRM 3169 | n | 52 | 52 | 52
Lab 3 | Feb 5, 2020 | 2 | 48 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: -2.3901e-06; SAMPLE 49Ti40Ar: -2.0651e-05; SAMPLE 90Zr: 20.3676; SAMPLE 91Zr: 4.5436; SAMPLE 92Zr: 7.1104; SAMPLE 94Zr: 7.5451; SAMPLE 96Zr: 1.2686 | 1 | -9.9666e-05 | -2.469e-05 | 20.3675 | 4.5436 | 7.1103 | 0 | 7.545 | 0 | 1.2685 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0977 | 20.3675 | 4.5436 | 7.1103 | 7.545 | 1.2685 | 0.22308 | 0.3491 | 0.37044 | 0.062281 | 1.5649 | 1.6606 | 0.27919 | 1.0611 | 0.17841 | 0.16813 | -1.5002 | -1.0524 | -0.99306 | -2.7761 | [10 empty] | 0 | 2.0977 | 0.21796 | 0.33335 | 0.3381 | 0.054383 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5235 | -1.0986 | -2.9117 | 0.42489 | 0.43904 | -1.3882 | 0.014149 | -1.8131 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.59% | 0.51886 | -7.3935 | 54.0279 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.98% | 0.30494 | -6.7606 | 0.091153 | SRM 3169 | 2SE | 1.2723 | 1.7311 | 6.0718
Lab 3 | Feb 5, 2020 | 2 | 49 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.6803e-05; SAMPLE 49Ti40Ar: 9.8459e-08; SAMPLE 90Zr: 19.8251; SAMPLE 91Zr: 4.4226; SAMPLE 92Zr: 6.9213; SAMPLE 94Zr: 7.3446; SAMPLE 96Zr: 1.2349 | 1 | -8.0474e-05 | -3.9406e-06 | 19.825 | 4.4226 | 6.9211 | 0 | 7.3444 | 0 | 1.2348 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0991 | 19.825 | 4.4226 | 6.9211 | 7.3444 | 1.2348 | 0.22308 | 0.34911 | 0.37046 | 0.062284 | 1.5649 | 1.6607 | 0.2792 | 1.0612 | 0.17841 | 0.16813 | -1.5002 | -1.0524 | -0.993 | -2.776 | [10 empty] | 0 | 2.0991 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014143 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.56% | -5.0593 | 0.11489 | -15.1197 | STD | Lab 3 | Feb 4, 2020 | 1 | 99.4% | -9.2645 | -1.396 | 8.6175 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 50 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.6945e-05; SAMPLE 49Ti40Ar: -9.768e-06; SAMPLE 90Zr: 20.2729; SAMPLE 91Zr: 4.5226; SAMPLE 92Zr: 7.0778; SAMPLE 94Zr: 7.5109; SAMPLE 96Zr: 1.2629 | 1 | -6.0331e-05 | -1.3807e-05 | 20.2729 | 4.5226 | 7.0777 | 0 | 7.5107 | 0 | 1.2628 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1003 | 20.2729 | 4.5226 | 7.0777 | 7.5107 | 1.2628 | 0.22309 | 0.34912 | 0.37048 | 0.062288 | 1.565 | 1.6607 | 0.27921 | 1.0612 | 0.17841 | 0.16813 | -1.5002 | -1.0523 | -0.99295 | -2.776 | [10 empty] | 0 | 2.1003 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014138 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.68% | 4.7135 | 6.4741 | -37.7691 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 101.3% | 6.4608 | 4.0978 | -2.3109 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 51 | iRM4 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 1.3688e-05; SAMPLE 49Ti40Ar: -6.369e-06; SAMPLE 90Zr: 19.7008; SAMPLE 91Zr: 4.395; SAMPLE 92Zr: 6.8783; SAMPLE 94Zr: 7.2994; SAMPLE 96Zr: 1.2274 | 1 | -8.3588e-05 | -1.0408e-05 | 19.7008 | 4.395 | 6.8782 | 0 | 7.2992 | 0 | 1.2273 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1016 | 19.7008 | 4.395 | 6.8782 | 7.2992 | 1.2273 | 0.22309 | 0.34913 | 0.3705 | 0.062295 | 1.565 | 1.6608 | 0.27924 | 1.0612 | 0.17843 | 0.16814 | -1.5002 | -1.0523 | -0.99289 | -2.7759 | [10 empty] | 0 | 2.1016 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014141 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Feb 5, 2020 | 2 | 97.09% | -4.1282 | -1.4992 | 16.3688 | STD | Lab 3 | Feb 4, 2020 | 1 | 98.87% | -5.1388 | 1.6008 | -17.5875 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 52 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 2.7456e-05; SAMPLE 49Ti40Ar: 3.1757e-05; SAMPLE 90Zr: 20.3095; SAMPLE 91Zr: 4.5309; SAMPLE 92Zr: 7.0909; SAMPLE 94Zr: 7.5252; SAMPLE 96Zr: 1.2654 | 1 | -6.982e-05 | 2.7717e-05 | 20.3095 | 4.5309 | 7.0908 | 0 | 7.5251 | 0 | 1.2653 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1027 | 20.3095 | 4.5309 | 7.0908 | 7.5251 | 1.2653 | 0.22309 | 0.34914 | 0.37052 | 0.062299 | 1.565 | 1.6608 | 0.27925 | 1.0612 | 0.17844 | 0.16814 | -1.5002 | -1.0523 | -0.99285 | -2.7758 | [10 empty] | 0 | 2.1027 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.9% | -3.4582 | 0.50637 | 7.9709 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100% | 3.9178 | -2.2499 | 17.5847 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 53 | iRM5 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 4.8912e-05; SAMPLE 49Ti40Ar: 7.5125e-06; SAMPLE 90Zr: 19.5619; SAMPLE 91Zr: 4.3642; SAMPLE 92Zr: 6.83; SAMPLE 94Zr: 7.2487; SAMPLE 96Zr: 1.2189 | 1 | -4.8364e-05 | 3.4735e-06 | 19.5618 | 4.3642 | 6.8299 | 0 | 7.2485 | 0 | 1.2188 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.104 | 19.5618 | 4.3642 | 6.8299 | 7.2485 | 1.2188 | 0.2231 | 0.34915 | 0.37054 | 0.062304 | 1.565 | 1.6609 | 0.27927 | 1.0613 | 0.17845 | 0.16814 | -1.5001 | -1.0523 | -0.99279 | -2.7757 | [10 empty] | 0 | 2.104 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0143 | 0.16314 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.42488 | 0.43903 | -1.3883 | 0.014151 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Feb 5, 2020 | 2 | 97.09% | 10.4157 | -9.4061 | -1.5772 | STD | Lab 3 | Feb 4, 2020 | 1 | 100.8% | -1.772 | 1.3651 | 1.7535 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 54 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 3.4452e-05; SAMPLE 49Ti40Ar: 5.2949e-06; SAMPLE 90Zr: 19.9857; SAMPLE 91Zr: 4.4588; SAMPLE 92Zr: 6.9782; SAMPLE 94Zr: 7.406; SAMPLE 96Zr: 1.2454 | 1 | -6.2824e-05 | 1.2559e-06 | 19.9856 | 4.4587 | 6.9781 | 0 | 7.4059 | 0 | 1.2453 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1051 | 19.9856 | 4.4587 | 6.9781 | 7.4059 | 1.2453 | 0.2231 | 0.34916 | 0.37056 | 0.062309 | 1.565 | 1.661 | 0.27929 | 1.0613 | 0.17846 | 0.16815 | -1.5001 | -1.0522 | -0.99274 | -2.7757 | [10 empty] | 0 | 2.1051 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014143 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.85% | 2.3504 | -0.90279 | 17.2432 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.25% | 1.6555 | -6.0196 | -10.116 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 55 | iRM6 | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: -3.6582e-07; SAMPLE 49Ti40Ar: 3.0408e-05; SAMPLE 90Zr: 19.3805; SAMPLE 91Zr: 4.3238; SAMPLE 92Zr: 6.7671; SAMPLE 94Zr: 7.1822; SAMPLE 96Zr: 1.2078 | 1 | -9.7642e-05 | 2.6369e-05 | 19.3805 | 4.3238 | 6.767 | 0 | 7.182 | 0 | 1.2077 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1063 | 19.3805 | 4.3238 | 6.767 | 7.182 | 1.2077 | 0.2231 | 0.34917 | 0.37058 | 0.062315 | 1.5651 | 1.661 | 0.27931 | 1.0613 | 0.17847 | 0.16815 | -1.5001 | -1.0522 | -0.99269 | -2.7756 | [10 empty] | 0 | 2.1063 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.014142 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Feb 5, 2020 | 2 | 97.62% | 2.604 | 2.6373 | 24.793 | STD | Lab 3 | Feb 4, 2020 | 1 | 100.8% | -4.3905 | 4.1697 | -5.0917 | SRM 3169
Lab 3 | Feb 5, 2020 | 2 | 56 | STD | OPZ 177Hf++: 9.7276e-05; OPZ 49Ti40Ar: 4.039e-06; OPZ 90Zr: 5.1238e-05; OPZ 91Zr: 3.0328e-05; OPZ 92Zr: 0.00010462; OPZ 94Zr: 0.0001778; OPZ 96Zr: 0.00010601; SAMPLE 177Hf++: 6.9729e-06; SAMPLE 49Ti40Ar: 1.0037e-05; SAMPLE 90Zr: 19.7217; SAMPLE 91Zr: 4.4; SAMPLE 92Zr: 6.8864; SAMPLE 94Zr: 7.309; SAMPLE 96Zr: 1.2291 | 1 | -9.0303e-05 | 5.9978e-06 | 19.7216 | 4.3999 | 6.8863 | 0 | 7.3088 | 0 | 1.229 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1075 | 19.7216 | 4.3999 | 6.8863 | 7.3088 | 1.229 | 0.2231 | 0.34918 | 0.3706 | 0.062317 | 1.5651 | 1.6611 | 0.27932 | 1.0614 | 0.17847 | 0.16815 | -1.5001 | -1.0522 | -0.99264 | -2.7755 | [10 empty] | 0 | 2.1075 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014145 | -1.8132 | -1.8273 | [22 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.39% | 5.3501 | 4.4688 | 20.7006 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 2 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -3.9451e-05; SAMPLE 49Ti40Ar: 2.4647e-05; SAMPLE 90Zr: 21.2452; SAMPLE 91Zr: 4.7311; SAMPLE 92Zr: 7.3908; SAMPLE 94Zr: 7.8151; SAMPLE 96Zr: 1.3094 | 1 | -0.00025887 | 3.6229e-05 | 21.2449 | 4.731 | 7.3905 | 0 | 7.8148 | 0 | 1.3092 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9362 | 21.2449 | 4.731 | 7.3905 | 7.8148 | 1.3092 | 0.22269 | 0.34787 | 0.36784 | 0.061623 | 1.5622 | 1.6518 | 0.27673 | 1.0574 | 0.17714 | 0.16753 | -1.502 | -1.0559 | -1.0001 | -2.7867 | [9 empty] | L3 JulS1 | 0 | 1.9362 | 0.21796 | 0.33336 | 0.3381 | 0.054374 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43901 | -1.3885 | 0.01411 | -1.8133 | -1.8275 | [22 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.2% | -2.3192 | -8.0462 | -16.8958 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 3 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -2.9899e-05; SAMPLE 49Ti40Ar: 3.3408e-05; SAMPLE 90Zr: 21.3198; SAMPLE 91Zr: 4.7477; SAMPLE 92Zr: 7.4167; SAMPLE 94Zr: 7.8427; SAMPLE 96Zr: 1.314 | 1 | -0.00024932 | 4.4991e-05 | 21.3195 | 4.7476 | 7.4165 | 0 | 7.8424 | 0 | 1.3138 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9366 | 21.3195 | 4.7476 | 7.4165 | 7.8424 | 1.3138 | 0.22269 | 0.34787 | 0.36785 | 0.061624 | 1.5622 | 1.6519 | 0.27673 | 1.0574 | 0.17715 | 0.16753 | -1.502 | -1.0559 | -1.0001 | -2.7867 | [10 empty] | 0 | 1.9366 | 0.21796 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.43901 | -1.3885 | 0.014121 | -1.8133 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 100.6% | -10.6541 | -10.058 | 3.7301 | STD | Lab 3 | Feb 4, 2020 | 1 | 99.84% | -2.7075 | 1.5819 | 1.1183 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 4 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -6.7386e-05; SAMPLE 49Ti40Ar: 3.8657e-05; SAMPLE 90Zr: 21.1526; SAMPLE 91Zr: 4.7105; SAMPLE 92Zr: 7.3586; SAMPLE 94Zr: 7.7812; SAMPLE 96Zr: 1.3037 | 1 | -0.0002868 | 5.024e-05 | 21.1522 | 4.7104 | 7.3584 | 0 | 7.7809 | 0 | 1.3035 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9368 | 21.1522 | 4.7104 | 7.3584 | 7.7809 | 1.3035 | 0.22269 | 0.34788 | 0.36785 | 0.061624 | 1.5622 | 1.6519 | 0.27673 | 1.0574 | 0.17714 | 0.16752 | -1.502 | -1.0559 | -1.0001 | -2.7867 | [10 empty] | 0 | 1.9368 | 0.21797 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.439 | -1.3885 | 0.014111 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.79% | 12.0378 | 3.7885 | -8.139 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.4% | 1.5758 | 6.3546 | -8.3733 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 5 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -1.1776e-05; SAMPLE 49Ti40Ar: 4.6022e-05; SAMPLE 90Zr: 21.0276; SAMPLE 91Zr: 4.6827; SAMPLE 92Zr: 7.3153; SAMPLE 94Zr: 7.7357; SAMPLE 96Zr: 1.2961 | 1 | -0.00023119 | 5.7604e-05 | 21.0272 | 4.6826 | 7.3151 | 0 | 7.7354 | 0 | 1.2959 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9381 | 21.0272 | 4.6826 | 7.3151 | 7.7354 | 1.2959 | 0.22269 | 0.34789 | 0.36787 | 0.06163 | 1.5622 | 1.652 | 0.27675 | 1.0575 | 0.17716 | 0.16753 | -1.502 | -1.0559 | -1 | -2.7866 | [10 empty] | 0 | 1.9381 | 0.21796 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014115 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 99.41% | -11.2534 | 0.062756 | 18.2876 | STD | Lab 3 | Feb 4, 2020 | 1 | 99.43% | 4.2847 | -5.4986 | 19.4169 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 6 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.1819e-05; SAMPLE 49Ti40Ar: 1.2961e-05; SAMPLE 90Zr: 21.1507; SAMPLE 91Zr: 4.7102; SAMPLE 92Zr: 7.3583; SAMPLE 94Zr: 7.7813; SAMPLE 96Zr: 1.3038 | 1 | -0.00027124 | 2.4543e-05 | 21.1504 | 4.7101 | 7.3581 | 0 | 7.781 | 0 | 1.3036 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9391 | 21.1504 | 4.7101 | 7.3581 | 7.781 | 1.3036 | 0.22269 | 0.34789 | 0.36789 | 0.061633 | 1.5622 | 1.652 | 0.27676 | 1.0575 | 0.17716 | 0.16753 | -1.502 | -1.0559 | -0.99997 | -2.7866 | [10 empty] | 0 | 1.9391 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.43901 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.1% | -5.4416 | -4.5445 | -12.3535 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.3% | -2.8833 | 1.7275 | 1.7652 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 7 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.514e-05; SAMPLE 49Ti40Ar: 2.0577e-05; SAMPLE 90Zr: 21.5483; SAMPLE 91Zr: 4.7988; SAMPLE 92Zr: 7.4969; SAMPLE 94Zr: 7.928; SAMPLE 96Zr: 1.3284 | 1 | -0.00027456 | 3.2159e-05 | 21.548 | 4.7987 | 7.4966 | 0 | 7.9277 | 0 | 1.3282 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9404 | 21.548 | 4.7987 | 7.4966 | 7.9277 | 1.3282 | 0.2227 | 0.3479 | 0.36791 | 0.061638 | 1.5622 | 1.6521 | 0.27678 | 1.0575 | 0.17717 | 0.16754 | -1.5019 | -1.0558 | -0.99991 | -2.7865 | [10 empty] | 0 | 1.9404 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.439 | -1.3885 | 0.014117 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 102% | 11.2114 | 1.1943 | -7.539 | STD | Lab 3 | Feb 4, 2020 | 1 | 100.1% | -3.2531 | 1.1435 | -5.9106 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 8 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -6.2746e-05; SAMPLE 49Ti40Ar: 2.3027e-05; SAMPLE 90Zr: 21.0981; SAMPLE 91Zr: 4.6986; SAMPLE 92Zr: 7.3403; SAMPLE 94Zr: 7.7627; SAMPLE 96Zr: 1.3007 | 1 | -0.00028216 | 3.461e-05 | 21.0977 | 4.6984 | 7.3401 | 0 | 7.7623 | 0 | 1.3005 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9412 | 21.0977 | 4.6984 | 7.3401 | 7.7623 | 1.3005 | 0.2227 | 0.34791 | 0.36792 | 0.061642 | 1.5622 | 1.6521 | 0.2768 | 1.0575 | 0.17718 | 0.16754 | -1.5019 | -1.0558 | -0.99988 | -2.7864 | [10 empty] | 0 | 1.9412 | 0.21796 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43901 | -1.3885 | 0.014118 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.94% | -4.8502 | -3.7329 | 25.0672 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100% | 2.3987 | -3.8894 | -5.4993 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 9 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.0632e-06; SAMPLE 49Ti40Ar: 2.6652e-05; SAMPLE 90Zr: 21.2823; SAMPLE 91Zr: 4.7397; SAMPLE 92Zr: 7.4047; SAMPLE 94Zr: 7.831; SAMPLE 96Zr: 1.3122 | 1 | -0.00022448 | 3.8235e-05 | 21.282 | 4.7396 | 7.4045 | 0 | 7.8307 | 0 | 1.312 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9427 | 21.282 | 4.7396 | 7.4045 | 7.8307 | 1.312 | 0.2227 | 0.34792 | 0.36795 | 0.061648 | 1.5623 | 1.6522 | 0.27682 | 1.0576 | 0.17719 | 0.16755 | -1.5019 | -1.0558 | -0.99981 | -2.7863 | [10 empty] | 0 | 1.9427 | 0.21796 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43901 | -1.3885 | 0.014113 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 100.9% | -3.4693 | 1.2326 | 14.4317 | STD | Lab 3 | Feb 4, 2020 | 1 | 99.66% | 0.80296 | -2.4608 | -12.4965 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 10 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.2009e-05; SAMPLE 49Ti40Ar: 3.3892e-05; SAMPLE 90Zr: 21.0705; SAMPLE 91Zr: 4.6925; SAMPLE 92Zr: 7.3311; SAMPLE 94Zr: 7.7531; SAMPLE 96Zr: 1.2991 | 1 | -0.00027143 | 4.5475e-05 | 21.0702 | 4.6924 | 7.3308 | 0 | 7.7528 | 0 | 1.2989 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9429 | 21.0702 | 4.6924 | 7.3308 | 7.7528 | 1.2989 | 0.2227 | 0.34792 | 0.36795 | 0.061647 | 1.5623 | 1.6522 | 0.27681 | 1.0576 | 0.17718 | 0.16754 | -1.5019 | -1.0558 | -0.99981 | -2.7863 | [10 empty] | 0 | 1.9429 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43901 | -1.3885 | 0.01411 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.2% | 9.7141 | 8.8534 | -16.4419 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 100.5% | -0.71756 | 11.0962 | 15.2357 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 11 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -3.4647e-05; SAMPLE 49Ti40Ar: 6.6422e-05; SAMPLE 90Zr: 20.8904; SAMPLE 91Zr: 4.6525; SAMPLE 92Zr: 7.2686; SAMPLE 94Zr: 7.6873; SAMPLE 96Zr: 1.2881 | 1 | -0.00025407 | 7.8004e-05 | 20.89 | 4.6524 | 7.2683 | 0 | 7.687 | 0 | 1.288 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9445 | 20.89 | 4.6524 | 7.2683 | 7.687 | 1.288 | 0.22271 | 0.34793 | 0.36798 | 0.061654 | 1.5623 | 1.6523 | 0.27684 | 1.0576 | 0.1772 | 0.16755 | -1.5019 | -1.0557 | -0.99974 | -2.7862 | [10 empty] | 0 | 1.9445 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.43901 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.45% | 4.1437 | -5.9899 | 11.9243 | STD | Lab 3 | Feb 4, 2020 | 1 | 99.98% | 5.3237 | -7.3137 | 0.004329 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 12 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -6.7895e-05; SAMPLE 49Ti40Ar: -6.7218e-06; SAMPLE 90Zr: 20.9432; SAMPLE 91Zr: 4.6643; SAMPLE 92Zr: 7.2871; SAMPLE 94Zr: 7.7071; SAMPLE 96Zr: 1.2915 | 1 | -0.00028731 | 4.8606e-06 | 20.9429 | 4.6641 | 7.2869 | 0 | 7.7068 | 0 | 1.2913 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9455 | 20.9429 | 4.6641 | 7.2869 | 7.7068 | 1.2913 | 0.22271 | 0.34794 | 0.36799 | 0.061657 | 1.5623 | 1.6524 | 0.27685 | 1.0576 | 0.17721 | 0.16755 | -1.5019 | -1.0557 | -0.99969 | -2.7862 | [10 empty] | 0 | 1.9455 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.76% | -7.0257 | -8.4807 | 6.0136 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.54% | -6.6654 | 0.60882 | 10.3551 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 13 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -3.2836e-05; SAMPLE 49Ti40Ar: 2.7607e-05; SAMPLE 90Zr: 21.33; SAMPLE 91Zr: 4.7505; SAMPLE 92Zr: 7.4219; SAMPLE 94Zr: 7.8498; SAMPLE 96Zr: 1.3154 | 1 | -0.00025225 | 3.919e-05 | 21.3296 | 4.7503 | 7.4217 | 0 | 7.8495 | 0 | 1.3152 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9466 | 21.3296 | 4.7503 | 7.4217 | 7.8495 | 1.3152 | 0.22271 | 0.34795 | 0.36801 | 0.061663 | 1.5623 | 1.6524 | 0.27687 | 1.0577 | 0.17722 | 0.16756 | -1.5019 | -1.0557 | -0.99964 | -2.7861 | [10 empty] | 0 | 1.9466 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014118 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 101.9% | -3.3751 | -2.6575 | 19.8375 | STD | Lab 3 | Feb 4, 2020 | 1 | 100.7% | 1.2097 | 2.2562 | -22.1904 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 14 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -7.1635e-05; SAMPLE 49Ti40Ar: -1.2946e-07; SAMPLE 90Zr: 20.9171; SAMPLE 91Zr: 4.6586; SAMPLE 92Zr: 7.2785; SAMPLE 94Zr: 7.6984; SAMPLE 96Zr: 1.2901 | 1 | -0.00029105 | 1.1453e-05 | 20.9168 | 4.6585 | 7.2783 | 0 | 7.6981 | 0 | 1.2899 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9481 | 20.9168 | 4.6585 | 7.2783 | 7.6981 | 1.2899 | 0.22271 | 0.34796 | 0.36803 | 0.061667 | 1.5624 | 1.6525 | 0.27689 | 1.0577 | 0.17722 | 0.16756 | -1.5019 | -1.0557 | -0.99958 | -2.786 | [10 empty] | 0 | 1.9481 | 0.21796 | 0.33336 | 0.3381 | 0.05437 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43902 | -1.3885 | 0.014112 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.4% | -3.8247 | 8.6646 | -16.9887 | [8 empty] | STD | Lab 3 | Feb 4, 2020 | 1 | 99.09% | -2.493 | -2.5694 | -1.5906 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 15 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -3.6311e-05; SAMPLE 49Ti40Ar: 1.7883e-05; SAMPLE 90Zr: 20.7552; SAMPLE 91Zr: 4.6227; SAMPLE 92Zr: 7.2226; SAMPLE 94Zr: 7.6396; SAMPLE 96Zr: 1.2803 | 1 | -0.00025573 | 2.9465e-05 | 20.7549 | 4.6226 | 7.2223 | 0 | 7.6393 | 0 | 1.2801 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9506 | 20.7549 | 4.6226 | 7.2223 | 7.6393 | 1.2801 | 0.22272 | 0.34798 | 0.36807 | 0.061678 | 1.5624 | 1.6526 | 0.27693 | 1.0577 | 0.17724 | 0.16757 | -1.5018 | -1.0556 | -0.99947 | -2.7858 | [10 empty] | 0 | 1.9506 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43901 | -1.3885 | 0.014113 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.68% | 4.2918 | 3.3048 | 8.1331 | STD | Lab 3 | Feb 4, 2020 | 1 | 100.8% | 2.9709 | 4.1042 | 22.3398 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 16 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -6.5581e-05; SAMPLE 49Ti40Ar: 2.4041e-05; SAMPLE 90Zr: 20.727; SAMPLE 91Zr: 4.6165; SAMPLE 92Zr: 7.2128; SAMPLE 94Zr: 7.6295; SAMPLE 96Zr: 1.2786 | 1 | -0.000285 | 3.5623e-05 | 20.7266 | 4.6163 | 7.2126 | 0 | 7.6292 | 0 | 1.2784 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9513 | 20.7266 | 4.6163 | 7.2126 | 7.6292 | 1.2784 | 0.22272 | 0.34799 | 0.36809 | 0.061681 | 1.5624 | 1.6526 | 0.27694 | 1.0578 | 0.17725 | 0.16757 | -1.5018 | -1.0556 | -0.99944 | -2.7858 | [10 empty] | 0 | 1.9513 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.43901 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.66% | 14.1624 | -5.6479 | 11.5354 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.5% | -4.8883 | 4.8246 | -8.6769 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 17 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: 1.0634e-05; SAMPLE 49Ti40Ar: 2.4382e-05; SAMPLE 90Zr: 20.5669; SAMPLE 91Zr: 4.5808; SAMPLE 92Zr: 7.1573; SAMPLE 94Zr: 7.5709; SAMPLE 96Zr: 1.2689 | 1 | -0.00020878 | 3.5964e-05 | 20.5665 | 4.5807 | 7.157 | 0 | 7.5706 | 0 | 1.2687 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9525 | 20.5665 | 4.5807 | 7.157 | 7.5706 | 1.2687 | 0.22272 | 0.34799 | 0.3681 | 0.061686 | 1.5624 | 1.6527 | 0.27696 | 1.0578 | 0.17726 | 0.16758 | -1.5018 | -1.0556 | -0.99939 | -2.7857 | [10 empty] | 0 | 1.9525 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43903 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 99.34% | -7.9832 | 0.6368 | 10.7505 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.03% | 7.8548 | -10.9401 | 21.6619 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 18 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -4.6526e-05; SAMPLE 49Ti40Ar: 2.6812e-05; SAMPLE 90Zr: 20.679; SAMPLE 91Zr: 4.6058; SAMPLE 92Zr: 7.1966; SAMPLE 94Zr: 7.6128; SAMPLE 96Zr: 1.2759 | 1 | -0.00026595 | 3.8394e-05 | 20.6787 | 4.6057 | 7.1963 | 0 | 7.6125 | 0 | 1.2757 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9542 | 20.6787 | 4.6057 | 7.1963 | 7.6125 | 1.2757 | 0.22273 | 0.34801 | 0.36813 | 0.061692 | 1.5625 | 1.6528 | 0.27698 | 1.0578 | 0.17727 | 0.16758 | -1.5018 | -1.0555 | -0.99931 | -2.7856 | [10 empty] | 0 | 1.9542 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43903 | -1.3885 | 0.01412 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.1% | -8.4448 | 3.0889 | -0.97592 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | -3.0579 | 14.0674 | -14.4267 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 19 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -2.4321e-05; SAMPLE 49Ti40Ar: 6.2659e-05; SAMPLE 90Zr: 21.0236; SAMPLE 91Zr: 4.6828; SAMPLE 92Zr: 7.3167; SAMPLE 94Zr: 7.74; SAMPLE 96Zr: 1.2973 | 1 | -0.00024374 | 7.4241e-05 | 21.0233 | 4.6826 | 7.3164 | 0 | 7.7397 | 0 | 1.2971 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9553 | 21.0233 | 4.6826 | 7.3164 | 7.7397 | 1.2971 | 0.22274 | 0.34802 | 0.36815 | 0.061698 | 1.5625 | 1.6529 | 0.277 | 1.0579 | 0.17728 | 0.16759 | -1.5018 | -1.0555 | -0.99927 | -2.7855 | [10 empty] | 0 | 1.9553 | 0.21797 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42488 | 0.43901 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 101.9% | 23.5368 | 0.63457 | 15.7721 | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | 1.4513 | -10.9498 | 0.53963 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 20 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -4.4528e-05; SAMPLE 49Ti40Ar: 3.3724e-05; SAMPLE 90Zr: 20.5935; SAMPLE 91Zr: 4.5869; SAMPLE 92Zr: 7.1672; SAMPLE 94Zr: 7.5821; SAMPLE 96Zr: 1.2708 | 1 | -0.00026395 | 4.5306e-05 | 20.5931 | 4.5868 | 7.1669 | 0 | 7.5818 | 0 | 1.2706 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9566 | 20.5931 | 4.5868 | 7.1669 | 7.5818 | 1.2706 | 0.22273 | 0.34802 | 0.36817 | 0.061702 | 1.5625 | 1.653 | 0.27702 | 1.0579 | 0.17729 | 0.16759 | -1.5018 | -1.0555 | -0.99921 | -2.7854 | [10 empty] | 0 | 1.9566 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43903 | -1.3885 | 0.014125 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.73% | -3.1548 | -5.2888 | 2.1347 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.87% | 0.8018 | 0.45953 | 1.9173 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 21 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: 1.9171e-05; SAMPLE 49Ti40Ar: 3.1188e-05; SAMPLE 90Zr: 20.7747; SAMPLE 91Zr: 4.6274; SAMPLE 92Zr: 7.2306; SAMPLE 94Zr: 7.6494; SAMPLE 96Zr: 1.2822 | 1 | -0.00020025 | 4.2771e-05 | 20.7744 | 4.6273 | 7.2303 | 0 | 7.6491 | 0 | 1.282 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9583 | 20.7744 | 4.6273 | 7.2303 | 7.6491 | 1.282 | 0.22274 | 0.34804 | 0.3682 | 0.061709 | 1.5625 | 1.653 | 0.27705 | 1.0579 | 0.17731 | 0.1676 | -1.5017 | -1.0554 | -0.99913 | -2.7853 | [10 empty] | 0 | 1.9583 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014118 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 100.8% | 11.5945 | 3.8951 | 11.4528 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.91% | 0.99418 | 7.3862 | -25.4678 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 22 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.5451e-05; SAMPLE 49Ti40Ar: 5.6933e-05; SAMPLE 90Zr: 20.6193; SAMPLE 91Zr: 4.5928; SAMPLE 92Zr: 7.1766; SAMPLE 94Zr: 7.5924; SAMPLE 96Zr: 1.2726 | 1 | -0.00027487 | 6.8515e-05 | 20.619 | 4.5927 | 7.1763 | 0 | 7.5921 | 0 | 1.2724 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9589 | 20.619 | 4.5927 | 7.1763 | 7.5921 | 1.2724 | 0.22274 | 0.34804 | 0.36821 | 0.061711 | 1.5626 | 1.6531 | 0.27705 | 1.0579 | 0.17731 | 0.1676 | -1.5018 | -1.0554 | -0.99911 | -2.7853 | [10 empty] | 0 | 1.9589 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43903 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.3% | -5.8908 | -0.81499 | -15.5471 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | -6.4906 | -7.3738 | 54.2709 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 23 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -4.8253e-05; SAMPLE 49Ti40Ar: 1.6379e-05; SAMPLE 90Zr: 20.4569; SAMPLE 91Zr: 4.5567; SAMPLE 92Zr: 7.1202; SAMPLE 94Zr: 7.5331; SAMPLE 96Zr: 1.2627 | 1 | -0.00026767 | 2.7962e-05 | 20.4566 | 4.5566 | 7.12 | 0 | 7.5328 | 0 | 1.2625 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9606 | 20.4566 | 4.5566 | 7.12 | 7.5328 | 1.2625 | 0.22274 | 0.34805 | 0.36823 | 0.061718 | 1.5626 | 1.6532 | 0.27708 | 1.058 | 0.17732 | 0.16761 | -1.5017 | -1.0554 | -0.99904 | -2.7852 | [10 empty] | 0 | 1.9606 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014128 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.49% | -4.9096 | -13.2281 | -1.9245 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.34% | 6.8507 | 4.419 | -33.9777 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 24 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.5988e-05; SAMPLE 49Ti40Ar: 3.5133e-05; SAMPLE 90Zr: 20.5023; SAMPLE 91Zr: 4.5669; SAMPLE 92Zr: 7.1364; SAMPLE 94Zr: 7.5503; SAMPLE 96Zr: 1.2657 | 1 | -0.00027541 | 4.6715e-05 | 20.5019 | 4.5668 | 7.1361 | 0 | 7.55 | 0 | 1.2655 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.962 | 20.5019 | 4.5668 | 7.1361 | 7.55 | 1.2655 | 0.22275 | 0.34807 | 0.36826 | 0.061725 | 1.5626 | 1.6532 | 0.2771 | 1.058 | 0.17733 | 0.16761 | -1.5017 | -1.0553 | -0.99897 | -2.7851 | [10 empty] | 0 | 1.962 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43901 | -1.3885 | 0.014111 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.76% | 18.2218 | 15.4158 | 36.1148 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.93% | -4.5354 | -7.4393 | -10.0301 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 25 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.7218e-05; SAMPLE 49Ti40Ar: 1.2571e-05; SAMPLE 90Zr: 20.7713; SAMPLE 91Zr: 4.6269; SAMPLE 92Zr: 7.2301; SAMPLE 94Zr: 7.6499; SAMPLE 96Zr: 1.2824 | 1 | -0.00027664 | 2.4154e-05 | 20.771 | 4.6267 | 7.2299 | 0 | 7.6496 | 0 | 1.2822 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9635 | 20.771 | 4.6267 | 7.2299 | 7.6496 | 1.2822 | 0.22275 | 0.34808 | 0.36828 | 0.06173 | 1.5626 | 1.6533 | 0.27713 | 1.058 | 0.17735 | 0.16762 | -1.5017 | -1.0553 | -0.99891 | -2.785 | [10 empty] | 0 | 1.9635 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43903 | -1.3885 | 0.014128 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 101.4% | -9.3041 | -5.1812 | 9.4346 | STD | Lab 3 | Feb 5, 2020 | 2 | 100.6% | 3.8094 | 13.2507 | 28.9739 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 26 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.3978e-05; SAMPLE 49Ti40Ar: 6.3295e-05; SAMPLE 90Zr: 20.4855; SAMPLE 91Zr: 4.5633; SAMPLE 92Zr: 7.1309; SAMPLE 94Zr: 7.5452; SAMPLE 96Zr: 1.2648 | 1 | -0.0002734 | 7.4877e-05 | 20.4852 | 4.5632 | 7.1306 | 0 | 7.5449 | 0 | 1.2646 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9653 | 20.4852 | 4.5632 | 7.1306 | 7.5449 | 1.2646 | 0.22275 | 0.34809 | 0.36831 | 0.061735 | 1.5627 | 1.6534 | 0.27714 | 1.0581 | 0.17735 | 0.16762 | -1.5017 | -1.0553 | -0.99883 | -2.7849 | [10 empty] | 0 | 1.9653 | 0.21796 | 0.33335 | 0.3381 | 0.05437 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43903 | -1.3885 | 0.014135 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.96% | -16.2128 | -21.3256 | -41.9616 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.87% | -1.3111 | -9.0457 | -36.6994 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 27 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -1.348e-05; SAMPLE 49Ti40Ar: 2.9541e-05; SAMPLE 90Zr: 20.5486; SAMPLE 91Zr: 4.5774; SAMPLE 92Zr: 7.1532; SAMPLE 94Zr: 7.5689; SAMPLE 96Zr: 1.2689 | 1 | -0.0002329 | 4.1124e-05 | 20.5483 | 4.5773 | 7.1529 | 0 | 7.5686 | 0 | 1.2687 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9667 | 20.5483 | 4.5773 | 7.1529 | 7.5686 | 1.2687 | 0.22276 | 0.3481 | 0.36833 | 0.061742 | 1.5627 | 1.6535 | 0.27717 | 1.0581 | 0.17737 | 0.16763 | -1.5017 | -1.0553 | -0.99877 | -2.7848 | [10 empty] | 0 | 1.9667 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43903 | -1.3885 | 0.014124 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 100.3% | -2.005 | 1.0024 | -0.24635 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.32% | -3.6179 | -1.2147 | 39.2921 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 28 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -8.6783e-05; SAMPLE 49Ti40Ar: 2.6899e-07; SAMPLE 90Zr: 20.4841; SAMPLE 91Zr: 4.5631; SAMPLE 92Zr: 7.1309; SAMPLE 94Zr: 7.5454; SAMPLE 96Zr: 1.265 | 1 | -0.0003062 | 1.1851e-05 | 20.4838 | 4.563 | 7.1306 | 0 | 7.5451 | 0 | 1.2648 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9676 | 20.4838 | 4.563 | 7.1306 | 7.5451 | 1.2648 | 0.22276 | 0.34811 | 0.36835 | 0.061747 | 1.5627 | 1.6535 | 0.27719 | 1.0581 | 0.17738 | 0.16763 | -1.5016 | -1.0552 | -0.99873 | -2.7847 | [10 empty] | 0 | 1.9676 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014115 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.4% | 6.2448 | 15.0859 | 23.3313 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 101.3% | 3.7659 | 4.6149 | -13.5627 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 29 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -3.4655e-05; SAMPLE 49Ti40Ar: 4.3582e-05; SAMPLE 90Zr: 20.2149; SAMPLE 91Zr: 4.5033; SAMPLE 92Zr: 7.0374; SAMPLE 94Zr: 7.4468; SAMPLE 96Zr: 1.2485 | 1 | -0.00025407 | 5.5164e-05 | 20.2145 | 4.5031 | 7.0371 | 0 | 7.4465 | 0 | 1.2483 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9694 | 20.2145 | 4.5031 | 7.0371 | 7.4465 | 1.2483 | 0.22277 | 0.34812 | 0.36838 | 0.061754 | 1.5627 | 1.6536 | 0.27721 | 1.0582 | 0.17739 | 0.16764 | -1.5016 | -1.0552 | -0.99865 | -2.7846 | [10 empty] | 0 | 1.9694 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.43902 | -1.3885 | 0.014125 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 99.12% | 3.2016 | -4.6299 | 4.6052 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.04% | -4.2096 | -2.2919 | -11.1291 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 30 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.8505e-05; SAMPLE 49Ti40Ar: 1.106e-05; SAMPLE 90Zr: 20.3049; SAMPLE 91Zr: 4.5234; SAMPLE 92Zr: 7.0689; SAMPLE 94Zr: 7.4804; SAMPLE 96Zr: 1.2542 | 1 | -0.00027792 | 2.2643e-05 | 20.3045 | 4.5232 | 7.0686 | 0 | 7.4801 | 0 | 1.254 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9706 | 20.3045 | 4.5232 | 7.0686 | 7.4801 | 1.254 | 0.22277 | 0.34813 | 0.3684 | 0.061758 | 1.5627 | 1.6537 | 0.27723 | 1.0582 | 0.1774 | 0.16764 | -1.5016 | -1.0552 | -0.9986 | -2.7845 | [10 empty] | 0 | 1.9706 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.43902 | -1.3885 | 0.014126 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.8% | 0.2744 | -8.8044 | -6.0968 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.4% | 6.4009 | -0.93341 | 10.8244 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 31 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -5.6195e-05; SAMPLE 49Ti40Ar: 3.1677e-05; SAMPLE 90Zr: 20.629; SAMPLE 91Zr: 4.5957; SAMPLE 92Zr: 7.1822; SAMPLE 94Zr: 7.6006; SAMPLE 96Zr: 1.2744 | 1 | -0.00027561 | 4.3259e-05 | 20.6287 | 4.5956 | 7.1819 | 0 | 7.6003 | 0 | 1.2742 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9729 | 20.6287 | 4.5956 | 7.1819 | 7.6003 | 1.2742 | 0.22277 | 0.34815 | 0.36843 | 0.061767 | 1.5628 | 1.6538 | 0.27726 | 1.0582 | 0.17741 | 0.16765 | -1.5016 | -1.0551 | -0.9985 | -2.7844 | [10 empty] | 0 | 1.9729 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5295 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43902 | -1.3885 | 0.014114 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 101.8% | -2.8513 | 8.5085 | -10.3395 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.78% | -4.7843 | -2.645 | -6.0441 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 32 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -6.062e-05; SAMPLE 49Ti40Ar: 1.7463e-05; SAMPLE 90Zr: 20.2055; SAMPLE 91Zr: 4.5014; SAMPLE 92Zr: 7.0349; SAMPLE 94Zr: 7.4449; SAMPLE 96Zr: 1.2483 | 1 | -0.00028004 | 2.9045e-05 | 20.2052 | 4.5013 | 7.0347 | 0 | 7.4446 | 0 | 1.2481 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9742 | 20.2052 | 4.5013 | 7.0347 | 7.4446 | 1.2481 | 0.22278 | 0.34816 | 0.36845 | 0.061772 | 1.5628 | 1.6539 | 0.27728 | 1.0583 | 0.17743 | 0.16765 | -1.5016 | -1.0551 | -0.99844 | -2.7843 | [10 empty] | 0 | 1.9742 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014119 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.89% | 4.2681 | 6.6349 | 3.491 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | 2.6893 | 7.1043 | 10.1579 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 33 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -1.5843e-05; SAMPLE 49Ti40Ar: 4.4805e-05; SAMPLE 90Zr: 20.2322; SAMPLE 91Zr: 4.5075; SAMPLE 92Zr: 7.0445; SAMPLE 94Zr: 7.4555; SAMPLE 96Zr: 1.2502 | 1 | -0.00023526 | 5.6387e-05 | 20.2319 | 4.5074 | 7.0443 | 0 | 7.4552 | 0 | 1.25 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9764 | 20.2319 | 4.5074 | 7.0443 | 7.4552 | 1.25 | 0.22279 | 0.34818 | 0.36849 | 0.061783 | 1.5628 | 1.654 | 0.27732 | 1.0583 | 0.17745 | 0.16767 | -1.5015 | -1.055 | -0.99835 | -2.7841 | [10 empty] | 0 | 1.9764 | 0.21796 | 0.33336 | 0.3381 | 0.054373 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.16311 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42489 | 0.43902 | -1.3885 | 0.014121 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 100.3% | 10.9911 | 1.8885 | 22.7935 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.42% | -4.4766 | -4.2412 | -21.3077 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 34 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -7.284e-05; SAMPLE 49Ti40Ar: -2.3158e-06; SAMPLE 90Zr: 20.1498; SAMPLE 91Zr: 4.4891; SAMPLE 92Zr: 7.016; SAMPLE 94Zr: 7.4255; SAMPLE 96Zr: 1.2451 | 1 | -0.00029226 | 9.2666e-06 | 20.1494 | 4.489 | 7.0157 | 0 | 7.4252 | 0 | 1.245 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9776 | 20.1494 | 4.489 | 7.0157 | 7.4252 | 1.245 | 0.22278 | 0.34819 | 0.36851 | 0.061786 | 1.5629 | 1.6541 | 0.27733 | 1.0584 | 0.17745 | 0.16767 | -1.5015 | -1.055 | -0.99829 | -2.7841 | [10 empty] | 0 | 1.9776 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43903 | -1.3885 | 0.014126 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.4% | -3.5453 | 0.65169 | 4.3975 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.9% | 6.2576 | 4.3813 | 26.6193 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 35 | iRM4 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -4.1121e-05; SAMPLE 49Ti40Ar: 2.1594e-05; SAMPLE 90Zr: 19.9343; SAMPLE 91Zr: 4.4412; SAMPLE 92Zr: 6.9413; SAMPLE 94Zr: 7.3467; SAMPLE 96Zr: 1.232 | 1 | -0.00026054 | 3.3177e-05 | 19.9339 | 4.4411 | 6.941 | 0 | 7.3464 | 0 | 1.2318 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9795 | 19.9339 | 4.4411 | 6.941 | 7.3464 | 1.2318 | 0.22279 | 0.3482 | 0.36854 | 0.061792 | 1.5629 | 1.6542 | 0.27736 | 1.0584 | 0.17746 | 0.16767 | -1.5015 | -1.055 | -0.99821 | -2.784 | [10 empty] | 0 | 1.9795 | 0.21796 | 0.33336 | 0.3381 | 0.05437 | 1.5295 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.42492 | 0.43903 | -1.3885 | 0.014117 | -1.8134 | -1.8275 | [14 empty] | iRM4 | Lab 3 | Jul 20, 2020 | 1 | 99.51% | -0.038047 | 12.9627 | -14.124 | STD | Lab 3 | Feb 5, 2020 | 2 | 99.22% | -1.6491 | -8.2986 | -27.9114 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 36 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -2.979e-05; SAMPLE 49Ti40Ar: 1.8539e-05; SAMPLE 90Zr: 19.9137; SAMPLE 91Zr: 4.4367; SAMPLE 92Zr: 6.9342; SAMPLE 94Zr: 7.3396; SAMPLE 96Zr: 1.2308 | 1 | -0.00024921 | 3.0121e-05 | 19.9133 | 4.4365 | 6.934 | 0 | 7.3393 | 0 | 1.2306 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.981 | 19.9133 | 4.4365 | 6.934 | 7.3393 | 1.2306 | 0.22279 | 0.34821 | 0.36856 | 0.061799 | 1.5629 | 1.6543 | 0.27738 | 1.0585 | 0.17748 | 0.16768 | -1.5015 | -1.055 | -0.99814 | -2.7839 | [10 empty] | 0 | 1.981 | 0.21796 | 0.33336 | 0.3381 | 0.054371 | 1.5294 | 1.5512 | 0.24945 | 1.0142 | 0.1631 | 0.16081 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43903 | -1.3885 | 0.014133 | -1.8134 | -1.8275 | [6 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.64% | -2.5118 | -8.3052 | -18.5897 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 100.2% | 2.8763 | 11.291 | 39.1196 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 37 | iRM5 | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -1.3218e-05; SAMPLE 49Ti40Ar: 2.2141e-05; SAMPLE 90Zr: 20.2511; SAMPLE 91Zr: 4.512; SAMPLE 92Zr: 7.0522; SAMPLE 94Zr: 7.4648; SAMPLE 96Zr: 1.2519 | 1 | -0.00023264 | 3.3723e-05 | 20.2508 | 4.5119 | 7.052 | 0 | 7.4645 | 0 | 1.2517 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9835 | 20.2508 | 4.5119 | 7.052 | 7.4645 | 1.2517 | 0.2228 | 0.34823 | 0.3686 | 0.06181 | 1.563 | 1.6544 | 0.27742 | 1.0585 | 0.1775 | 0.16769 | -1.5015 | -1.0549 | -0.99804 | -2.7837 | [10 empty] | 0 | 1.9835 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.4249 | 0.43902 | -1.3885 | 0.014122 | -1.8134 | -1.8275 | [14 empty] | iRM5 | Lab 3 | Jul 20, 2020 | 1 | 101.9% | 10.5067 | 7.44 | 10.4908 | STD | Lab 3 | Feb 5, 2020 | 2 | 100.3% | -6.7233 | -4.4336 | -50.0868 | SRM 3169
Lab 3 | Jul 20, 2020 | 1 | 38 | STD | OPZ 177Hf++: 0.00021942; OPZ 49Ti40Ar: -1.1582e-05; OPZ 90Zr: 0.0003423; OPZ 91Zr: 0.00011827; OPZ 92Zr: 0.00025133; OPZ 94Zr: 0.00030291; OPZ 96Zr: 0.00019274; SAMPLE 177Hf++: -6.3241e-05; SAMPLE 49Ti40Ar: 2.2343e-05; SAMPLE 90Zr: 19.8213; SAMPLE 91Zr: 4.4162; SAMPLE 92Zr: 6.9026; SAMPLE 94Zr: 7.3066; SAMPLE 96Zr: 1.2254 | 1 | -0.00028266 | 3.3925e-05 | 19.8209 | 4.4161 | 6.9024 | 0 | 7.3063 | 0 | 1.2252 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 1.9844 | 19.8209 | 4.4161 | 6.9024 | 7.3063 | 1.2252 | 0.2228 | 0.34824 | 0.36862 | 0.061814 | 1.563 | 1.6545 | 0.27744 | 1.0585 | 0.17751 | 0.16769 | -1.5015 | -1.0549 | -0.998 | -2.7836 | [10 empty] | 0 | 1.9844 | 0.21796 | 0.33336 | 0.3381 | 0.054372 | 1.5294 | 1.5512 | 0.24946 | 1.0142 | 0.1631 | 0.16082 | -1.5234 | -1.0985 | -2.9119 | 0.42491 | 0.43903 | -1.3885 | 0.014125 | -1.8134 | -1.8275 | [22 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.59% | 0.51886 | -7.3935 | 54.0279 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 2 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 1.9479e-05; SAMPLE 49Ti40Ar: 1.7585e-05; SAMPLE 90Zr: 17.4706; SAMPLE 91Zr: 3.8961; SAMPLE 92Zr: 6.0954; SAMPLE 94Zr: 6.464; SAMPLE 96Zr: 1.0861 | 1 | -0.00013189 | 2.647e-05 | 17.4703 | 3.8961 | 6.0952 | 0 | 6.4638 | 0 | 1.086 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0697 | 17.4703 | 3.8961 | 6.0952 | 6.4638 | 1.086 | 0.22301 | 0.34889 | 0.36999 | 0.062163 | 1.5645 | 1.6591 | 0.27875 | 1.0605 | 0.17817 | 0.16801 | -1.5005 | -1.053 | -0.99428 | -2.778 | [9 empty] | L3 JulS2 | 0 | 2.0697 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014131 | -1.8132 | -1.8274 | [22 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.68% | 4.7135 | 6.4741 | -37.7691 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 3 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 5.7924e-06; SAMPLE 49Ti40Ar: 1.2913e-05; SAMPLE 90Zr: 17.4494; SAMPLE 91Zr: 3.8915; SAMPLE 92Zr: 6.0883; SAMPLE 94Zr: 6.4569; SAMPLE 96Zr: 1.085 | 1 | -0.00014558 | 2.1797e-05 | 17.4492 | 3.8914 | 6.0882 | 0 | 6.4567 | 0 | 1.0849 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0722 | 17.4492 | 3.8914 | 6.0882 | 6.4567 | 1.0849 | 0.22301 | 0.34891 | 0.37003 | 0.062172 | 1.5645 | 1.6592 | 0.27878 | 1.0605 | 0.17819 | 0.16802 | -1.5005 | -1.0529 | -0.99417 | -2.7778 | [10 empty] | 0 | 2.0722 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014132 | -1.8133 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -0.84304 | 1.1443 | -8.3872 | STD | Lab 3 | Feb 5, 2020 | 2 | 100.9% | -3.4582 | 0.50637 | 7.9709 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 4 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -8.8677e-06; SAMPLE 49Ti40Ar: 1.3733e-05; SAMPLE 90Zr: 17.3798; SAMPLE 91Zr: 3.876; SAMPLE 92Zr: 6.0642; SAMPLE 94Zr: 6.4316; SAMPLE 96Zr: 1.0808 | 1 | -0.00016024 | 2.2617e-05 | 17.3795 | 3.8759 | 6.0641 | 0 | 6.4314 | 0 | 1.0807 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.074 | 17.3795 | 3.8759 | 6.0641 | 6.4314 | 1.0807 | 0.22302 | 0.34892 | 0.37006 | 0.06218 | 1.5645 | 1.6593 | 0.27881 | 1.0606 | 0.17821 | 0.16803 | -1.5005 | -1.0529 | -0.9941 | -2.7777 | [10 empty] | 0 | 2.074 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.9% | -10.3757 | -1.3011 | -13.101 | [8 empty] | STD | Lab 3 | Feb 5, 2020 | 2 | 99.85% | 2.3504 | -0.90279 | 17.2432 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 5 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 2.012e-05; SAMPLE 49Ti40Ar: 3.5011e-05; SAMPLE 90Zr: 17.3255; SAMPLE 91Zr: 3.864; SAMPLE 92Zr: 6.0456; SAMPLE 94Zr: 6.4122; SAMPLE 96Zr: 1.0776 | 1 | -0.00013125 | 4.3895e-05 | 17.3253 | 3.8639 | 6.0455 | 0 | 6.412 | 0 | 1.0774 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0765 | 17.3253 | 3.8639 | 6.0455 | 6.412 | 1.0774 | 0.22302 | 0.34894 | 0.3701 | 0.062189 | 1.5646 | 1.6595 | 0.27885 | 1.0606 | 0.17822 | 0.16803 | -1.5005 | -1.0528 | -0.99399 | -2.7776 | [10 empty] | 0 | 2.0765 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.01413 | -1.8133 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | -5.4111 | 7.6793 | -19.1403
Lab 3 | Jul 21, 2020 | 2 | 6 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 2.846e-05; SAMPLE 49Ti40Ar: 2.4341e-05; SAMPLE 90Zr: 16.9773; SAMPLE 91Zr: 3.7865; SAMPLE 92Zr: 5.9243; SAMPLE 94Zr: 6.2838; SAMPLE 96Zr: 1.056 | 1 | -0.00012291 | 3.3225e-05 | 16.9771 | 3.7864 | 5.9242 | 0 | 6.2836 | 0 | 1.0559 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0779 | 16.9771 | 3.7864 | 5.9242 | 6.2836 | 1.0559 | 0.22303 | 0.34895 | 0.37012 | 0.062197 | 1.5646 | 1.6595 | 0.27887 | 1.0607 | 0.17824 | 0.16804 | -1.5005 | -1.0528 | -0.99393 | -2.7775 | [10 empty] | 0 | 2.0779 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.01414 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 98.99% | 11.5787 | -5.6038 | 12.0157 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.79% | 12.0378 | 3.7885 | -8.139 | SRM 3169 | Avg | 0.21049 | -0.11461 | -0.85194
Lab 3 | Jul 21, 2020 | 2 | 7 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.702e-05; SAMPLE 49Ti40Ar: 9.0585e-06; SAMPLE 90Zr: 17.1907; SAMPLE 91Zr: 3.8342; SAMPLE 92Zr: 5.9992; SAMPLE 94Zr: 6.3635; SAMPLE 96Zr: 1.0695 | 1 | -0.00016839 | 1.7943e-05 | 17.1905 | 3.8341 | 5.9991 | 0 | 6.3633 | 0 | 1.0694 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0807 | 17.1905 | 3.8341 | 5.9991 | 6.3633 | 1.0694 | 0.22303 | 0.34898 | 0.37017 | 0.062208 | 1.5647 | 1.6597 | 0.27891 | 1.0607 | 0.17826 | 0.16805 | -1.5004 | -1.0527 | -0.9938 | -2.7773 | [10 empty] | 0 | 2.0807 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.014122 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.4% | -1.894 | 13.443 | -2.5636 | STD | Lab 3 | Jul 20, 2020 | 1 | 100.1% | -5.4416 | -4.5445 | -12.3535 | SRM 3169 | 2SD | 21.9135 | 15.2279 | 53.2724
Lab 3 | Jul 21, 2020 | 2 | 8 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 2.0023e-06; SAMPLE 49Ti40Ar: -7.2423e-06; SAMPLE 90Zr: 16.9202; SAMPLE 91Zr: 3.7739; SAMPLE 92Zr: 5.905; SAMPLE 94Zr: 6.2638; SAMPLE 96Zr: 1.0528 | 1 | -0.00014937 | 1.6418e-06 | 16.92 | 3.7738 | 5.9048 | 0 | 6.2636 | 0 | 1.0527 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0822 | 16.92 | 3.7738 | 5.9048 | 6.2636 | 1.0527 | 0.22304 | 0.34899 | 0.37019 | 0.062214 | 1.5647 | 1.6598 | 0.27894 | 1.0608 | 0.17827 | 0.16806 | -1.5004 | -1.0527 | -0.99374 | -2.7772 | [10 empty] | 0 | 2.0822 | 0.21795 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014132 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.4% | -15.0758 | 7.1453 | -5.9172 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.94% | -4.8502 | -3.7329 | 25.0672 | SRM 3169 | n | 77 | 77 | 77
Lab 3 | Jul 21, 2020 | 2 | 9 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 9.8101e-05; SAMPLE 49Ti40Ar: 9.676e-06; SAMPLE 90Zr: 16.4968; SAMPLE 91Zr: 3.6795; SAMPLE 92Zr: 5.7574; SAMPLE 94Zr: 6.1076; SAMPLE 96Zr: 1.0266 | 1 | -5.3269e-05 | 1.856e-05 | 16.4966 | 3.6794 | 5.7573 | 0 | 6.1074 | 0 | 1.0265 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0841 | 16.4966 | 3.6794 | 5.7573 | 6.1074 | 1.0265 | 0.22304 | 0.349 | 0.37022 | 0.062222 | 1.5647 | 1.6599 | 0.27897 | 1.0608 | 0.17829 | 0.16807 | -1.5004 | -1.0527 | -0.99366 | -2.777 | [10 empty] | 0 | 2.0841 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014137 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.05% | -6.6903 | -2.7174 | 10.7629 | STD | Lab 3 | Jul 20, 2020 | 1 | 100.2% | 9.7141 | 8.8534 | -16.4419 | SRM 3169 | 2SE | 2.4973 | 1.7354 | 6.071
Lab 3 | Jul 21, 2020 | 2 | 10 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -4.539e-06; SAMPLE 49Ti40Ar: 2.0763e-05; SAMPLE 90Zr: 16.7302; SAMPLE 91Zr: 3.7317; SAMPLE 92Zr: 5.839; SAMPLE 94Zr: 6.1944; SAMPLE 96Zr: 1.0412 | 1 | -0.00015591 | 2.9647e-05 | 16.73 | 3.7316 | 5.8389 | 0 | 6.1942 | 0 | 1.0411 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0856 | 16.73 | 3.7316 | 5.8389 | 6.1942 | 1.0411 | 0.22305 | 0.34901 | 0.37024 | 0.062228 | 1.5647 | 1.6599 | 0.27899 | 1.0608 | 0.1783 | 0.16807 | -1.5004 | -1.0527 | -0.99359 | -2.777 | [10 empty] | 0 | 2.0856 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014138 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.51% | 13.6884 | -7.3291 | 4.3174 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.76% | -7.0257 | -8.4807 | 6.0136 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 11 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.3643e-06; SAMPLE 49Ti40Ar: 2.6102e-05; SAMPLE 90Zr: 16.8659; SAMPLE 91Zr: 3.762; SAMPLE 92Zr: 5.8868; SAMPLE 94Zr: 6.2453; SAMPLE 96Zr: 1.0498 | 1 | -0.00015473 | 3.4987e-05 | 16.8657 | 3.762 | 5.8867 | 0 | 6.2451 | 0 | 1.0497 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0882 | 16.8657 | 3.762 | 5.8867 | 6.2451 | 1.0497 | 0.22305 | 0.34903 | 0.37029 | 0.062237 | 1.5648 | 1.6601 | 0.27902 | 1.0609 | 0.17831 | 0.16808 | -1.5003 | -1.0526 | -0.99348 | -2.7768 | [10 empty] | 0 | 2.0882 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014129 | -1.8133 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | 0.46612 | 4.4106 | -23.9164 | STD | Lab 3 | Jul 20, 2020 | 1 | 100.4% | -3.8247 | 8.6646 | -16.9887 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 12 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -7.0034e-06; SAMPLE 49Ti40Ar: 3.2804e-05; SAMPLE 90Zr: 16.7046; SAMPLE 91Zr: 3.7261; SAMPLE 92Zr: 5.8307; SAMPLE 94Zr: 6.1859; SAMPLE 96Zr: 1.0398 | 1 | -0.00015837 | 4.1688e-05 | 16.7044 | 3.726 | 5.8305 | 0 | 6.1857 | 0 | 1.0397 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0893 | 16.7044 | 3.726 | 5.8305 | 6.1857 | 1.0397 | 0.22306 | 0.34904 | 0.3703 | 0.062243 | 1.5648 | 1.6601 | 0.27905 | 1.0609 | 0.17832 | 0.16809 | -1.5003 | -1.0526 | -0.99343 | -2.7767 | [10 empty] | 0 | 2.0893 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014129 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 101.1% | -3.5286 | 11.1873 | 12.8108 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.66% | 14.1624 | -5.6479 | 11.5354 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 13 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 2.5958e-06; SAMPLE 49Ti40Ar: 2.693e-05; SAMPLE 90Zr: 16.6029; SAMPLE 91Zr: 3.7035; SAMPLE 92Zr: 5.7954; SAMPLE 94Zr: 6.1487; SAMPLE 96Zr: 1.0336 | 1 | -0.00014877 | 3.5814e-05 | 16.6027 | 3.7034 | 5.7952 | 0 | 6.1485 | 0 | 1.0335 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0911 | 16.6027 | 3.7034 | 5.7952 | 6.1485 | 1.0335 | 0.22306 | 0.34905 | 0.37033 | 0.062249 | 1.5648 | 1.6602 | 0.27907 | 1.061 | 0.17834 | 0.16809 | -1.5003 | -1.0525 | -0.99335 | -2.7766 | [10 empty] | 0 | 2.0911 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014134 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.6% | 8.7377 | 2.8308 | 0.8286 | STD | Lab 3 | Jul 20, 2020 | 1 | 100.1% | -8.4448 | 3.0889 | -0.97592 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 14 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.6527e-05; SAMPLE 49Ti40Ar: 2.3649e-05; SAMPLE 90Zr: 16.3032; SAMPLE 91Zr: 3.6367; SAMPLE 92Zr: 5.6909; SAMPLE 94Zr: 6.0381; SAMPLE 96Zr: 1.015 | 1 | -0.0001679 | 3.2533e-05 | 16.303 | 3.6366 | 5.6907 | 0 | 6.0379 | 0 | 1.0149 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0924 | 16.303 | 3.6366 | 5.6907 | 6.0379 | 1.0149 | 0.22306 | 0.34906 | 0.37035 | 0.062254 | 1.5648 | 1.6603 | 0.27909 | 1.061 | 0.17835 | 0.16809 | -1.5003 | -1.0525 | -0.99329 | -2.7765 | [10 empty] | 0 | 2.0924 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014143 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 98.72% | 2.0034 | -11.9192 | -18.2555 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.73% | -3.1548 | -5.2888 | 2.1347 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 15 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -2.3993e-05; SAMPLE 49Ti40Ar: 2.5386e-05; SAMPLE 90Zr: 16.5604; SAMPLE 91Zr: 3.6941; SAMPLE 92Zr: 5.781; SAMPLE 94Zr: 6.1339; SAMPLE 96Zr: 1.0312 | 1 | -0.00017536 | 3.427e-05 | 16.5601 | 3.6941 | 5.7808 | 0 | 6.1338 | 0 | 1.0311 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0948 | 16.5601 | 3.6941 | 5.7808 | 6.1338 | 1.0311 | 0.22307 | 0.34908 | 0.37039 | 0.062262 | 1.5649 | 1.6604 | 0.27911 | 1.061 | 0.17836 | 0.1681 | -1.5003 | -1.0524 | -0.99319 | -2.7764 | [10 empty] | 0 | 2.0948 | 0.21796 | 0.33336 | 0.3381 | 0.054377 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014132 | -1.8133 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.5% | 11.0165 | 6.3856 | -17.258 | STD | Lab 3 | Jul 20, 2020 | 1 | 100.3% | -5.8908 | -0.81499 | -15.5471 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 16 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 1.4108e-05; SAMPLE 49Ti40Ar: 5.7794e-05; SAMPLE 90Zr: 16.3246; SAMPLE 91Zr: 3.6415; SAMPLE 92Zr: 5.6988; SAMPLE 94Zr: 6.0469; SAMPLE 96Zr: 1.0166 | 1 | -0.00013726 | 6.6678e-05 | 16.3244 | 3.6415 | 5.6987 | 0 | 6.0468 | 0 | 1.0165 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.096 | 16.3244 | 3.6415 | 5.6987 | 6.0468 | 1.0165 | 0.22307 | 0.34909 | 0.37041 | 0.062269 | 1.5649 | 1.6605 | 0.27914 | 1.0611 | 0.17837 | 0.16811 | -1.5003 | -1.0524 | -0.99314 | -2.7763 | [10 empty] | 0 | 2.096 | 0.21795 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43906 | -1.3883 | 0.014134 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.7% | -7.773 | 0.98641 | 16.0038 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.76% | 18.2218 | 15.4158 | 36.1148 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 17 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 7.3007e-05; SAMPLE 49Ti40Ar: 3.6572e-06; SAMPLE 90Zr: 15.8378; SAMPLE 91Zr: 3.5331; SAMPLE 92Zr: 5.5291; SAMPLE 94Zr: 5.8672; SAMPLE 96Zr: 0.98644 | 1 | -7.8362e-05 | 1.2541e-05 | 15.8376 | 3.533 | 5.529 | 0 | 5.867 | 0 | 0.98634 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0982 | 15.8376 | 3.533 | 5.529 | 5.867 | 0.98634 | 0.22308 | 0.3491 | 0.37045 | 0.062278 | 1.565 | 1.6606 | 0.27918 | 1.0611 | 0.17839 | 0.16812 | -1.5002 | -1.0524 | -0.99304 | -2.7761 | [10 empty] | 0 | 2.0982 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014142 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.66% | -1.961 | -12.6449 | 25.1645 | STD | Lab 3 | Jul 20, 2020 | 1 | 99.96% | -16.2128 | -21.3256 | -41.9616 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 18 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 1.5018e-05; SAMPLE 49Ti40Ar: 3.5984e-05; SAMPLE 90Zr: 16.1101; SAMPLE 91Zr: 3.5939; SAMPLE 92Zr: 5.6244; SAMPLE 94Zr: 5.9684; SAMPLE 96Zr: 1.0035 | 1 | -0.00013635 | 4.4868e-05 | 16.1099 | 3.5938 | 5.6243 | 0 | 5.9682 | 0 | 1.0033 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.0996 | 16.1099 | 3.5938 | 5.6243 | 5.9682 | 1.0033 | 0.22308 | 0.34912 | 0.37047 | 0.062282 | 1.565 | 1.6607 | 0.27919 | 1.0612 | 0.1784 | 0.16811 | -1.5002 | -1.0523 | -0.99298 | -2.7761 | [10 empty] | 0 | 2.0996 | 0.21795 | 0.33336 | 0.3381 | 0.054377 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43906 | -1.3883 | 0.014126 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.65% | 7.018 | 11.86 | -50.2406 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 100.4% | 6.2448 | 15.0859 | 23.3313 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 19 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -6.3522e-06; SAMPLE 49Ti40Ar: 3.961e-05; SAMPLE 90Zr: 16.2116; SAMPLE 91Zr: 3.6166; SAMPLE 92Zr: 5.66; SAMPLE 94Zr: 6.0065; SAMPLE 96Zr: 1.0099 | 1 | -0.00015772 | 4.8494e-05 | 16.2113 | 3.6165 | 5.6599 | 0 | 6.0063 | 0 | 1.0098 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1015 | 16.2113 | 3.6165 | 5.6599 | 6.0063 | 1.0098 | 0.22309 | 0.34913 | 0.3705 | 0.06229 | 1.565 | 1.6608 | 0.27922 | 1.0612 | 0.17841 | 0.16812 | -1.5002 | -1.0523 | -0.9929 | -2.776 | [10 empty] | 0 | 2.1015 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.01414 | -1.8132 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | 5.0183 | -6.6714 | -28.0897 | STD | Lab 3 | Jul 20, 2020 | 1 | 99.8% | 0.2744 | -8.8044 | -6.0968 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 20 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.4932e-05; SAMPLE 49Ti40Ar: -1.1785e-05; SAMPLE 90Zr: 16.0093; SAMPLE 91Zr: 3.5715; SAMPLE 92Zr: 5.5895; SAMPLE 94Zr: 5.9318; SAMPLE 96Zr: 0.99743 | 1 | -0.0001663 | -2.9007e-06 | 16.0091 | 3.5714 | 5.5893 | 0 | 5.9316 | 0 | 0.99732 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1022 | 16.0091 | 3.5714 | 5.5893 | 5.9316 | 0.99732 | 0.22309 | 0.34914 | 0.37051 | 0.062297 | 1.565 | 1.6609 | 0.27925 | 1.0612 | 0.17843 | 0.16814 | -1.5002 | -1.0523 | -0.99287 | -2.7758 | [10 empty] | 0 | 2.1022 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5513 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42492 | 0.43906 | -1.3883 | 0.014141 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -7.9251 | -9.1266 | 54.7355 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.89% | 4.2681 | 6.6349 | 3.491 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 21 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 4.0959e-06; SAMPLE 49Ti40Ar: 4.2089e-05; SAMPLE 90Zr: 16.0759; SAMPLE 91Zr: 3.5864; SAMPLE 92Zr: 5.613; SAMPLE 94Zr: 5.957; SAMPLE 96Zr: 1.0016 | 1 | -0.00014727 | 5.0973e-05 | 16.0757 | 3.5863 | 5.6129 | 0 | 5.9568 | 0 | 1.0015 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1043 | 16.0757 | 3.5863 | 5.6129 | 5.9568 | 1.0015 | 0.22309 | 0.34915 | 0.37055 | 0.062302 | 1.5651 | 1.661 | 0.27927 | 1.0613 | 0.17844 | 0.16813 | -1.5002 | -1.0522 | -0.99278 | -2.7758 | [10 empty] | 0 | 2.1043 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014139 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.8% | -6.6617 | 1.6107 | -44.6196 | STD | Lab 3 | Jul 20, 2020 | 1 | 100.4% | -3.5453 | 0.65169 | 4.3975 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 22 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 2.3405e-05; SAMPLE 49Ti40Ar: 5.1188e-05; SAMPLE 90Zr: 15.8731; SAMPLE 91Zr: 3.5413; SAMPLE 92Zr: 5.5424; SAMPLE 94Zr: 5.8822; SAMPLE 96Zr: 0.98914 | 1 | -0.00012796 | 6.0072e-05 | 15.8728 | 3.5412 | 5.5422 | 0 | 5.882 | 0 | 0.98903 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1059 | 15.8728 | 3.5412 | 5.5422 | 5.882 | 0.98903 | 0.2231 | 0.34916 | 0.37057 | 0.06231 | 1.5651 | 1.661 | 0.2793 | 1.0613 | 0.17845 | 0.16815 | -1.5001 | -1.0522 | -0.99271 | -2.7756 | [10 empty] | 0 | 2.1059 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014139 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 10.1497 | 3.3934 | -15.5016 | [8 empty] | STD | Lab 3 | Jul 20, 2020 | 1 | 99.64% | -2.5118 | -8.3052 | -18.5897 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 23 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.2489e-05; SAMPLE 49Ti40Ar: 9.0391e-06; SAMPLE 90Zr: 15.9818; SAMPLE 91Zr: 3.5655; SAMPLE 92Zr: 5.5806; SAMPLE 94Zr: 5.9231; SAMPLE 96Zr: 0.99605 | 1 | -0.00018386 | 1.7923e-05 | 15.9816 | 3.5655 | 5.5805 | 0 | 5.9229 | 0 | 0.99594 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1081 | 15.9816 | 3.5655 | 5.5805 | 5.9229 | 0.99594 | 0.2231 | 0.34918 | 0.37061 | 0.062318 | 1.5652 | 1.6612 | 0.27933 | 1.0614 | 0.17847 | 0.16815 | -1.5001 | -1.0522 | -0.99261 | -2.7755 | [10 empty] | 0 | 2.1081 | 0.21795 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43907 | -1.3883 | 0.014132 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.6% | -9.7713 | 8.7516 | -6.3123 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.9% | -10.3757 | -1.3011 | -13.101 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 24 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.2429e-05; SAMPLE 49Ti40Ar: 4.3002e-05; SAMPLE 90Zr: 15.5889; SAMPLE 91Zr: 3.4779; SAMPLE 92Zr: 5.4435; SAMPLE 94Zr: 5.7777; SAMPLE 96Zr: 0.97163 | 1 | -0.0001838 | 5.1886e-05 | 15.5887 | 3.4779 | 5.4434 | 0 | 5.7775 | 0 | 0.97152 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.109 | 15.5887 | 3.4779 | 5.4434 | 5.7775 | 0.97152 | 0.2231 | 0.34919 | 0.37062 | 0.062322 | 1.5651 | 1.6612 | 0.27934 | 1.0614 | 0.17848 | 0.16816 | -1.5001 | -1.0521 | -0.99257 | -2.7754 | [10 empty] | 0 | 2.109 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014143 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.48% | -3.9196 | 1.0671 | 15.8956 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 98.99% | 11.5787 | -5.6038 | 12.0157 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 25 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 8.1436e-05; SAMPLE 49Ti40Ar: 6.7455e-06; SAMPLE 90Zr: 15.1597; SAMPLE 91Zr: 3.3822; SAMPLE 92Zr: 5.2938; SAMPLE 94Zr: 5.619; SAMPLE 96Zr: 0.94499 | 1 | -6.9934e-05 | 1.563e-05 | 15.1595 | 3.3822 | 5.2937 | 0 | 5.6188 | 0 | 0.94488 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1106 | 15.1595 | 3.3822 | 5.2937 | 5.6188 | 0.94488 | 0.2231 | 0.3492 | 0.37065 | 0.06233 | 1.5652 | 1.6613 | 0.27937 | 1.0614 | 0.17849 | 0.16816 | -1.5001 | -1.0521 | -0.9925 | -2.7753 | [10 empty] | 0 | 2.1106 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42493 | 0.43907 | -1.3883 | 0.014139 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.63% | -4.8627 | 7.8351 | 25.0743 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.4% | -15.0758 | 7.1453 | -5.9172 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 26 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.1824e-06; SAMPLE 49Ti40Ar: 1.0043e-05; SAMPLE 90Zr: 15.4666; SAMPLE 91Zr: 3.4508; SAMPLE 92Zr: 5.4013; SAMPLE 94Zr: 5.7334; SAMPLE 96Zr: 0.96424 | 1 | -0.00015255 | 1.8927e-05 | 15.4664 | 3.4507 | 5.4011 | 0 | 5.7332 | 0 | 0.96413 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1132 | 15.4664 | 3.4507 | 5.4011 | 5.7332 | 0.96413 | 0.22311 | 0.34922 | 0.37069 | 0.062337 | 1.5652 | 1.6615 | 0.2794 | 1.0615 | 0.17851 | 0.16817 | -1.5001 | -1.0521 | -0.99239 | -2.7752 | [10 empty] | 0 | 2.1132 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.24949 | 1.0143 | 0.16313 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.01415 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100% | -6.1907 | -7.928 | -28.8386 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.51% | 13.6884 | -7.3291 | 4.3174 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 27 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.7315e-05; SAMPLE 49Ti40Ar: -5.811e-06; SAMPLE 90Zr: 15.4324; SAMPLE 91Zr: 3.4432; SAMPLE 92Zr: 5.3896; SAMPLE 94Zr: 5.7211; SAMPLE 96Zr: 0.96224 | 1 | -0.00016868 | 3.0731e-06 | 15.4322 | 3.4432 | 5.3894 | 0 | 5.7209 | 0 | 0.96213 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1147 | 15.4322 | 3.4432 | 5.3894 | 5.7209 | 0.96213 | 0.22312 | 0.34923 | 0.37071 | 0.062346 | 1.5653 | 1.6615 | 0.27943 | 1.0615 | 0.17852 | 0.16818 | -1.5001 | -1.052 | -0.99232 | -2.7751 | [10 empty] | 0 | 2.1147 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42493 | 0.43907 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.2% | -5.1419 | 5.6344 | 27.1017 | STD | Lab 3 | Jul 21, 2020 | 2 | 101.1% | -3.5286 | 11.1873 | 12.8108 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 28 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -5.4733e-05; SAMPLE 49Ti40Ar: 2.1649e-05; SAMPLE 90Zr: 15.3362; SAMPLE 91Zr: 3.4219; SAMPLE 92Zr: 5.3562; SAMPLE 94Zr: 5.6859; SAMPLE 96Zr: 0.95634 | 1 | -0.0002061 | 3.0533e-05 | 15.336 | 3.4218 | 5.356 | 0 | 5.6857 | 0 | 0.95623 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1165 | 15.336 | 3.4218 | 5.356 | 5.6857 | 0.95623 | 0.22312 | 0.34925 | 0.37074 | 0.062352 | 1.5653 | 1.6616 | 0.27945 | 1.0616 | 0.17853 | 0.16818 | -1.5 | -1.052 | -0.99225 | -2.775 | [10 empty] | 0 | 2.1165 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.01414 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -4.2894 | 1.9865 | 15.2521 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 98.72% | 2.0034 | -11.9192 | -18.2555 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 29 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.4046e-05; SAMPLE 49Ti40Ar: 2.9311e-05; SAMPLE 90Zr: 15.3156; SAMPLE 91Zr: 3.4173; SAMPLE 92Zr: 5.3492; SAMPLE 94Zr: 5.6787; SAMPLE 96Zr: 0.95513 | 1 | -0.00016542 | 3.8196e-05 | 15.3154 | 3.4172 | 5.349 | 0 | 5.6785 | 0 | 0.95502 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1181 | 15.3154 | 3.4172 | 5.349 | 5.6785 | 0.95502 | 0.22312 | 0.34926 | 0.37077 | 0.062357 | 1.5653 | 1.6617 | 0.27947 | 1.0616 | 0.17854 | 0.16818 | -1.5 | -1.0519 | -0.99218 | -2.7749 | [10 empty] | 0 | 2.1181 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42493 | 0.43907 | -1.3883 | 0.014141 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.4% | -22.3547 | -3.2295 | -25.9543 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.7% | -7.773 | 0.98641 | 16.0038 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 30 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.5891e-05; SAMPLE 49Ti40Ar: 4.6151e-05; SAMPLE 90Zr: 15.166; SAMPLE 91Zr: 3.3841; SAMPLE 92Zr: 5.2971; SAMPLE 94Zr: 5.6235; SAMPLE 96Zr: 0.94591 | 1 | -0.00016726 | 5.5035e-05 | 15.1658 | 3.384 | 5.297 | 0 | 5.6233 | 0 | 0.9458 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1195 | 15.1658 | 3.384 | 5.297 | 5.6233 | 0.9458 | 0.22313 | 0.34927 | 0.37079 | 0.062364 | 1.5653 | 1.6617 | 0.27949 | 1.0616 | 0.17855 | 0.16819 | -1.5 | -1.0519 | -0.99211 | -2.7748 | [10 empty] | 0 | 2.1195 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014135 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.63% | 17.6851 | 4.7225 | -9.9209 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.65% | 7.018 | 11.86 | -50.2406 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 31 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.9623e-05; SAMPLE 49Ti40Ar: 1.4868e-05; SAMPLE 90Zr: 15.3118; SAMPLE 91Zr: 3.4166; SAMPLE 92Zr: 5.3482; SAMPLE 94Zr: 5.6781; SAMPLE 96Zr: 0.95511 | 1 | -0.00017099 | 2.3752e-05 | 15.3116 | 3.4165 | 5.3481 | 0 | 5.6779 | 0 | 0.955 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1215 | 15.3116 | 3.4165 | 5.3481 | 5.6779 | 0.955 | 0.22313 | 0.34928 | 0.37082 | 0.062371 | 1.5654 | 1.6619 | 0.27952 | 1.0617 | 0.17857 | 0.1682 | -1.5 | -1.0519 | -0.99203 | -2.7747 | [10 empty] | 0 | 2.1215 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.014144 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.2% | -18.6456 | -6.5614 | -22.9547 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -7.9251 | -9.1266 | 54.7355 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 32 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.8254e-05; SAMPLE 49Ti40Ar: 3.3001e-05; SAMPLE 90Zr: 15.1087; SAMPLE 91Zr: 3.3714; SAMPLE 92Zr: 5.2775; SAMPLE 94Zr: 5.6031; SAMPLE 96Zr: 0.94256 | 1 | -0.00016962 | 4.1885e-05 | 15.1085 | 3.3713 | 5.2773 | 0 | 5.6029 | 0 | 0.94245 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1229 | 15.1085 | 3.3713 | 5.2773 | 5.6029 | 0.94245 | 0.22314 | 0.3493 | 0.37085 | 0.062379 | 1.5654 | 1.662 | 0.27955 | 1.0617 | 0.17858 | 0.16821 | -1.5 | -1.0518 | -0.99197 | -2.7745 | [10 empty] | 0 | 2.1229 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014139 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | -9.7824 | 6.0333 | 0.21916 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 10.1497 | 3.3934 | -15.5016 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 33 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 9.8266e-05; SAMPLE 49Ti40Ar: 4.352e-06; SAMPLE 90Zr: 14.6874; SAMPLE 91Zr: 3.2774; SAMPLE 92Zr: 5.1305; SAMPLE 94Zr: 5.4472; SAMPLE 96Zr: 0.91635 | 1 | -5.3104e-05 | 1.3236e-05 | 14.6872 | 3.2773 | 5.1304 | 0 | 5.447 | 0 | 0.91624 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1244 | 14.6872 | 3.2773 | 5.1304 | 5.447 | 0.91624 | 0.22314 | 0.34931 | 0.37087 | 0.062384 | 1.5654 | 1.6621 | 0.27957 | 1.0617 | 0.17859 | 0.16821 | -1.5 | -1.0518 | -0.9919 | -2.7745 | [10 empty] | 0 | 2.1244 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43906 | -1.3883 | 0.014134 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.86% | -9.752 | 12.6985 | -22.1612 | STD | Lab 3 | Jul 21, 2020 | 2 | 99.48% | -3.9196 | 1.0671 | 15.8956 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 34 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -4.9804e-05; SAMPLE 49Ti40Ar: 9.2142e-07; SAMPLE 90Zr: 14.9089; SAMPLE 91Zr: 3.3269; SAMPLE 92Zr: 5.2079; SAMPLE 94Zr: 5.5297; SAMPLE 96Zr: 0.93028 | 1 | -0.00020117 | 9.8055e-06 | 14.9087 | 3.3268 | 5.2078 | 0 | 5.5295 | 0 | 0.93017 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1258 | 14.9087 | 3.3268 | 5.2078 | 5.5295 | 0.93017 | 0.22315 | 0.34931 | 0.37089 | 0.062391 | 1.5654 | 1.6621 | 0.2796 | 1.0618 | 0.17861 | 0.16822 | -1.4999 | -1.0518 | -0.99184 | -2.7743 | [10 empty] | 0 | 2.1258 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43905 | -1.3883 | 0.014155 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.27% | 3.9293 | -9.1001 | 5.3902 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100% | -6.1907 | -7.928 | -28.8386 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 35 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -4.3933e-06; SAMPLE 49Ti40Ar: 1.3687e-05; SAMPLE 90Zr: 14.9976; SAMPLE 91Zr: 3.3467; SAMPLE 92Zr: 5.2392; SAMPLE 94Zr: 5.563; SAMPLE 96Zr: 0.93592 | 1 | -0.00015576 | 2.2571e-05 | 14.9973 | 3.3466 | 5.239 | 0 | 5.5629 | 0 | 0.93581 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1276 | 14.9973 | 3.3466 | 5.239 | 5.5629 | 0.93581 | 0.22315 | 0.34933 | 0.37092 | 0.062398 | 1.5655 | 1.6622 | 0.27963 | 1.0618 | 0.17862 | 0.16822 | -1.4999 | -1.0517 | -0.99176 | -2.7742 | [10 empty] | 0 | 2.1276 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.014144 | -1.8132 | -1.8273 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.5% | -1.6577 | 10.2111 | -4.4936 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -4.2894 | 1.9865 | 15.2521 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 36 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.4721e-05; SAMPLE 49Ti40Ar: 2.24e-05; SAMPLE 90Zr: 14.9286; SAMPLE 91Zr: 3.3314; SAMPLE 92Zr: 5.2151; SAMPLE 94Zr: 5.5377; SAMPLE 96Zr: 0.93168 | 1 | -0.00016609 | 3.1285e-05 | 14.9284 | 3.3313 | 5.215 | 0 | 5.5375 | 0 | 0.93157 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1287 | 14.9284 | 3.3313 | 5.215 | 5.5375 | 0.93157 | 0.22315 | 0.34934 | 0.37094 | 0.062403 | 1.5655 | 1.6623 | 0.27964 | 1.0618 | 0.17863 | 0.16823 | -1.4999 | -1.0517 | -0.99171 | -2.7741 | [10 empty] | 0 | 2.1287 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.24951 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42491 | 0.43906 | -1.3883 | 0.014153 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.7% | -5.2268 | -2.984 | 11.4765 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.63% | 17.6851 | 4.7225 | -9.9209 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 37 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -2.5766e-05; SAMPLE 49Ti40Ar: 4.3567e-05; SAMPLE 90Zr: 14.8816; SAMPLE 91Zr: 3.321; SAMPLE 92Zr: 5.1989; SAMPLE 94Zr: 5.5207; SAMPLE 96Zr: 0.92881 | 1 | -0.00017714 | 5.2451e-05 | 14.8813 | 3.3209 | 5.1988 | 0 | 5.5205 | 0 | 0.9287 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1306 | 14.8813 | 3.3209 | 5.1988 | 5.5205 | 0.9287 | 0.22316 | 0.34935 | 0.37097 | 0.062407 | 1.5655 | 1.6624 | 0.27965 | 1.0619 | 0.17864 | 0.16823 | -1.4999 | -1.0517 | -0.99163 | -2.7741 | [10 empty] | 0 | 2.1306 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5294 | 1.5512 | 0.24949 | 1.0143 | 0.16313 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014155 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.3% | 6.7257 | -5.7961 | -51.8264 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | -9.7824 | 6.0333 | 0.21916 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 38 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.327e-05; SAMPLE 49Ti40Ar: 1.6487e-05; SAMPLE 90Zr: 14.7416; SAMPLE 91Zr: 3.2898; SAMPLE 92Zr: 5.1502; SAMPLE 94Zr: 5.4691; SAMPLE 96Zr: 0.92019 | 1 | -0.00018464 | 2.5371e-05 | 14.7414 | 3.2897 | 5.1501 | 0 | 5.469 | 0 | 0.92008 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1319 | 14.7414 | 3.2897 | 5.1501 | 5.469 | 0.92008 | 0.22316 | 0.34936 | 0.37099 | 0.062415 | 1.5655 | 1.6624 | 0.27969 | 1.0619 | 0.17865 | 0.16824 | -1.4999 | -1.0516 | -0.99157 | -2.774 | [10 empty] | 0 | 2.1319 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014145 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.98% | 5.7518 | 2.8457 | 10.5227 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.27% | 3.9293 | -9.1001 | 5.3902 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 39 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -5.969e-06; SAMPLE 49Ti40Ar: 2.7798e-05; SAMPLE 90Zr: 14.813; SAMPLE 91Zr: 3.3057; SAMPLE 92Zr: 5.1754; SAMPLE 94Zr: 5.496; SAMPLE 96Zr: 0.92471 | 1 | -0.00015734 | 3.6682e-05 | 14.8128 | 3.3057 | 5.1752 | 0 | 5.4958 | 0 | 0.9246 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1333 | 14.8128 | 3.3057 | 5.1752 | 5.4958 | 0.9246 | 0.22316 | 0.34938 | 0.37101 | 0.062419 | 1.5656 | 1.6625 | 0.2797 | 1.0619 | 0.17866 | 0.16824 | -1.4999 | -1.0516 | -0.99151 | -2.7739 | [10 empty] | 0 | 2.1333 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42493 | 0.43906 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.1% | -5.5718 | 6.7614 | -5.7334 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.7% | -5.2268 | -2.984 | 11.4765 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 40 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -5.904e-05; SAMPLE 49Ti40Ar: 1.2433e-05; SAMPLE 90Zr: 14.5617; SAMPLE 91Zr: 3.2497; SAMPLE 92Zr: 5.0877; SAMPLE 94Zr: 5.4031; SAMPLE 96Zr: 0.9091 | 1 | -0.00021041 | 2.1317e-05 | 14.5615 | 3.2496 | 5.0876 | 0 | 5.4029 | 0 | 0.90899 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1349 | 14.5615 | 3.2496 | 5.0876 | 5.4029 | 0.90899 | 0.22317 | 0.34939 | 0.37104 | 0.062425 | 1.5656 | 1.6626 | 0.27972 | 1.062 | 0.17867 | 0.16824 | -1.4998 | -1.0516 | -0.99144 | -2.7738 | [10 empty] | 0 | 2.1349 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014143 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100% | -5.9921 | 4.367 | -32.9611 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.98% | 5.7518 | 2.8457 | 10.5227 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 41 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 6.3033e-05; SAMPLE 49Ti40Ar: 1.3028e-05; SAMPLE 90Zr: 14.1688; SAMPLE 91Zr: 3.1621; SAMPLE 92Zr: 4.9506; SAMPLE 94Zr: 5.2578; SAMPLE 96Zr: 0.8847 | 1 | -8.8336e-05 | 2.1912e-05 | 14.1686 | 3.162 | 4.9505 | 0 | 5.2576 | 0 | 0.88459 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1369 | 14.1686 | 3.162 | 4.9505 | 5.2576 | 0.88459 | 0.22317 | 0.3494 | 0.37107 | 0.062434 | 1.5656 | 1.6627 | 0.27976 | 1.062 | 0.17869 | 0.16825 | -1.4998 | -1.0515 | -0.99136 | -2.7737 | [10 empty] | 0 | 2.1369 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.014156 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 97.92% | -14.6631 | -9.8627 | 3.4383 | STD | Lab 3 | Jul 21, 2020 | 2 | 100% | -5.9921 | 4.367 | -32.9611 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 42 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -4.4323e-05; SAMPLE 49Ti40Ar: -7.9018e-07; SAMPLE 90Zr: 14.3783; SAMPLE 91Zr: 3.2089; SAMPLE 92Zr: 5.024; SAMPLE 94Zr: 5.3358; SAMPLE 96Zr: 0.89787 | 1 | -0.00019569 | 8.0939e-06 | 14.378 | 3.2088 | 5.0238 | 0 | 5.3356 | 0 | 0.89776 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1382 | 14.378 | 3.2088 | 5.0238 | 5.3356 | 0.89776 | 0.22318 | 0.34941 | 0.37109 | 0.062439 | 1.5656 | 1.6628 | 0.27978 | 1.0621 | 0.1787 | 0.16826 | -1.4998 | -1.0515 | -0.9913 | -2.7736 | [10 empty] | 0 | 2.1382 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.014149 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.52% | 9.08 | 0.86551 | 4.9676 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.52% | 9.08 | 0.86551 | 4.9676 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 43 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -2.1865e-05; SAMPLE 49Ti40Ar: 2.6277e-05; SAMPLE 90Zr: 14.4123; SAMPLE 91Zr: 3.2166; SAMPLE 92Zr: 5.0361; SAMPLE 94Zr: 5.3488; SAMPLE 96Zr: 0.90009 | 1 | -0.00017324 | 3.5161e-05 | 14.4121 | 3.2165 | 5.0359 | 0 | 5.3486 | 0 | 0.89998 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1397 | 14.4121 | 3.2165 | 5.0359 | 5.3486 | 0.89998 | 0.22318 | 0.34942 | 0.37112 | 0.062446 | 1.5657 | 1.6629 | 0.2798 | 1.0621 | 0.17871 | 0.16827 | -1.4998 | -1.0515 | -0.99123 | -2.7734 | [10 empty] | 0 | 2.1397 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014139 | -1.8132 | -1.8273 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.4% | 7.7493 | 14.335 | 7.543 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.6% | -3.1159 | -5.9507 | 29.2064 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 44 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.6325e-06; SAMPLE 49Ti40Ar: 4.2759e-05; SAMPLE 90Zr: 14.3344; SAMPLE 91Zr: 3.1992; SAMPLE 92Zr: 5.0089; SAMPLE 94Zr: 5.3201; SAMPLE 96Zr: 0.8953 | 1 | -0.000153 | 5.1643e-05 | 14.3342 | 3.1991 | 5.0088 | 0 | 5.3199 | 0 | 0.89519 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1409 | 14.3342 | 3.1991 | 5.0088 | 5.3199 | 0.89519 | 0.22318 | 0.34943 | 0.37114 | 0.062451 | 1.5657 | 1.663 | 0.27983 | 1.0621 | 0.17872 | 0.16827 | -1.4998 | -1.0515 | -0.99118 | -2.7734 | [10 empty] | 0 | 2.1409 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43907 | -1.3883 | 0.014157 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.6% | -3.1159 | -5.9507 | 29.2064 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.83% | -13.3364 | 1.6431 | -42.1054 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 45 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -2.4574e-05; SAMPLE 49Ti40Ar: 2.1282e-05; SAMPLE 90Zr: 14.2648; SAMPLE 91Zr: 3.1837; SAMPLE 92Zr: 4.9848; SAMPLE 94Zr: 5.2947; SAMPLE 96Zr: 0.89104 | 1 | -0.00017594 | 3.0166e-05 | 14.2646 | 3.1836 | 4.9847 | 0 | 5.2945 | 0 | 0.89093 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1427 | 14.2646 | 3.1836 | 4.9847 | 5.2945 | 0.89093 | 0.22318 | 0.34944 | 0.37117 | 0.062458 | 1.5657 | 1.6631 | 0.27985 | 1.0622 | 0.17873 | 0.16827 | -1.4998 | -1.0514 | -0.9911 | -2.7733 | [10 empty] | 0 | 2.1427 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.01415 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100.3% | 2.8183 | 5.2652 | 5.0518 | STD | Lab 3 | Jul 21, 2020 | 2 | 99.7% | 10.7546 | -4.2134 | 15.1947 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 46 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -6.0599e-05; SAMPLE 49Ti40Ar: 1.9321e-06; SAMPLE 90Zr: 14.1176; SAMPLE 91Zr: 3.1509; SAMPLE 92Zr: 4.9335; SAMPLE 94Zr: 5.2404; SAMPLE 96Zr: 0.8819 | 1 | -0.00021197 | 1.0816e-05 | 14.1174 | 3.1508 | 4.9334 | 0 | 5.2402 | 0 | 0.88179 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.144 | 14.1174 | 3.1508 | 4.9334 | 5.2402 | 0.88179 | 0.22319 | 0.34945 | 0.37119 | 0.062461 | 1.5657 | 1.6631 | 0.27986 | 1.0622 | 0.17874 | 0.16827 | -1.4998 | -1.0514 | -0.99105 | -2.7732 | [10 empty] | 0 | 2.144 | 0.21795 | 0.33335 | 0.3381 | 0.054377 | 1.5295 | 1.5513 | 0.24949 | 1.0143 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43908 | -1.3883 | 0.014153 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.83% | -13.3364 | 1.6431 | -42.1054 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.72% | 5.1691 | 4.6409 | 17.1469 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 47 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -5.6305e-05; SAMPLE 49Ti40Ar: 7.7919e-06; SAMPLE 90Zr: 14.1523; SAMPLE 91Zr: 3.1587; SAMPLE 92Zr: 4.9459; SAMPLE 94Zr: 5.2536; SAMPLE 96Zr: 0.88415 | 1 | -0.00020768 | 1.6676e-05 | 14.1521 | 3.1586 | 4.9457 | 0 | 5.2534 | 0 | 0.88404 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1454 | 14.1521 | 3.1586 | 4.9457 | 5.2534 | 0.88404 | 0.22319 | 0.34947 | 0.37121 | 0.062467 | 1.5658 | 1.6632 | 0.27988 | 1.0622 | 0.17875 | 0.16828 | -1.4997 | -1.0513 | -0.99099 | -2.7731 | [10 empty] | 0 | 2.1454 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42493 | 0.43907 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 100.9% | -3.6032 | 16.4609 | -9.2216 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.6% | -6.4231 | 2.019 | -16.881 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 48 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.774e-05; SAMPLE 49Ti40Ar: 2.0416e-05; SAMPLE 90Zr: 13.9476; SAMPLE 91Zr: 3.1131; SAMPLE 92Zr: 4.8744; SAMPLE 94Zr: 5.1779; SAMPLE 96Zr: 0.87145 | 1 | -0.00016911 | 2.93e-05 | 13.9474 | 3.113 | 4.8742 | 0 | 5.1777 | 0 | 0.87134 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1465 | 13.9474 | 3.113 | 4.8742 | 5.1777 | 0.87134 | 0.2232 | 0.34947 | 0.37123 | 0.062474 | 1.5658 | 1.6632 | 0.2799 | 1.0623 | 0.17877 | 0.16829 | -1.4997 | -1.0513 | -0.99094 | -2.773 | [10 empty] | 0 | 2.1465 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43906 | -1.3883 | 0.014153 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.7% | 10.7546 | -4.2134 | 15.1947 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.89% | -0.89344 | -2.7324 | 14.5291 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 49 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 0.00011432; SAMPLE 49Ti40Ar: 2.1689e-05; SAMPLE 90Zr: 13.6538; SAMPLE 91Zr: 3.0475; SAMPLE 92Zr: 4.7719; SAMPLE 94Zr: 5.0691; SAMPLE 96Zr: 0.85318 | 1 | -3.7051e-05 | 3.0573e-05 | 13.6535 | 3.0474 | 4.7717 | 0 | 5.0689 | 0 | 0.85307 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1482 | 13.6535 | 3.0474 | 4.7717 | 5.0689 | 0.85307 | 0.2232 | 0.34949 | 0.37125 | 0.06248 | 1.5658 | 1.6634 | 0.27993 | 1.0623 | 0.17878 | 0.16829 | -1.4997 | -1.0513 | -0.99087 | -2.7729 | [10 empty] | 0 | 2.1482 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42493 | 0.43907 | -1.3883 | 0.014149 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.19% | -18.6254 | -1.0449 | -10.4323 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -2.9933 | 9.0597 | -20.6788 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 50 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -4.7706e-05; SAMPLE 49Ti40Ar: 2.2614e-05; SAMPLE 90Zr: 13.8619; SAMPLE 91Zr: 3.094; SAMPLE 92Zr: 4.8448; SAMPLE 94Zr: 5.1467; SAMPLE 96Zr: 0.86627 | 1 | -0.00019908 | 3.1498e-05 | 13.8617 | 3.094 | 4.8446 | 0 | 5.1465 | 0 | 0.86616 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1494 | 13.8617 | 3.094 | 4.8446 | 5.1465 | 0.86616 | 0.2232 | 0.3495 | 0.37128 | 0.062486 | 1.5658 | 1.6634 | 0.27995 | 1.0623 | 0.17879 | 0.1683 | -1.4997 | -1.0513 | -0.99081 | -2.7728 | [10 empty] | 0 | 2.1494 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.014144 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.72% | 5.1691 | 4.6409 | 17.1469 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.56% | 5.4528 | -10.4399 | 10.4885 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 51 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.566e-05; SAMPLE 49Ti40Ar: 1.6705e-05; SAMPLE 90Zr: 13.9342; SAMPLE 91Zr: 3.1102; SAMPLE 92Zr: 4.8702; SAMPLE 94Zr: 5.1739; SAMPLE 96Zr: 0.8709 | 1 | -0.00018703 | 2.5589e-05 | 13.934 | 3.1102 | 4.8701 | 0 | 5.1737 | 0 | 0.87079 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1511 | 13.934 | 3.1102 | 4.8701 | 5.1737 | 0.87079 | 0.22321 | 0.34951 | 0.3713 | 0.062494 | 1.5659 | 1.6635 | 0.27998 | 1.0623 | 0.1788 | 0.16831 | -1.4997 | -1.0512 | -0.99074 | -2.7727 | [10 empty] | 0 | 2.1511 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42492 | 0.43906 | -1.3883 | 0.014139 | -1.8132 | -1.8273 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.6% | 2.4594 | 4.838 | 29.3741 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 6.6744 | -2.6477 | -9.5758 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 52 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.7904e-05; SAMPLE 49Ti40Ar: 1.1633e-05; SAMPLE 90Zr: 13.8528; SAMPLE 91Zr: 3.0921; SAMPLE 92Zr: 4.8419; SAMPLE 94Zr: 5.144; SAMPLE 96Zr: 0.86585 | 1 | -0.00018927 | 2.0518e-05 | 13.8526 | 3.092 | 4.8418 | 0 | 5.1438 | 0 | 0.86574 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1524 | 13.8526 | 3.092 | 4.8418 | 5.1438 | 0.86574 | 0.22321 | 0.34952 | 0.37132 | 0.062497 | 1.5659 | 1.6636 | 0.27999 | 1.0624 | 0.17881 | 0.16831 | -1.4997 | -1.0512 | -0.99068 | -2.7726 | [10 empty] | 0 | 2.1524 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.014144 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.6% | -6.4231 | 2.019 | -16.881 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.94% | -11.7 | 4.7065 | 26.0442 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 53 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 8.0434e-06; SAMPLE 49Ti40Ar: 4.1279e-05; SAMPLE 90Zr: 13.6622; SAMPLE 91Zr: 3.0496; SAMPLE 92Zr: 4.7753; SAMPLE 94Zr: 5.0734; SAMPLE 96Zr: 0.85397 | 1 | -0.00014333 | 5.0163e-05 | 13.6619 | 3.0495 | 4.7752 | 0 | 5.0732 | 0 | 0.85386 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1535 | 13.6619 | 3.0495 | 4.7752 | 5.0732 | 0.85386 | 0.22321 | 0.34953 | 0.37134 | 0.062499 | 1.5659 | 1.6636 | 0.28 | 1.0624 | 0.17881 | 0.16831 | -1.4996 | -1.0512 | -0.99063 | -2.7726 | [10 empty] | 0 | 2.1535 | 0.21795 | 0.33335 | 0.3381 | 0.054377 | 1.5295 | 1.5513 | 0.24949 | 1.0143 | 0.16313 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43907 | -1.3883 | 0.014154 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 99.23% | 2.4161 | -7.9842 | -33.4991 | STD | Lab 3 | Jul 21, 2020 | 2 | 100% | 10.586 | -1.3842 | -9.5796 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 54 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -2.343e-05; SAMPLE 49Ti40Ar: 2.0533e-05; SAMPLE 90Zr: 13.6835; SAMPLE 91Zr: 3.0544; SAMPLE 92Zr: 4.783; SAMPLE 94Zr: 5.0817; SAMPLE 96Zr: 0.85542 | 1 | -0.0001748 | 2.9418e-05 | 13.6833 | 3.0543 | 4.7828 | 0 | 5.0815 | 0 | 0.85531 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.155 | 13.6833 | 3.0543 | 4.7828 | 5.0815 | 0.85531 | 0.22321 | 0.34954 | 0.37137 | 0.062508 | 1.5659 | 1.6637 | 0.28003 | 1.0624 | 0.17883 | 0.16832 | -1.4996 | -1.0511 | -0.99057 | -2.7725 | [10 empty] | 0 | 2.155 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.014149 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.89% | -0.89344 | -2.7324 | 14.5291 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.39% | -8.7199 | 3.4833 | -23.0128 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 55 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -6.1352e-05; SAMPLE 49Ti40Ar: 1.7886e-05; SAMPLE 90Zr: 13.7194; SAMPLE 91Zr: 3.0624; SAMPLE 92Zr: 4.7957; SAMPLE 94Zr: 5.0954; SAMPLE 96Zr: 0.85772 | 1 | -0.00021272 | 2.6771e-05 | 13.7191 | 3.0624 | 4.7956 | 0 | 5.0952 | 0 | 0.85761 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1567 | 13.7191 | 3.0624 | 4.7956 | 5.0952 | 0.85761 | 0.22322 | 0.34955 | 0.37139 | 0.062512 | 1.566 | 1.6638 | 0.28005 | 1.0625 | 0.17883 | 0.16832 | -1.4996 | -1.0511 | -0.9905 | -2.7724 | [10 empty] | 0 | 2.1567 | 0.21795 | 0.33335 | 0.3381 | 0.054377 | 1.5295 | 1.5513 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.014145 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 100.8% | 0.73283 | 3.3364 | -30.8964 | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 0.13293 | -0.3846 | 17.7775 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 56 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.5326e-05; SAMPLE 49Ti40Ar: 5.5699e-06; SAMPLE 90Zr: 13.5453; SAMPLE 91Zr: 3.0236; SAMPLE 92Zr: 4.7349; SAMPLE 94Zr: 5.0309; SAMPLE 96Zr: 0.84691 | 1 | -0.0001867 | 1.4454e-05 | 13.545 | 3.0235 | 4.7348 | 0 | 5.0307 | 0 | 0.8468 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1577 | 13.545 | 3.0235 | 4.7348 | 5.0307 | 0.8468 | 0.22322 | 0.34956 | 0.37141 | 0.062517 | 1.566 | 1.6639 | 0.28007 | 1.0625 | 0.17885 | 0.16833 | -1.4996 | -1.0511 | -0.99045 | -2.7723 | [10 empty] | 0 | 2.1577 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014148 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | -2.9933 | 9.0597 | -20.6788 | [8 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | 2.4783 | -5.8966 | -12.2242 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 57 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 5.9546e-05; SAMPLE 49Ti40Ar: 2.6037e-05; SAMPLE 90Zr: 13.1949; SAMPLE 91Zr: 2.9454; SAMPLE 92Zr: 4.6126; SAMPLE 94Zr: 4.9011; SAMPLE 96Zr: 0.82509 | 1 | -9.1824e-05 | 3.4921e-05 | 13.1946 | 2.9453 | 4.6125 | 0 | 4.9009 | 0 | 0.82499 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1591 | 13.1946 | 2.9453 | 4.6125 | 4.9009 | 0.82499 | 0.22322 | 0.34957 | 0.37143 | 0.062524 | 1.566 | 1.664 | 0.2801 | 1.0625 | 0.17886 | 0.16833 | -1.4996 | -1.051 | -0.99039 | -2.7722 | [10 empty] | 0 | 2.1591 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42494 | 0.43908 | -1.3883 | 0.014146 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.01% | -23.3863 | 10.4738 | 7.1893 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.67% | 1.8721 | -6.215 | -51.0738 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 58 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.5521e-05; SAMPLE 49Ti40Ar: 5.7248e-06; SAMPLE 90Zr: 13.3793; SAMPLE 91Zr: 2.9867; SAMPLE 92Zr: 4.6771; SAMPLE 94Zr: 4.9699; SAMPLE 96Zr: 0.8367 | 1 | -0.00018689 | 1.4609e-05 | 13.3791 | 2.9866 | 4.677 | 0 | 4.9697 | 0 | 0.83659 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1603 | 13.3791 | 2.9866 | 4.677 | 4.9697 | 0.83659 | 0.22323 | 0.34957 | 0.37145 | 0.06253 | 1.566 | 1.664 | 0.28011 | 1.0626 | 0.17887 | 0.16834 | -1.4996 | -1.051 | -0.99034 | -2.7721 | [10 empty] | 0 | 2.1603 | 0.21795 | 0.33334 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42489 | 0.43906 | -1.3883 | 0.014165 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.56% | 5.4528 | -10.4399 | 10.4885 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.1% | 13.2586 | 2.3909 | -10.5123 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 59 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -5.8067e-05; SAMPLE 49Ti40Ar: 4.1509e-05; SAMPLE 90Zr: 13.3821; SAMPLE 91Zr: 2.9874; SAMPLE 92Zr: 4.6784; SAMPLE 94Zr: 4.9712; SAMPLE 96Zr: 0.83693 | 1 | -0.00020944 | 5.0394e-05 | 13.3819 | 2.9873 | 4.6782 | 0 | 4.971 | 0 | 0.83682 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1614 | 13.3819 | 2.9873 | 4.6782 | 4.971 | 0.83682 | 0.22323 | 0.34959 | 0.37147 | 0.062534 | 1.566 | 1.664 | 0.28013 | 1.0626 | 0.17888 | 0.16834 | -1.4995 | -1.051 | -0.99029 | -2.772 | [10 empty] | 0 | 2.1614 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014134 | -1.8132 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.2% | 3.7215 | 28.2253 | -3.2707 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.87% | -20.6068 | -0.44348 | 31.2658 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 60 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.5564e-05; SAMPLE 49Ti40Ar: 4.3399e-05; SAMPLE 90Zr: 13.3313; SAMPLE 91Zr: 2.9761; SAMPLE 92Zr: 4.6607; SAMPLE 94Zr: 4.9527; SAMPLE 96Zr: 0.83385 | 1 | -0.00018693 | 5.2283e-05 | 13.3311 | 2.976 | 4.6605 | 0 | 4.9525 | 0 | 0.83375 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1632 | 13.3311 | 2.976 | 4.6605 | 4.9525 | 0.83375 | 0.22324 | 0.3496 | 0.3715 | 0.062541 | 1.566 | 1.6641 | 0.28016 | 1.0626 | 0.1789 | 0.16835 | -1.4995 | -1.051 | -0.99021 | -2.7719 | [10 empty] | 0 | 2.1632 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43906 | -1.3883 | 0.014161 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 6.6744 | -2.6477 | -9.5758 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.4% | 9.392 | -0.10217 | -22.3734 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 61 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.2114e-05; SAMPLE 49Ti40Ar: 8.4075e-06; SAMPLE 90Zr: 13.1298; SAMPLE 91Zr: 2.9311; SAMPLE 92Zr: 4.5904; SAMPLE 94Zr: 4.878; SAMPLE 96Zr: 0.82132 | 1 | -0.00018348 | 1.7292e-05 | 13.1295 | 2.931 | 4.5902 | 0 | 4.8778 | 0 | 0.82121 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1644 | 13.1295 | 2.931 | 4.5902 | 4.8778 | 0.82121 | 0.22324 | 0.34961 | 0.37152 | 0.062547 | 1.5661 | 1.6642 | 0.28018 | 1.0627 | 0.1789 | 0.16835 | -1.4995 | -1.0509 | -0.99016 | -2.7718 | [10 empty] | 0 | 2.1644 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.24951 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.01415 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 99.14% | -1.241 | 6.6038 | -1.8746 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.67% | 3.8746 | 3.4626 | 28.6507 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 62 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.4945e-05; SAMPLE 49Ti40Ar: 1.5608e-05; SAMPLE 90Zr: 13.1562; SAMPLE 91Zr: 2.937; SAMPLE 92Zr: 4.5998; SAMPLE 94Zr: 4.8883; SAMPLE 96Zr: 0.82308 | 1 | -0.00018631 | 2.4493e-05 | 13.156 | 2.937 | 4.5997 | 0 | 4.8881 | 0 | 0.82297 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1661 | 13.156 | 2.937 | 4.5997 | 4.8881 | 0.82297 | 0.22324 | 0.34962 | 0.37155 | 0.062555 | 1.5661 | 1.6643 | 0.28021 | 1.0627 | 0.17892 | 0.16836 | -1.4995 | -1.0509 | -0.99008 | -2.7717 | [10 empty] | 0 | 2.1661 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5513 | 0.24951 | 1.0143 | 0.16314 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42492 | 0.43907 | -1.3883 | 0.014152 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.94% | -11.7 | 4.7065 | 26.0442 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100% | -10.4098 | -2.293 | -39.0791 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 63 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.6327e-05; SAMPLE 49Ti40Ar: 7.7472e-06; SAMPLE 90Zr: 13.1697; SAMPLE 91Zr: 2.9401; SAMPLE 92Zr: 4.6047; SAMPLE 94Zr: 4.8936; SAMPLE 96Zr: 0.82398 | 1 | -0.0001877 | 1.6631e-05 | 13.1695 | 2.94 | 4.6045 | 0 | 4.8934 | 0 | 0.82387 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1677 | 13.1695 | 2.94 | 4.6045 | 4.8934 | 0.82387 | 0.22324 | 0.34964 | 0.37157 | 0.062559 | 1.5662 | 1.6644 | 0.28023 | 1.0627 | 0.17893 | 0.16836 | -1.4995 | -1.0509 | -0.99002 | -2.7716 | [10 empty] | 0 | 2.1677 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43908 | -1.3883 | 0.014152 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 100.7% | -5.4567 | -0.41162 | -10.5982 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.97% | 13.5086 | -2.3051 | 49.8694 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 64 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -4.2061e-05; SAMPLE 49Ti40Ar: 6.0218e-05; SAMPLE 90Zr: 12.9962; SAMPLE 91Zr: 2.9014; SAMPLE 92Zr: 4.5441; SAMPLE 94Zr: 4.8293; SAMPLE 96Zr: 0.81317 | 1 | -0.00019343 | 6.9102e-05 | 12.996 | 2.9013 | 4.544 | 0 | 4.8291 | 0 | 0.81306 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1686 | 12.996 | 2.9013 | 4.544 | 4.8291 | 0.81306 | 0.22325 | 0.34964 | 0.37158 | 0.062562 | 1.5662 | 1.6644 | 0.28024 | 1.0628 | 0.17893 | 0.16837 | -1.4995 | -1.0508 | -0.98998 | -2.7716 | [10 empty] | 0 | 2.1686 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43907 | -1.3883 | 0.014152 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100% | 10.586 | -1.3842 | -9.5796 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | -15.5064 | 2.4266 | -44.8974 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 65 | iRM6 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: 8.5607e-05; SAMPLE 49Ti40Ar: 8.1943e-06; SAMPLE 90Zr: 12.6866; SAMPLE 91Zr: 2.8323; SAMPLE 92Zr: 4.436; SAMPLE 94Zr: 4.7145; SAMPLE 96Zr: 0.79386 | 1 | -6.5763e-05 | 1.7078e-05 | 12.6864 | 2.8323 | 4.4359 | 0 | 4.7143 | 0 | 0.79375 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1698 | 12.6864 | 2.8323 | 4.4359 | 4.7143 | 0.79375 | 0.22325 | 0.34966 | 0.37161 | 0.062567 | 1.5662 | 1.6645 | 0.28025 | 1.0628 | 0.17894 | 0.16837 | -1.4995 | -1.0508 | -0.98992 | -2.7715 | [10 empty] | 0 | 2.1698 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014141 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 21, 2020 | 2 | 98.22% | 8.2655 | 9.9206 | -1.7241 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.77% | 22.0973 | 8.9988 | 26.2171 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 66 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -2.1413e-05; SAMPLE 49Ti40Ar: 3.3711e-05; SAMPLE 90Zr: 12.8371; SAMPLE 91Zr: 2.8659; SAMPLE 92Zr: 4.4887; SAMPLE 94Zr: 4.7707; SAMPLE 96Zr: 0.80334 | 1 | -0.00017278 | 4.2595e-05 | 12.8369 | 2.8659 | 4.4886 | 0 | 4.7705 | 0 | 0.80323 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1712 | 12.8369 | 2.8659 | 4.4886 | 4.7705 | 0.80323 | 0.22325 | 0.34966 | 0.37163 | 0.062572 | 1.5662 | 1.6646 | 0.28028 | 1.0628 | 0.17895 | 0.16837 | -1.4995 | -1.0508 | -0.98986 | -2.7714 | [10 empty] | 0 | 2.1712 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42493 | 0.43908 | -1.3883 | 0.014149 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 99.39% | -8.7199 | 3.4833 | -23.0128 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100% | -23.5833 | -21.5761 | -9.171 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 67 | iRM10 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.5362e-05; SAMPLE 49Ti40Ar: 2.6345e-05; SAMPLE 90Zr: 12.9041; SAMPLE 91Zr: 2.881; SAMPLE 92Zr: 4.5123; SAMPLE 94Zr: 4.796; SAMPLE 96Zr: 0.80763 | 1 | -0.00018673 | 3.5229e-05 | 12.9039 | 2.8809 | 4.5122 | 0 | 4.7958 | 0 | 0.80752 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1729 | 12.9039 | 2.8809 | 4.5122 | 4.7958 | 0.80752 | 0.22326 | 0.34967 | 0.37165 | 0.062579 | 1.5662 | 1.6647 | 0.2803 | 1.0629 | 0.17897 | 0.16838 | -1.4994 | -1.0508 | -0.98979 | -2.7713 | [10 empty] | 0 | 2.1729 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5513 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43907 | -1.3883 | 0.014155 | -1.8132 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 3.9118 | -3.7416 | -9.8633 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.65% | 15.5657 | 4.6474 | -8.8755 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 68 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -4.2743e-05; SAMPLE 49Ti40Ar: 5.285e-06; SAMPLE 90Zr: 12.8355; SAMPLE 91Zr: 2.8657; SAMPLE 92Zr: 4.4884; SAMPLE 94Zr: 4.7707; SAMPLE 96Zr: 0.80342 | 1 | -0.00019411 | 1.4169e-05 | 12.8352 | 2.8656 | 4.4883 | 0 | 4.7705 | 0 | 0.80331 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1742 | 12.8352 | 2.8656 | 4.4883 | 4.7705 | 0.80331 | 0.22326 | 0.34968 | 0.37168 | 0.062586 | 1.5663 | 1.6648 | 0.28033 | 1.0629 | 0.17898 | 0.16839 | -1.4994 | -1.0507 | -0.98973 | -2.7712 | [10 empty] | 0 | 2.1742 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5513 | 0.24951 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42492 | 0.43907 | -1.3883 | 0.014154 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.5% | 0.13293 | -0.3846 | 17.7775 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.5% | -9.3278 | 10.0943 | 3.2543 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 69 | iRM4 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -1.7496e-05; SAMPLE 49Ti40Ar: 1.0438e-05; SAMPLE 90Zr: 12.7767; SAMPLE 91Zr: 2.8527; SAMPLE 92Zr: 4.468; SAMPLE 94Zr: 4.7492; SAMPLE 96Zr: 0.79979 | 1 | -0.00016887 | 1.9322e-05 | 12.7764 | 2.8526 | 4.4679 | 0 | 4.749 | 0 | 0.79968 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1758 | 12.7764 | 2.8526 | 4.4679 | 4.749 | 0.79968 | 0.22327 | 0.3497 | 0.3717 | 0.06259 | 1.5663 | 1.6648 | 0.28033 | 1.0629 | 0.17898 | 0.16839 | -1.4994 | -1.0507 | -0.98966 | -2.7711 | [10 empty] | 0 | 2.1758 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014149 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 21, 2020 | 2 | 100% | 18.7534 | 6.5304 | -45.8332 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.67% | 11.6939 | 5.2525 | 10.457 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 70 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.3844e-05; SAMPLE 49Ti40Ar: -1.1013e-05; SAMPLE 90Zr: 12.7183; SAMPLE 91Zr: 2.8396; SAMPLE 92Zr: 4.4477; SAMPLE 94Zr: 4.7277; SAMPLE 96Zr: 0.79622 | 1 | -0.00018521 | -2.1289e-06 | 12.7181 | 2.8395 | 4.4476 | 0 | 4.7275 | 0 | 0.79611 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1766 | 12.7181 | 2.8395 | 4.4476 | 4.7275 | 0.79611 | 0.22327 | 0.3497 | 0.37172 | 0.062596 | 1.5663 | 1.6649 | 0.28036 | 1.0629 | 0.179 | 0.1684 | -1.4994 | -1.0507 | -0.98963 | -2.771 | [10 empty] | 0 | 2.1766 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.24951 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43907 | -1.3883 | 0.014157 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 21, 2020 | 2 | 100.1% | 2.4783 | -5.8966 | -12.2242 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | -7.7485 | -12.0909 | -13.4382 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 71 | iRM11 | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.2622e-05; SAMPLE 49Ti40Ar: 2.8091e-05; SAMPLE 90Zr: 12.7994; SAMPLE 91Zr: 2.8577; SAMPLE 92Zr: 4.4763; SAMPLE 94Zr: 4.7583; SAMPLE 96Zr: 0.80139 | 1 | -0.00018399 | 3.6975e-05 | 12.7992 | 2.8577 | 4.4762 | 0 | 4.7581 | 0 | 0.80128 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1788 | 12.7992 | 2.8577 | 4.4762 | 4.7581 | 0.80128 | 0.22327 | 0.34972 | 0.37175 | 0.062604 | 1.5664 | 1.665 | 0.2804 | 1.063 | 0.17901 | 0.1684 | -1.4994 | -1.0506 | -0.98953 | -2.7709 | [10 empty] | 0 | 2.1788 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42494 | 0.43908 | -1.3883 | 0.014149 | -1.8132 | -1.8273 | [14 empty] | iRM11 | Lab 3 | Jul 21, 2020 | 2 | 101.3% | -13.2272 | 4.6327 | -21.8299 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.7% | -2.6515 | 3.5631 | 42.6389 | SRM 3169
Lab 3 | Jul 21, 2020 | 2 | 72 | STD | OPZ 177Hf++: 0.00015137; OPZ 49Ti40Ar: -8.8841e-06; OPZ 90Zr: 0.0002141; OPZ 91Zr: 7.2305e-05; OPZ 92Zr: 0.00013289; OPZ 94Zr: 0.00019201; OPZ 96Zr: 0.00010929; SAMPLE 177Hf++: -3.4598e-05; SAMPLE 49Ti40Ar: 2.8911e-06; SAMPLE 90Zr: 12.5634; SAMPLE 91Zr: 2.8051; SAMPLE 92Zr: 4.3938; SAMPLE 94Zr: 4.6708; SAMPLE 96Zr: 0.7867 | 1 | -0.00018597 | 1.1775e-05 | 12.5632 | 2.805 | 4.3937 | 0 | 4.6706 | 0 | 0.78659 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.1798 | 12.5632 | 2.805 | 4.3937 | 4.6706 | 0.78659 | 0.22328 | 0.34973 | 0.37177 | 0.062611 | 1.5664 | 1.6651 | 0.28042 | 1.063 | 0.17903 | 0.16841 | -1.4994 | -1.0506 | -0.98949 | -2.7708 | [10 empty] | 0 | 2.1798 | 0.21795 | 0.33335 | 0.3381 | 0.054382 | 1.5295 | 1.5513 | 0.24951 | 1.0142 | 0.16314 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42492 | 0.43907 | -1.3882 | 0.014149 | -1.8132 | -1.8273 | [22 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 1.4018 | 5.9744 | -50.8231 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 2 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 3.5716e-05; SAMPLE 49Ti40Ar: 2.9029e-05; SAMPLE 90Zr: 12.661; SAMPLE 91Zr: 2.8279; SAMPLE 92Zr: 4.4311; SAMPLE 94Zr: 4.7134; SAMPLE 96Zr: 0.79433 | 1 | -5.1606e-05 | 4.1712e-05 | 12.6608 | 2.8279 | 4.431 | 0 | 4.7132 | 0 | 0.79425 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2106 | 12.6608 | 2.8279 | 4.431 | 4.7132 | 0.79425 | 0.22336 | 0.34997 | 0.37227 | 0.062733 | 1.5669 | 1.6667 | 0.28087 | 1.0637 | 0.17925 | 0.16852 | -1.499 | -1.0499 | -0.98815 | -2.7689 | [9 empty] | L3 JulS3 | 0 | 2.2106 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.01413 | -1.8132 | -1.8274 | [22 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.49% | 8.4502 | -11.1115 | 23.8202 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 3 | iRM11 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.5217e-05; SAMPLE 49Ti40Ar: 3.5236e-05; SAMPLE 90Zr: 12.592; SAMPLE 91Zr: 2.8126; SAMPLE 92Zr: 4.4071; SAMPLE 94Zr: 4.688; SAMPLE 96Zr: 0.79009 | 1 | -7.2105e-05 | 4.7919e-05 | 12.5918 | 2.8125 | 4.4069 | 0 | 4.6878 | 0 | 0.79001 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2124 | 12.5918 | 2.8125 | 4.4069 | 4.6878 | 0.79001 | 0.22336 | 0.34999 | 0.37229 | 0.06274 | 1.5669 | 1.6668 | 0.28089 | 1.0637 | 0.17927 | 0.16852 | -1.499 | -1.0499 | -0.98807 | -2.7687 | [10 empty] | 0 | 2.2124 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014135 | -1.8132 | -1.8273 | [14 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 99.98% | 6.0863 | -0.85961 | 46.419 | STD | Lab 3 | Jul 22, 2020 | 3 | 100.6% | -11.5156 | 3.8864 | 8.0104 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 4 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.1241e-05; SAMPLE 49Ti40Ar: 3.055e-05; SAMPLE 90Zr: 12.5269; SAMPLE 91Zr: 2.7981; SAMPLE 92Zr: 4.3844; SAMPLE 94Zr: 4.664; SAMPLE 96Zr: 0.78601 | 1 | -7.6081e-05 | 4.3232e-05 | 12.5267 | 2.798 | 4.3843 | 0 | 4.6639 | 0 | 0.78592 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2136 | 12.5267 | 2.798 | 4.3843 | 4.6639 | 0.78592 | 0.22337 | 0.34999 | 0.37231 | 0.06274 | 1.5669 | 1.6668 | 0.28089 | 1.0638 | 0.17926 | 0.16851 | -1.4989 | -1.0498 | -0.98802 | -2.7688 | [10 empty] | 0 | 2.2136 | 0.21796 | 0.33335 | 0.3381 | 0.054375 | 1.5294 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43904 | -1.3884 | 0.014138 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.67% | 1.8721 | -6.215 | -51.0738 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.64% | 1.2396 | 5.3977 | -25.2005 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 5 | iRM5 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.1064e-05; SAMPLE 49Ti40Ar: 3.6806e-05; SAMPLE 90Zr: 12.5385; SAMPLE 91Zr: 2.8007; SAMPLE 92Zr: 4.3886; SAMPLE 94Zr: 4.6685; SAMPLE 96Zr: 0.7868 | 1 | -7.6258e-05 | 4.9488e-05 | 12.5383 | 2.8007 | 4.3885 | 0 | 4.6684 | 0 | 0.78672 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2144 | 12.5383 | 2.8007 | 4.3885 | 4.6684 | 0.78672 | 0.22337 | 0.35001 | 0.37233 | 0.062745 | 1.5669 | 1.6669 | 0.2809 | 1.0638 | 0.17927 | 0.16852 | -1.4989 | -1.0498 | -0.98798 | -2.7687 | [10 empty] | 0 | 2.2144 | 0.21796 | 0.33336 | 0.3381 | 0.054377 | 1.5295 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5234 | -1.0985 | -2.9118 | 0.42491 | 0.43903 | -1.3884 | 0.014119 | -1.8133 | -1.8274 | [14 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 7.9236 | 16.2301 | 21.3491 | STD | Lab 3 | Jul 22, 2020 | 3 | 100.1% | 13.0493 | 4.3262 | -11.2225 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 6 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 6.6237e-06; SAMPLE 49Ti40Ar: 5.9905e-05; SAMPLE 90Zr: 12.475; SAMPLE 91Zr: 2.7866; SAMPLE 92Zr: 4.3664; SAMPLE 94Zr: 4.6451; SAMPLE 96Zr: 0.78286 | 1 | -8.0698e-05 | 7.2587e-05 | 12.4748 | 2.7865 | 4.3663 | 0 | 4.6449 | 0 | 0.78278 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2153 | 12.4748 | 2.7865 | 4.3663 | 4.6449 | 0.78278 | 0.22337 | 0.35001 | 0.37234 | 0.062748 | 1.5669 | 1.6669 | 0.28092 | 1.0638 | 0.17928 | 0.16852 | -1.4989 | -1.0498 | -0.98794 | -2.7686 | [10 empty] | 0 | 2.2153 | 0.21796 | 0.33336 | 0.3381 | 0.054376 | 1.5294 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.4249 | 0.43904 | -1.3884 | 0.014133 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.1% | 13.2586 | 2.3909 | -10.5123 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.8% | -10.1967 | -10.1231 | 34.5574 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 7 | iRM6 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 3.8818e-05; SAMPLE 49Ti40Ar: 3.58e-05; SAMPLE 90Zr: 12.5164; SAMPLE 91Zr: 2.7958; SAMPLE 92Zr: 4.3811; SAMPLE 94Zr: 4.6608; SAMPLE 96Zr: 0.78558 | 1 | -4.8504e-05 | 4.8482e-05 | 12.5162 | 2.7958 | 4.381 | 0 | 4.6606 | 0 | 0.7855 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2169 | 12.5162 | 2.7958 | 4.381 | 4.6606 | 0.7855 | 0.22337 | 0.35002 | 0.37237 | 0.062759 | 1.567 | 1.667 | 0.28096 | 1.0638 | 0.1793 | 0.16854 | -1.4989 | -1.0498 | -0.98787 | -2.7685 | [10 empty] | 0 | 2.2169 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014122 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.6% | 7.9493 | 10.581 | 39.476 | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 3.3176 | -3.4025 | -32.9452 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 8 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 4.5215e-05; SAMPLE 49Ti40Ar: 3.4358e-05; SAMPLE 90Zr: 12.4056; SAMPLE 91Zr: 2.771; SAMPLE 92Zr: 4.3423; SAMPLE 94Zr: 4.6196; SAMPLE 96Zr: 0.77864 | 1 | -4.2107e-05 | 4.704e-05 | 12.4054 | 2.771 | 4.3422 | 0 | 4.6195 | 0 | 0.77856 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2174 | 12.4054 | 2.771 | 4.3422 | 4.6195 | 0.77856 | 0.22337 | 0.35003 | 0.37238 | 0.06276 | 1.567 | 1.6671 | 0.28097 | 1.0639 | 0.1793 | 0.16854 | -1.4989 | -1.0498 | -0.98785 | -2.7684 | [10 empty] | 0 | 2.2174 | 0.21795 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43906 | -1.3883 | 0.014133 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.87% | -20.6068 | -0.44348 | 31.2658 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.68% | -15.1423 | 13.3923 | 69.7256 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 9 | iRM10 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 6.1313e-05; SAMPLE 49Ti40Ar: 3.1976e-05; SAMPLE 90Zr: 12.5859; SAMPLE 91Zr: 2.8114; SAMPLE 92Zr: 4.4055; SAMPLE 94Zr: 4.687; SAMPLE 96Zr: 0.79001 | 1 | -2.6009e-05 | 4.4658e-05 | 12.5857 | 2.8114 | 4.4054 | 0 | 4.6869 | 0 | 0.78993 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2186 | 12.5857 | 2.8114 | 4.4054 | 4.6869 | 0.78993 | 0.22338 | 0.35003 | 0.37239 | 0.062764 | 1.567 | 1.6671 | 0.28098 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0497 | -0.9878 | -2.7684 | [10 empty] | 0 | 2.2186 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014144 | -1.8132 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.6% | 5.1132 | -10.9547 | 2.2026 | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 24.2058 | -10.0698 | -79.2046 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 10 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 5.261e-05; SAMPLE 49Ti40Ar: 5.227e-05; SAMPLE 90Zr: 12.3693; SAMPLE 91Zr: 2.7631; SAMPLE 92Zr: 4.3298; SAMPLE 94Zr: 4.6065; SAMPLE 96Zr: 0.77643 | 1 | -3.4713e-05 | 6.4952e-05 | 12.3691 | 2.763 | 4.3297 | 0 | 4.6063 | 0 | 0.77635 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2191 | 12.3691 | 2.763 | 4.3297 | 4.6063 | 0.77635 | 0.22338 | 0.35004 | 0.3724 | 0.062765 | 1.567 | 1.6671 | 0.28098 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0497 | -0.98778 | -2.7684 | [10 empty] | 0 | 2.2191 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014132 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.4% | 9.392 | -0.10217 | -22.3734 | [8 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.1% | -24.4534 | -3.3746 | 24.2955 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 11 | iRM11 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 3.898e-05; SAMPLE 49Ti40Ar: 2.7608e-05; SAMPLE 90Zr: 12.3112; SAMPLE 91Zr: 2.7501; SAMPLE 92Zr: 4.3095; SAMPLE 94Zr: 4.5851; SAMPLE 96Zr: 0.77283 | 1 | -4.8342e-05 | 4.029e-05 | 12.311 | 2.7501 | 4.3094 | 0 | 4.5849 | 0 | 0.77274 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2202 | 12.311 | 2.7501 | 4.3094 | 4.5849 | 0.77274 | 0.22338 | 0.35004 | 0.37242 | 0.062769 | 1.567 | 1.6672 | 0.28099 | 1.0639 | 0.17932 | 0.16854 | -1.4989 | -1.0497 | -0.98773 | -2.7683 | [10 empty] | 0 | 2.2202 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5294 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3884 | 0.014143 | -1.8133 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100% | -0.89466 | -11.422 | -34.0219 | STD | Lab 3 | Jul 22, 2020 | 3 | 99.88% | 25.5306 | 12.1591 | 20.2636 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 12 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.9072e-05; SAMPLE 49Ti40Ar: 1.9219e-05; SAMPLE 90Zr: 12.2441; SAMPLE 91Zr: 2.7351; SAMPLE 92Zr: 4.2861; SAMPLE 94Zr: 4.5602; SAMPLE 96Zr: 0.76869 | 1 | -6.825e-05 | 3.1901e-05 | 12.2439 | 2.7351 | 4.286 | 0 | 4.56 | 0 | 0.76861 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.221 | 12.2439 | 2.7351 | 4.286 | 4.56 | 0.76861 | 0.22338 | 0.35005 | 0.37243 | 0.062775 | 1.567 | 1.6672 | 0.28102 | 1.0639 | 0.17933 | 0.16855 | -1.4989 | -1.0497 | -0.98769 | -2.7682 | [10 empty] | 0 | 2.221 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014131 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.67% | 3.8746 | 3.4626 | 28.6507
Lab 3 | Jul 22, 2020 | 3 | 13 | iRM5 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.4567e-05; SAMPLE 49Ti40Ar: 3.8923e-05; SAMPLE 90Zr: 12.2964; SAMPLE 91Zr: 2.7468; SAMPLE 92Zr: 4.3045; SAMPLE 94Zr: 4.5799; SAMPLE 96Zr: 0.77204 | 1 | -6.2755e-05 | 5.1605e-05 | 12.2962 | 2.7468 | 4.3043 | 0 | 4.5798 | 0 | 0.77195 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2222 | 12.2962 | 2.7468 | 4.3043 | 4.5798 | 0.77195 | 0.22338 | 0.35005 | 0.37245 | 0.06278 | 1.567 | 1.6673 | 0.28104 | 1.064 | 0.17934 | 0.16856 | -1.4989 | -1.0497 | -0.98764 | -2.7681 | [10 empty] | 0 | 2.2222 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014154 | -1.8132 | -1.8274 | [14 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.6% | -7.2536 | -20.1022 | 20.5925 | STD | Lab 3 | Jul 29, 2020 | 4 | 99.28% | 14.2258 | 11.3455 | -21.205 | SRM 3169 | Avg | 0.52064 | 0.14497 | 0.35594
Lab 3 | Jul 22, 2020 | 3 | 14 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.1836e-05; SAMPLE 49Ti40Ar: -9.8615e-06; SAMPLE 90Zr: 12.199; SAMPLE 91Zr: 2.7251; SAMPLE 92Zr: 4.2705; SAMPLE 94Zr: 4.5438; SAMPLE 96Zr: 0.76594 | 1 | -7.5486e-05 | 2.8208e-06 | 12.1988 | 2.7251 | 4.2704 | 0 | 4.5437 | 0 | 0.76586 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2232 | 12.1988 | 2.7251 | 4.2704 | 4.5437 | 0.76586 | 0.22339 | 0.35007 | 0.37247 | 0.062782 | 1.5671 | 1.6674 | 0.28104 | 1.064 | 0.17934 | 0.16856 | -1.4988 | -1.0496 | -0.9876 | -2.7681 | [10 empty] | 0 | 2.2232 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014137 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100% | -10.4098 | -2.293 | -39.0791 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.9% | -10.0618 | -8.4872 | 68.7184 | SRM 3169 | 2SD | 21.0125 | 13.1287 | 58.9569
Lab 3 | Jul 22, 2020 | 3 | 15 | iRM6 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 4.1846e-05; SAMPLE 49Ti40Ar: 5.8057e-06; SAMPLE 90Zr: 12.2735; SAMPLE 91Zr: 2.7418; SAMPLE 92Zr: 4.2967; SAMPLE 94Zr: 4.5717; SAMPLE 96Zr: 0.77067 | 1 | -4.5476e-05 | 1.8488e-05 | 12.2733 | 2.7417 | 4.2965 | 0 | 4.5716 | 0 | 0.77059 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2239 | 12.2733 | 2.7417 | 4.2965 | 4.5716 | 0.77059 | 0.22339 | 0.35007 | 0.37248 | 0.062786 | 1.5671 | 1.6674 | 0.28106 | 1.064 | 0.17935 | 0.16856 | -1.4988 | -1.0496 | -0.98757 | -2.768 | [10 empty] | 0 | 2.2239 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014138 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | -15.4762 | -0.22744 | -2.3322 | STD | Lab 3 | Jul 29, 2020 | 4 | 100.1% | 7.4027 | 0.8896 | -30.8248 | SRM 3169 | n | 35 | 35 | 35
Lab 3 | Jul 22, 2020 | 3 | 16 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 4.4997e-05; SAMPLE 49Ti40Ar: 3.2681e-05; SAMPLE 90Zr: 12.1504; SAMPLE 91Zr: 2.7144; SAMPLE 92Zr: 4.2537; SAMPLE 94Zr: 4.5261; SAMPLE 96Zr: 0.76301 | 1 | -4.2325e-05 | 4.5363e-05 | 12.1502 | 2.7143 | 4.2536 | 0 | 4.5259 | 0 | 0.76293 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2249 | 12.1502 | 2.7143 | 4.2536 | 4.5259 | 0.76293 | 0.2234 | 0.35008 | 0.3725 | 0.062792 | 1.5671 | 1.6674 | 0.28108 | 1.064 | 0.17936 | 0.16857 | -1.4988 | -1.0496 | -0.98753 | -2.7679 | [10 empty] | 0 | 2.2249 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9118 | 0.42489 | 0.43903 | -1.3883 | 0.014138 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.97% | 13.5086 | -2.3051 | 49.8694 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.09% | -4.8772 | 0.51574 | -13.31 | SRM 3169 | 2SE | 3.5518 | 2.2192 | 9.9655
Lab 3 | Jul 22, 2020 | 3 | 17 | iRM10 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 7.1588e-05; SAMPLE 49Ti40Ar: 9.5595e-06; SAMPLE 90Zr: 12.3288; SAMPLE 91Zr: 2.7542; SAMPLE 92Zr: 4.3163; SAMPLE 94Zr: 4.5928; SAMPLE 96Zr: 0.77425 | 1 | -1.5734e-05 | 2.2242e-05 | 12.3286 | 2.7542 | 4.3162 | 0 | 4.5926 | 0 | 0.77417 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2262 | 12.3286 | 2.7542 | 4.3162 | 4.5926 | 0.77417 | 0.2234 | 0.3501 | 0.37252 | 0.062794 | 1.5672 | 1.6675 | 0.28109 | 1.064 | 0.17936 | 0.16857 | -1.4988 | -1.0495 | -0.98747 | -2.7679 | [10 empty] | 0 | 2.2262 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014123 | -1.8133 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.6% | -6.8493 | 14.0345 | -11.5376 | STD | Lab 3 | Jul 29, 2020 | 4 | 101.3% | 4.0171 | 1.3975 | 13.9083 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 18 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 6.9974e-06; SAMPLE 49Ti40Ar: 2.908e-05; SAMPLE 90Zr: 12.1084; SAMPLE 91Zr: 2.705; SAMPLE 92Zr: 4.2392; SAMPLE 94Zr: 4.5109; SAMPLE 96Zr: 0.76044 | 1 | -8.0325e-05 | 4.1763e-05 | 12.1082 | 2.7049 | 4.2391 | 0 | 4.5107 | 0 | 0.76036 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2271 | 12.1082 | 2.7049 | 4.2391 | 4.5107 | 0.76036 | 0.2234 | 0.3501 | 0.37253 | 0.062797 | 1.5672 | 1.6676 | 0.2811 | 1.0641 | 0.17937 | 0.16857 | -1.4988 | -1.0495 | -0.98743 | -2.7678 | [10 empty] | 0 | 2.2271 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3884 | 0.014135 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | -15.5064 | 2.4266 | -44.8974 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.33% | -5.7183 | 1.4522 | -13.9916 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 19 | iRM11 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.0759e-05; SAMPLE 49Ti40Ar: 2.4666e-05; SAMPLE 90Zr: 12.0587; SAMPLE 91Zr: 2.6939; SAMPLE 92Zr: 4.2218; SAMPLE 94Zr: 4.4925; SAMPLE 96Zr: 0.75736 | 1 | -7.6564e-05 | 3.7349e-05 | 12.0585 | 2.6939 | 4.2217 | 0 | 4.4924 | 0 | 0.75728 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2279 | 12.0585 | 2.6939 | 4.2217 | 4.4924 | 0.75728 | 0.2234 | 0.3501 | 0.37255 | 0.062801 | 1.5671 | 1.6676 | 0.28111 | 1.0641 | 0.17938 | 0.16857 | -1.4988 | -1.0495 | -0.98739 | -2.7678 | [10 empty] | 0 | 2.2279 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014144 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.1% | -6.1346 | -7.9869 | -14.8614 | STD | Lab 3 | Jul 29, 2020 | 4 | 99.48% | -4.6024 | 2.5527 | 18.8966 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 20 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.4546e-05; SAMPLE 49Ti40Ar: 1.0965e-05; SAMPLE 90Zr: 11.9912; SAMPLE 91Zr: 2.6789; SAMPLE 92Zr: 4.1983; SAMPLE 94Zr: 4.4676; SAMPLE 96Zr: 0.75321 | 1 | -6.2776e-05 | 2.3647e-05 | 11.991 | 2.6789 | 4.1982 | 0 | 4.4674 | 0 | 0.75313 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.229 | 11.991 | 2.6789 | 4.1982 | 4.4674 | 0.75313 | 0.22341 | 0.35011 | 0.37256 | 0.062807 | 1.5672 | 1.6677 | 0.28113 | 1.0641 | 0.17939 | 0.16858 | -1.4988 | -1.0495 | -0.98734 | -2.7677 | [10 empty] | 0 | 2.229 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9118 | 0.42489 | 0.43903 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.77% | 22.0973 | 8.9988 | 26.2171 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.8% | 23.5674 | -11.2699 | -3.254 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 21 | iRM5 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.9895e-05; SAMPLE 49Ti40Ar: 2.267e-05; SAMPLE 90Zr: 12.0637; SAMPLE 91Zr: 2.6951; SAMPLE 92Zr: 4.2238; SAMPLE 94Zr: 4.4947; SAMPLE 96Zr: 0.75778 | 1 | -6.7428e-05 | 3.5352e-05 | 12.0635 | 2.695 | 4.2236 | 0 | 4.4946 | 0 | 0.7577 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2297 | 12.0635 | 2.695 | 4.2236 | 4.4946 | 0.7577 | 0.2234 | 0.35012 | 0.37257 | 0.062809 | 1.5672 | 1.6677 | 0.28115 | 1.0641 | 0.1794 | 0.16858 | -1.4988 | -1.0495 | -0.98732 | -2.7677 | [10 empty] | 0 | 2.2297 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.014142 | -1.8132 | -1.8274 | [14 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.9% | -19.3296 | 5.027 | -1.6762 | STD | Lab 3 | Jul 29, 2020 | 4 | 99.83% | -19.3185 | 6.3704 | -21.9874 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 22 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 9.8966e-06; SAMPLE 49Ti40Ar: -9.7981e-07; SAMPLE 90Zr: 11.9283; SAMPLE 91Zr: 2.6649; SAMPLE 92Zr: 4.1764; SAMPLE 94Zr: 4.4445; SAMPLE 96Zr: 0.74933 | 1 | -7.7426e-05 | 1.1703e-05 | 11.9281 | 2.6648 | 4.1763 | 0 | 4.4443 | 0 | 0.74925 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2308 | 11.9281 | 2.6648 | 4.1763 | 4.4443 | 0.74925 | 0.22341 | 0.35012 | 0.37259 | 0.062814 | 1.5672 | 1.6678 | 0.28116 | 1.0642 | 0.17941 | 0.16859 | -1.4988 | -1.0495 | -0.98727 | -2.7676 | [10 empty] | 0 | 2.2308 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014156 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100% | -23.5833 | -21.5761 | -9.171 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.33% | 4.1333 | -2.806 | 25.4785 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 23 | iRM6 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 3.1969e-05; SAMPLE 49Ti40Ar: 1.2969e-05; SAMPLE 90Zr: 11.9725; SAMPLE 91Zr: 2.6748; SAMPLE 92Zr: 4.192; SAMPLE 94Zr: 4.4611; SAMPLE 96Zr: 0.75216 | 1 | -5.5354e-05 | 2.5652e-05 | 11.9723 | 2.6747 | 4.1919 | 0 | 4.461 | 0 | 0.75208 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2317 | 11.9723 | 2.6747 | 4.1919 | 4.461 | 0.75208 | 0.22341 | 0.35013 | 0.37261 | 0.062819 | 1.5672 | 1.6678 | 0.28118 | 1.0642 | 0.17941 | 0.16859 | -1.4988 | -1.0494 | -0.98723 | -2.7675 | [10 empty] | 0 | 2.2317 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014141 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.7% | -4.3567 | 3.9379 | 22.0271 | STD | Lab 3 | Jul 29, 2020 | 4 | 101.1% | 3.8899 | 10.5944 | -15.0945 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 24 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.7922e-05; SAMPLE 49Ti40Ar: 6.2617e-06; SAMPLE 90Zr: 11.8543; SAMPLE 91Zr: 2.6484; SAMPLE 92Zr: 4.1507; SAMPLE 94Zr: 4.4172; SAMPLE 96Zr: 0.74475 | 1 | -5.94e-05 | 1.8944e-05 | 11.8541 | 2.6484 | 4.1506 | 0 | 4.417 | 0 | 0.74467 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2322 | 11.8541 | 2.6484 | 4.1506 | 4.417 | 0.74467 | 0.22341 | 0.35014 | 0.37262 | 0.062819 | 1.5672 | 1.6678 | 0.28118 | 1.0642 | 0.17941 | 0.16859 | -1.4987 | -1.0494 | -0.98721 | -2.7675 | [10 empty] | 0 | 2.2322 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9118 | 0.4249 | 0.43903 | -1.3883 | 0.014133 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.65% | 15.5657 | 4.6474 | -8.8755 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.5% | -12.9156 | -11.1492 | 11.599 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 25 | iRM10 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 6.0653e-05; SAMPLE 49Ti40Ar: 3.1931e-05; SAMPLE 90Zr: 12.0382; SAMPLE 91Zr: 2.6895; SAMPLE 92Zr: 4.2151; SAMPLE 94Zr: 4.4857; SAMPLE 96Zr: 0.75629 | 1 | -2.6669e-05 | 4.4613e-05 | 12.038 | 2.6894 | 4.2149 | 0 | 4.4855 | 0 | 0.75621 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.232 | 12.038 | 2.6894 | 4.2149 | 4.4855 | 0.75621 | 0.22341 | 0.35014 | 0.37261 | 0.062819 | 1.5672 | 1.6678 | 0.28118 | 1.0642 | 0.17941 | 0.16859 | -1.4987 | -1.0494 | -0.98722 | -2.7675 | [10 empty] | 0 | 2.232 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014134 | -1.8132 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.5% | -1.3583 | -7.9294 | -7.6742 | STD | Lab 3 | Jul 29, 2020 | 4 | 99.67% | 19.0927 | 5.524 | -5.1049 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 26 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.2757e-05; SAMPLE 49Ti40Ar: 2.5475e-05; SAMPLE 90Zr: 11.8643; SAMPLE 91Zr: 2.6506; SAMPLE 92Zr: 4.1542; SAMPLE 94Zr: 4.4208; SAMPLE 96Zr: 0.74537 | 1 | -6.4566e-05 | 3.8158e-05 | 11.8641 | 2.6506 | 4.1541 | 0 | 4.4207 | 0 | 0.74529 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2318 | 11.8641 | 2.6506 | 4.1541 | 4.4207 | 0.74529 | 0.22341 | 0.35014 | 0.37261 | 0.062819 | 1.5673 | 1.6678 | 0.28118 | 1.0642 | 0.17941 | 0.16859 | -1.4987 | -1.0494 | -0.98723 | -2.7675 | [10 empty] | 0 | 2.2318 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014118 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.5% | -9.3278 | 10.0943 | 3.2543 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.9% | -10.9878 | 1.2578 | -14.4083 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 27 | iRM11 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.9479e-05; SAMPLE 49Ti40Ar: 2.7375e-05; SAMPLE 90Zr: 11.7626; SAMPLE 91Zr: 2.628; SAMPLE 92Zr: 4.1188; SAMPLE 94Zr: 4.3834; SAMPLE 96Zr: 0.73909 | 1 | -6.7843e-05 | 4.0057e-05 | 11.7624 | 2.6279 | 4.1187 | 0 | 4.3832 | 0 | 0.73901 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2341 | 11.7624 | 2.6279 | 4.1187 | 4.3832 | 0.73901 | 0.22342 | 0.35015 | 0.37265 | 0.062828 | 1.5673 | 1.6679 | 0.28121 | 1.0642 | 0.17943 | 0.1686 | -1.4987 | -1.0494 | -0.98712 | -2.7674 | [10 empty] | 0 | 2.2341 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014132 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 99.58% | 0.62147 | -10.3504 | -20.2272 | STD | Lab 3 | Jul 29, 2020 | 4 | 99.53% | 0.20491 | -7.0959 | 4.4461 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 28 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.2509e-05; SAMPLE 49Ti40Ar: -2.6536e-06; SAMPLE 90Zr: 11.7608; SAMPLE 91Zr: 2.6276; SAMPLE 92Zr: 4.1182; SAMPLE 94Zr: 4.3828; SAMPLE 96Zr: 0.73901 | 1 | -6.4813e-05 | 1.0029e-05 | 11.7606 | 2.6275 | 4.1181 | 0 | 4.3826 | 0 | 0.73893 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2346 | 11.7606 | 2.6275 | 4.1181 | 4.3826 | 0.73893 | 0.22342 | 0.35016 | 0.37265 | 0.062831 | 1.5673 | 1.668 | 0.28123 | 1.0642 | 0.17944 | 0.1686 | -1.4987 | -1.0494 | -0.9871 | -2.7673 | [10 empty] | 0 | 2.2346 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42491 | 0.43903 | -1.3883 | 0.014124 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.67% | 11.6939 | 5.2525 | 10.457 | [8 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.19% | -5.208 | -0.58526 | 29.6219 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 29 | iRM5 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.3516e-05; SAMPLE 49Ti40Ar: 1.1262e-05; SAMPLE 90Zr: 11.8417; SAMPLE 91Zr: 2.6457; SAMPLE 92Zr: 4.1465; SAMPLE 94Zr: 4.413; SAMPLE 96Zr: 0.74411 | 1 | -6.3806e-05 | 2.3944e-05 | 11.8415 | 2.6456 | 4.1464 | 0 | 4.4128 | 0 | 0.74403 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2346 | 11.8415 | 2.6456 | 4.1464 | 4.4128 | 0.74403 | 0.22342 | 0.35016 | 0.37266 | 0.062832 | 1.5672 | 1.6679 | 0.28123 | 1.0643 | 0.17944 | 0.16861 | -1.4987 | -1.0494 | -0.9871 | -2.7673 | [10 empty] | 0 | 2.2346 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9117 | 0.42489 | 0.43902 | -1.3883 | 0.014136 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | 18.2122 | -3.5481 | 19.6356 | STD | Lab 3 | Jul 29, 2020 | 4 | 100.9% | 7.479 | 8.2588 | -36.2426 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 30 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: -1.2094e-05; SAMPLE 49Ti40Ar: 2.4688e-05; SAMPLE 90Zr: 11.7358; SAMPLE 91Zr: 2.622; SAMPLE 92Zr: 4.1095; SAMPLE 94Zr: 4.3736; SAMPLE 96Zr: 0.73748 | 1 | -9.9416e-05 | 3.737e-05 | 11.7356 | 2.6219 | 4.1093 | 0 | 4.3735 | 0 | 0.7374 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2354 | 11.7356 | 2.6219 | 4.1093 | 4.3735 | 0.7374 | 0.22342 | 0.35016 | 0.37267 | 0.062834 | 1.5673 | 1.668 | 0.28124 | 1.0643 | 0.17944 | 0.16861 | -1.4987 | -1.0494 | -0.98707 | -2.7673 | [10 empty] | 0 | 2.2354 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014141 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | -7.7485 | -12.0909 | -13.4382 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.3% | 5.8025 | -3.3423 | 19.6871 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 31 | iRM6 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.5778e-05; SAMPLE 49Ti40Ar: 4.9801e-05; SAMPLE 90Zr: 11.7972; SAMPLE 91Zr: 2.6357; SAMPLE 92Zr: 4.131; SAMPLE 94Zr: 4.3967; SAMPLE 96Zr: 0.74138 | 1 | -6.1544e-05 | 6.2484e-05 | 11.797 | 2.6357 | 4.1309 | 0 | 4.3965 | 0 | 0.7413 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2364 | 11.797 | 2.6357 | 4.1309 | 4.3965 | 0.7413 | 0.22342 | 0.35017 | 0.37268 | 0.062838 | 1.5673 | 1.6681 | 0.28126 | 1.0643 | 0.17945 | 0.16861 | -1.4987 | -1.0493 | -0.98702 | -2.7672 | [10 empty] | 0 | 2.2364 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014146 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 101% | -1.4651 | -8.9293 | -15.6156 | STD | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -8.4781 | -0.67832 | -12.308 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 32 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 4.2818e-05; SAMPLE 49Ti40Ar: 4.6517e-05; SAMPLE 90Zr: 11.6337; SAMPLE 91Zr: 2.5993; SAMPLE 92Zr: 4.074; SAMPLE 94Zr: 4.3361; SAMPLE 96Zr: 0.73119 | 1 | -4.4504e-05 | 5.9199e-05 | 11.6335 | 2.5992 | 4.0739 | 0 | 4.3359 | 0 | 0.73111 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2378 | 11.6335 | 2.5992 | 4.0739 | 4.3359 | 0.73111 | 0.22342 | 0.35018 | 0.37271 | 0.062845 | 1.5673 | 1.6682 | 0.28128 | 1.0643 | 0.17946 | 0.16862 | -1.4987 | -1.0493 | -0.98696 | -2.7671 | [10 empty] | 0 | 2.2378 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42492 | 0.43905 | -1.3883 | 0.014133 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.7% | -2.6515 | 3.5631 | 42.6389 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.1% | 7.9021 | -1.896 | -4.5155 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 33 | iRM10 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 7.3087e-05; SAMPLE 49Ti40Ar: 5.2084e-05; SAMPLE 90Zr: 11.7898; SAMPLE 91Zr: 2.6342; SAMPLE 92Zr: 4.1287; SAMPLE 94Zr: 4.3945; SAMPLE 96Zr: 0.74105 | 1 | -1.4235e-05 | 6.4767e-05 | 11.7896 | 2.6341 | 4.1286 | 0 | 4.3943 | 0 | 0.74097 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2393 | 11.7896 | 2.6341 | 4.1286 | 4.3943 | 0.74097 | 0.22343 | 0.35019 | 0.37273 | 0.06285 | 1.5673 | 1.6682 | 0.2813 | 1.0644 | 0.17947 | 0.16862 | -1.4987 | -1.0493 | -0.9869 | -2.767 | [10 empty] | 0 | 2.2393 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014146 | -1.8132 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.5% | 0.44756 | -13.2895 | 3.6109 | STD | Lab 3 | Jul 30, 2020 | 5 | 99.01% | 5.5028 | 6.6802 | 34.6036 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 34 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 5.6095e-06; SAMPLE 49Ti40Ar: 3.5569e-05; SAMPLE 90Zr: 11.6012; SAMPLE 91Zr: 2.5921; SAMPLE 92Zr: 4.0628; SAMPLE 94Zr: 4.3244; SAMPLE 96Zr: 0.72922 | 1 | -8.1713e-05 | 4.8252e-05 | 11.601 | 2.592 | 4.0627 | 0 | 4.3243 | 0 | 0.72914 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2404 | 11.601 | 2.592 | 4.0627 | 4.3243 | 0.72914 | 0.22343 | 0.3502 | 0.37275 | 0.062852 | 1.5674 | 1.6683 | 0.2813 | 1.0644 | 0.17947 | 0.16862 | -1.4987 | -1.0492 | -0.98685 | -2.767 | [10 empty] | 0 | 2.2404 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014132 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 1.4018 | 5.9744 | -50.8231 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.8% | -13.1029 | 2.8269 | -7.7566 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 35 | iRM11 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.112e-06; SAMPLE 49Ti40Ar: 3.0377e-05; SAMPLE 90Zr: 11.5634; SAMPLE 91Zr: 2.5836; SAMPLE 92Zr: 4.0496; SAMPLE 94Zr: 4.3104; SAMPLE 96Zr: 0.72689 | 1 | -8.621e-05 | 4.306e-05 | 11.5632 | 2.5836 | 4.0494 | 0 | 4.3102 | 0 | 0.72681 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2409 | 11.5632 | 2.5836 | 4.0494 | 4.3102 | 0.72681 | 0.22343 | 0.3502 | 0.37276 | 0.062856 | 1.5674 | 1.6683 | 0.28132 | 1.0644 | 0.17949 | 0.16862 | -1.4987 | -1.0492 | -0.98683 | -2.7669 | [10 empty] | 0 | 2.2409 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014147 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.1% | -7.3496 | -9.4811 | 16.3832 | STD | Lab 3 | Jul 30, 2020 | 5 | 100% | 13.4545 | -16.3101 | -47.1196 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 36 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.757e-05; SAMPLE 49Ti40Ar: 2.8338e-05; SAMPLE 90Zr: 11.5009; SAMPLE 91Zr: 2.5697; SAMPLE 92Zr: 4.0278; SAMPLE 94Zr: 4.2873; SAMPLE 96Zr: 0.723 | 1 | -5.9752e-05 | 4.102e-05 | 11.5007 | 2.5697 | 4.0276 | 0 | 4.2871 | 0 | 0.72292 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2417 | 11.5007 | 2.5697 | 4.0276 | 4.2871 | 0.72292 | 0.22344 | 0.35021 | 0.37277 | 0.062859 | 1.5674 | 1.6684 | 0.28133 | 1.0644 | 0.17949 | 0.16863 | -1.4986 | -1.0492 | -0.98679 | -2.7669 | [10 empty] | 0 | 2.2417 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014144 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.49% | 8.4502 | -11.1115 | 23.8202 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 99.96% | -17.29 | 11.4632 | 78.321 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 37 | iRM5 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.7464e-05; SAMPLE 49Ti40Ar: 3.6378e-05; SAMPLE 90Zr: 11.613; SAMPLE 91Zr: 2.5947; SAMPLE 92Zr: 4.067; SAMPLE 94Zr: 4.329; SAMPLE 96Zr: 0.73003 | 1 | -6.9858e-05 | 4.9061e-05 | 11.6128 | 2.5947 | 4.0669 | 0 | 4.3288 | 0 | 0.72995 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2411 | 11.6128 | 2.5947 | 4.0669 | 4.3288 | 0.72995 | 0.22343 | 0.35021 | 0.37276 | 0.062857 | 1.5674 | 1.6683 | 0.28133 | 1.0644 | 0.17949 | 0.16863 | -1.4986 | -1.0492 | -0.98682 | -2.7669 | [10 empty] | 0 | 2.2411 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014137 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.9% | -0.12937 | 1.6245 | 8.1355 | STD | Lab 3 | Jul 30, 2020 | 5 | 99.41% | 6.4852 | 1.5307 | -74.3031 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 38 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.5639e-05; SAMPLE 49Ti40Ar: -6.9076e-06; SAMPLE 90Zr: 11.5175; SAMPLE 91Zr: 2.5734; SAMPLE 92Zr: 4.0336; SAMPLE 94Zr: 4.2935; SAMPLE 96Zr: 0.72405 | 1 | -6.1683e-05 | 5.7747e-06 | 11.5173 | 2.5733 | 4.0335 | 0 | 4.2933 | 0 | 0.72397 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2418 | 11.5173 | 2.5733 | 4.0335 | 4.2933 | 0.72397 | 0.22343 | 0.35021 | 0.37277 | 0.062859 | 1.5674 | 1.6684 | 0.28133 | 1.0644 | 0.17949 | 0.16863 | -1.4986 | -1.0492 | -0.98679 | -2.7669 | [10 empty] | 0 | 2.2418 | 0.21795 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014133 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.6% | -11.5156 | 3.8864 | 8.0104 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.3% | 12.1667 | -6.7645 | 39.1375 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 39 | iRM6 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 3.6436e-05; SAMPLE 49Ti40Ar: 3.8577e-05; SAMPLE 90Zr: 11.5536; SAMPLE 91Zr: 2.5815; SAMPLE 92Zr: 4.0463; SAMPLE 94Zr: 4.3072; SAMPLE 96Zr: 0.72639 | 1 | -5.0886e-05 | 5.126e-05 | 11.5534 | 2.5814 | 4.0462 | 0 | 4.307 | 0 | 0.72631 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2431 | 11.5534 | 2.5814 | 4.0462 | 4.307 | 0.72631 | 0.22344 | 0.35022 | 0.37279 | 0.062865 | 1.5674 | 1.6685 | 0.28136 | 1.0645 | 0.1795 | 0.16863 | -1.4986 | -1.0492 | -0.98673 | -2.7668 | [10 empty] | 0 | 2.2431 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014139 | -1.8132 | -1.8273 | [14 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | -3.9123 | -7.9507 | 31.9688 | STD | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -12.1159 | 5.4901 | -3.8008 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 40 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.6693e-05; SAMPLE 49Ti40Ar: 3.0189e-05; SAMPLE 90Zr: 11.3967; SAMPLE 91Zr: 2.5465; SAMPLE 92Zr: 3.9916; SAMPLE 94Zr: 4.249; SAMPLE 96Zr: 0.71656 | 1 | -7.063e-05 | 4.2871e-05 | 11.3965 | 2.5465 | 3.9915 | 0 | 4.2488 | 0 | 0.71648 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2445 | 11.3965 | 2.5465 | 3.9915 | 4.2488 | 0.71648 | 0.22344 | 0.35023 | 0.37282 | 0.062868 | 1.5675 | 1.6685 | 0.28136 | 1.0645 | 0.1795 | 0.16863 | -1.4986 | -1.0492 | -0.98667 | -2.7667 | [10 empty] | 0 | 2.2445 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.01413 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.64% | 1.2396 | 5.3977 | -25.2005 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.2% | -5.4377 | -1.6996 | -9.5743 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 41 | iRM10 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 8.815e-05; SAMPLE 49Ti40Ar: 1.1245e-05; SAMPLE 90Zr: 11.5396; SAMPLE 91Zr: 2.5784; SAMPLE 92Zr: 4.0417; SAMPLE 94Zr: 4.3024; SAMPLE 96Zr: 0.72561 | 1 | 8.2737e-07 | 2.3927e-05 | 11.5394 | 2.5784 | 4.0415 | 0 | 4.3022 | 0 | 0.72553 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2454 | 11.5394 | 2.5784 | 4.0415 | 4.3022 | 0.72553 | 0.22344 | 0.35024 | 0.37283 | 0.062874 | 1.5675 | 1.6686 | 0.28139 | 1.0645 | 0.17952 | 0.16864 | -1.4986 | -1.0491 | -0.98663 | -2.7666 | [10 empty] | 0 | 2.2454 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014135 | -1.8132 | -1.8273 | [14 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 101.4% | -13.1974 | -1.8582 | 30.0214 | STD | Lab 3 | Jul 30, 2020 | 5 | 98.84% | 14.0945 | -3.4966 | -11.9222 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 42 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 4.8183e-07; SAMPLE 49Ti40Ar: 2.6042e-05; SAMPLE 90Zr: 11.3594; SAMPLE 91Zr: 2.5382; SAMPLE 92Zr: 3.9786; SAMPLE 94Zr: 4.2354; SAMPLE 96Zr: 0.71431 | 1 | -8.684e-05 | 3.8724e-05 | 11.3592 | 2.5382 | 3.9785 | 0 | 4.2352 | 0 | 0.71423 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2463 | 11.3592 | 2.5382 | 3.9785 | 4.2352 | 0.71423 | 0.22345 | 0.35025 | 0.37284 | 0.062877 | 1.5675 | 1.6686 | 0.2814 | 1.0645 | 0.17952 | 0.16864 | -1.4986 | -1.0491 | -0.98659 | -2.7666 | [10 empty] | 0 | 2.2463 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.1% | 13.0493 | 4.3262 | -11.2225 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -2.3071 | 0.22484 | 17.2229 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 43 | iRM11 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 6.4685e-06; SAMPLE 49Ti40Ar: -3.1671e-06; SAMPLE 90Zr: 11.3499; SAMPLE 91Zr: 2.5361; SAMPLE 92Zr: 3.9754; SAMPLE 94Zr: 4.232; SAMPLE 96Zr: 0.71376 | 1 | -8.0854e-05 | 9.5152e-06 | 11.3497 | 2.5361 | 3.9752 | 0 | 4.2319 | 0 | 0.71368 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2473 | 11.3497 | 2.5361 | 3.9752 | 4.2319 | 0.71368 | 0.22345 | 0.35025 | 0.37286 | 0.062881 | 1.5675 | 1.6687 | 0.28141 | 1.0646 | 0.17953 | 0.16865 | -1.4986 | -1.0491 | -0.98655 | -2.7665 | [10 empty] | 0 | 2.2473 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014145 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.2% | 2.2467 | 0.86291 | -21.6014 | STD | Lab 3 | Jul 30, 2020 | 5 | 100.5% | -2.0475 | 6.1145 | -4.7415 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 44 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 3.3355e-06; SAMPLE 49Ti40Ar: 1.5524e-05; SAMPLE 90Zr: 11.2882; SAMPLE 91Zr: 2.5223; SAMPLE 92Zr: 3.9538; SAMPLE 94Zr: 4.2092; SAMPLE 96Zr: 0.70995 | 1 | -8.3987e-05 | 2.8206e-05 | 11.288 | 2.5223 | 3.9537 | 0 | 4.209 | 0 | 0.70987 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2479 | 11.288 | 2.5223 | 3.9537 | 4.209 | 0.70987 | 0.22345 | 0.35025 | 0.37287 | 0.062887 | 1.5675 | 1.6687 | 0.28144 | 1.0646 | 0.17955 | 0.16866 | -1.4986 | -1.0491 | -0.98652 | -2.7664 | [10 empty] | 0 | 2.2479 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.24951 | 1.0143 | 0.16314 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43906 | -1.3883 | 0.014154 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.8% | -10.1967 | -10.1231 | 34.5574 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.1% | -5.6146 | -6.1394 | 8.1469 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 45 | iRM5 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 5.3119e-06; SAMPLE 49Ti40Ar: 1.0164e-06; SAMPLE 90Zr: 11.3635; SAMPLE 91Zr: 2.5392; SAMPLE 92Zr: 3.9803; SAMPLE 94Zr: 4.2373; SAMPLE 96Zr: 0.71469 | 1 | -8.201e-05 | 1.3699e-05 | 11.3633 | 2.5391 | 3.9801 | 0 | 4.2372 | 0 | 0.71461 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2486 | 11.3633 | 2.5391 | 3.9801 | 4.2372 | 0.71461 | 0.22345 | 0.35026 | 0.37288 | 0.062888 | 1.5675 | 1.6687 | 0.28144 | 1.0646 | 0.17955 | 0.16865 | -1.4986 | -1.0491 | -0.98649 | -2.7664 | [10 empty] | 0 | 2.2486 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014141 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | 8.0448 | 10.9871 | -10.6413 | STD | Lab 3 | Jul 30, 2020 | 5 | 99.05% | 7.3317 | 6.3134 | -4.173 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 46 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: -4.758e-06; SAMPLE 49Ti40Ar: 1.5282e-07; SAMPLE 90Zr: 11.2615; SAMPLE 91Zr: 2.5164; SAMPLE 92Zr: 3.9446; SAMPLE 94Zr: 4.1995; SAMPLE 96Zr: 0.70832 | 1 | -9.208e-05 | 1.2835e-05 | 11.2613 | 2.5163 | 3.9445 | 0 | 4.1993 | 0 | 0.70824 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2496 | 11.2613 | 2.5163 | 3.9445 | 4.1993 | 0.70824 | 0.22345 | 0.35027 | 0.3729 | 0.062892 | 1.5675 | 1.6688 | 0.28146 | 1.0646 | 0.17955 | 0.16866 | -1.4986 | -1.0491 | -0.98645 | -2.7663 | [10 empty] | 0 | 2.2496 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.01415 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 3.3176 | -3.4025 | -32.9452 | [8 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 101.1% | -6.4061 | -4.0978 | 17.5595 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 47 | iRM6 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.3845e-05; SAMPLE 49Ti40Ar: 3.4113e-05; SAMPLE 90Zr: 11.3102; SAMPLE 91Zr: 2.5273; SAMPLE 92Zr: 3.9618; SAMPLE 94Zr: 4.2178; SAMPLE 96Zr: 0.71142 | 1 | -6.3477e-05 | 4.6796e-05 | 11.31 | 2.5273 | 3.9616 | 0 | 4.2176 | 0 | 0.71133 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2505 | 11.31 | 2.5273 | 3.9616 | 4.2176 | 0.71133 | 0.22345 | 0.35028 | 0.37291 | 0.062894 | 1.5676 | 1.6689 | 0.28146 | 1.0646 | 0.17956 | 0.16866 | -1.4986 | -1.049 | -0.98641 | -2.7663 | [10 empty] | 0 | 2.2505 | 0.21795 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.01414 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 100.8% | 4.3298 | 4.0267 | -43.2574 | STD | Lab 3 | Jul 30, 2020 | 5 | 99.06% | 7.9592 | 0.089571 | -19.2518 | SRM 3169
Lab 3 | Jul 22, 2020 | 3 | 48 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 6.246e-06; SAMPLE 49Ti40Ar: 5.1353e-05; SAMPLE 90Zr: 11.1772; SAMPLE 91Zr: 2.4976; SAMPLE 92Zr: 3.9152; SAMPLE 94Zr: 4.1682; SAMPLE 96Zr: 0.7031 | 1 | -8.1076e-05 | 6.4035e-05 | 11.177 | 2.4975 | 3.9151 | 0 | 4.1681 | 0 | 0.70302 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2506 | 11.177 | 2.4975 | 3.9151 | 4.1681 | 0.70302 | 0.22345 | 0.35028 | 0.37292 | 0.062899 | 1.5676 | 1.6689 | 0.28149 | 1.0646 | 0.17957 | 0.16867 | -1.4986 | -1.049 | -0.9864 | -2.7662 | [10 empty] | 0 | 2.2506 | 0.21795 | 0.33335 | 0.3381 | 0.054382 | 1.5295 | 1.5513 | 0.24951 | 1.0142 | 0.16314 | 0.16085 | -1.5235 | -1.0986 | -2.9117 | 0.42492 | 0.43906 | -1.3882 | 0.014139 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.68% | -15.1423 | 13.3923 | 69.7256
Lab 3 | Jul 22, 2020 | 3 | 49 | iRM10 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 9.2237e-05; SAMPLE 49Ti40Ar: 2.1264e-05; SAMPLE 90Zr: 9.4829; SAMPLE 91Zr: 2.119; SAMPLE 92Zr: 3.3218; SAMPLE 94Zr: 3.5366; SAMPLE 96Zr: 0.59651 | 1 | 4.9144e-06 | 3.3947e-05 | 9.4827 | 2.119 | 3.3217 | 0 | 3.5364 | 0 | 0.59643 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2517 | 9.4827 | 2.119 | 3.3217 | 3.5364 | 0.59643 | 0.22346 | 0.35029 | 0.37293 | 0.062897 | 1.5676 | 1.6689 | 0.28147 | 1.0646 | 0.17956 | 0.16865 | -1.4985 | -1.049 | -0.98636 | -2.7663 | [10 empty] | 0 | 2.2517 | 0.21795 | 0.33335 | 0.3381 | 0.054377 | 1.5295 | 1.5513 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42492 | 0.43906 | -1.3883 | 0.01414 | -1.8133 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 22, 2020 | 3 | 84.89% | -11.6225 | 6.7419 | -53.7544
Lab 3 | Jul 22, 2020 | 3 | 50 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 1.3349e-05; SAMPLE 49Ti40Ar: 3.1063e-05; SAMPLE 90Zr: 11.1639; SAMPLE 91Zr: 2.4947; SAMPLE 92Zr: 3.9106; SAMPLE 94Zr: 4.1636; SAMPLE 96Zr: 0.70227 | 1 | -7.3974e-05 | 4.3745e-05 | 11.1637 | 2.4947 | 3.9105 | 0 | 4.1634 | 0 | 0.70219 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2524 | 11.1637 | 2.4947 | 3.9105 | 4.1634 | 0.70219 | 0.22346 | 0.35029 | 0.37294 | 0.0629 | 1.5675 | 1.6689 | 0.28148 | 1.0647 | 0.17957 | 0.16866 | -1.4985 | -1.049 | -0.98633 | -2.7662 | [10 empty] | 0 | 2.2524 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5294 | 1.5512 | 0.24948 | 1.0143 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42488 | 0.43904 | -1.3884 | 0.014155 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.3% | 24.2058 | -10.0698 | -79.2046
Lab 3 | Jul 22, 2020 | 3 | 51 | iRM11 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: -1.8335e-05; SAMPLE 49Ti40Ar: 4.4451e-05; SAMPLE 90Zr: 11.151; SAMPLE 91Zr: 2.4918; SAMPLE 92Zr: 3.9062; SAMPLE 94Zr: 4.1588; SAMPLE 96Zr: 0.70152 | 1 | -0.00010566 | 5.7133e-05 | 11.1508 | 2.4918 | 3.9061 | 0 | 4.1587 | 0 | 0.70144 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2526 | 11.1508 | 2.4918 | 3.9061 | 4.1587 | 0.70144 | 0.22346 | 0.35029 | 0.37295 | 0.062905 | 1.5676 | 1.669 | 0.2815 | 1.0647 | 0.17958 | 0.16867 | -1.4985 | -1.049 | -0.98632 | -2.7661 | [10 empty] | 0 | 2.2526 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42491 | 0.43905 | -1.3883 | 0.014137 | -1.8132 | -1.8273 | [14 empty] | iRM11 | Lab 3 | Jul 22, 2020 | 3 | 100.2% | 0.87435 | 15.8633 | 34.215
Lab 3 | Jul 22, 2020 | 3 | 52 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 9.0046e-06; SAMPLE 49Ti40Ar: 6.8002e-06; SAMPLE 90Zr: 11.0931; SAMPLE 91Zr: 2.4789; SAMPLE 92Zr: 3.8859; SAMPLE 94Zr: 4.1374; SAMPLE 96Zr: 0.69792 | 1 | -7.8318e-05 | 1.9482e-05 | 11.0929 | 2.4788 | 3.8858 | 0 | 4.1372 | 0 | 0.69784 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2536 | 11.0929 | 2.4788 | 3.8858 | 4.1372 | 0.69784 | 0.22346 | 0.3503 | 0.37296 | 0.062909 | 1.5676 | 1.669 | 0.28152 | 1.0647 | 0.17959 | 0.16867 | -1.4985 | -1.049 | -0.98627 | -2.7661 | [10 empty] | 0 | 2.2536 | 0.21795 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.24951 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42491 | 0.43906 | -1.3883 | 0.014151 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 100.1% | -24.4534 | -3.3746 | 24.2955
Lab 3 | Jul 22, 2020 | 3 | 53 | iRM5 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 4.2963e-05; SAMPLE 49Ti40Ar: 4.8988e-05; SAMPLE 90Zr: 11.1517; SAMPLE 91Zr: 2.492; SAMPLE 92Zr: 3.9066; SAMPLE 94Zr: 4.1594; SAMPLE 96Zr: 0.70164 | 1 | -4.4359e-05 | 6.1671e-05 | 11.1515 | 2.4919 | 3.9064 | 0 | 4.1593 | 0 | 0.70156 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2545 | 11.1515 | 2.4919 | 3.9064 | 4.1593 | 0.70156 | 0.22346 | 0.35031 | 0.37298 | 0.062912 | 1.5676 | 1.6691 | 0.28153 | 1.0647 | 0.17959 | 0.16868 | -1.4985 | -1.0489 | -0.98624 | -2.766 | [10 empty] | 0 | 2.2545 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.014145 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Jul 22, 2020 | 3 | 100.9% | -8.688 | 0.67367 | -9.6369
Lab 3 | Jul 22, 2020 | 3 | 54 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: -5.4469e-06; SAMPLE 49Ti40Ar: 2.0765e-05; SAMPLE 90Zr: 11.0023; SAMPLE 91Zr: 2.4587; SAMPLE 92Zr: 3.8543; SAMPLE 94Zr: 4.1038; SAMPLE 96Zr: 0.69228 | 1 | -9.2769e-05 | 3.3447e-05 | 11.0021 | 2.4586 | 3.8542 | 0 | 4.1036 | 0 | 0.6922 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2549 | 11.0021 | 2.4586 | 3.8542 | 4.1036 | 0.6922 | 0.22347 | 0.35031 | 0.37299 | 0.062915 | 1.5676 | 1.6691 | 0.28154 | 1.0647 | 0.1796 | 0.16868 | -1.4985 | -1.0489 | -0.98622 | -2.766 | [10 empty] | 0 | 2.2549 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9117 | 0.4249 | 0.43904 | -1.3883 | 0.01414 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 22, 2020 | 3 | 99.88% | 25.5306 | 12.1591 | 20.2636
Lab 3 | Jul 22, 2020 | 3 | 55 | iRM6 | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.1028e-05; SAMPLE 49Ti40Ar: 6.6802e-07; SAMPLE 90Zr: 11.0968; SAMPLE 91Zr: 2.4798; SAMPLE 92Zr: 3.8875; SAMPLE 94Zr: 4.1393; SAMPLE 96Zr: 0.69825 | 1 | -6.6295e-05 | 1.335e-05 | 11.0966 | 2.4798 | 3.8874 | 0 | 4.1391 | 0 | 0.69816 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2563 | 11.0966 | 2.4798 | 3.8874 | 4.1391 | 0.69816 | 0.22347 | 0.35032 | 0.37301 | 0.062917 | 1.5677 | 1.6692 | 0.28154 | 1.0647 | 0.1796 | 0.16867 | -1.4985 | -1.0489 | -0.98616 | -2.7659 | [10 empty] | 0 | 2.2563 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014132 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 22, 2020 | 3 | 101.2% | 3.5013 | 14.5076 | -41.3192
Lab 3 | Jul 22, 2020 | 3 | 56 | STD | OPZ 177Hf++: 8.7322e-05; OPZ 49Ti40Ar: -1.2682e-05; OPZ 90Zr: 0.00019603; OPZ 91Zr: 5.2133e-05; OPZ 92Zr: 0.00012668; OPZ 94Zr: 0.00017101; OPZ 96Zr: 8.0888e-05; SAMPLE 177Hf++: 2.0989e-05; SAMPLE 49Ti40Ar: 2.0123e-05; SAMPLE 90Zr: 10.9377; SAMPLE 91Zr: 2.4442; SAMPLE 92Zr: 3.8318; SAMPLE 94Zr: 4.08; SAMPLE 96Zr: 0.68828 | 1 | -6.6333e-05 | 3.2806e-05 | 10.9375 | 2.4442 | 3.8316 | 0 | 4.0799 | 0 | 0.6882 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2569 | 10.9375 | 2.4442 | 3.8316 | 4.0799 | 0.6882 | 0.22347 | 0.35032 | 0.37302 | 0.062921 | 1.5677 | 1.6692 | 0.28157 | 1.0648 | 0.17961 | 0.16868 | -1.4985 | -1.0489 | -0.98613 | -2.7659 | [10 empty] | 0 | 2.2569 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5513 | 0.2495 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.014153 | -1.8132 | -1.8273
Lab 3 | Jul 29, 2020 | 4 | 2 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -5.1852e-06; SAMPLE 49Ti40Ar: 4.8276e-05; SAMPLE 90Zr: 17.2141; SAMPLE 91Zr: 3.8451; SAMPLE 92Zr: 6.0251; SAMPLE 94Zr: 6.4097; SAMPLE 96Zr: 1.0802 | 1 | -0.0001363 | 5.3804e-05 | 17.214 | 3.845 | 6.025 | 0 | 6.4095 | 0 | 1.0802 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2154 | 17.214 | 3.845 | 6.025 | 6.4095 | 1.0802 | 0.22337 | 0.35001 | 0.37234 | 0.062749 | 1.567 | 1.667 | 0.28092 | 1.0638 | 0.17928 | 0.16852 | -1.4989 | -1.0498 | -0.98794 | -2.7686 | [9 empty] | L3 JulS4 | 0 | 2.2154 | 0.21795 | 0.33335 | 0.3381 | 0.054376 | 1.5295 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3884 | 0.014142 | -1.8133 | -1.8274
Lab 3 | Jul 29, 2020 | 4 | 3 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -2.1972e-05; SAMPLE 49Ti40Ar: 3.3638e-05; SAMPLE 90Zr: 16.9787; SAMPLE 91Zr: 3.7926; SAMPLE 92Zr: 5.9428; SAMPLE 94Zr: 6.3221; SAMPLE 96Zr: 1.0655 | 1 | -0.00015308 | 3.9165e-05 | 16.9786 | 3.7926 | 5.9427 | 0 | 6.322 | 0 | 1.0654 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2159 | 16.9786 | 3.7926 | 5.9427 | 6.322 | 1.0654 | 0.22337 | 0.35001 | 0.37235 | 0.062752 | 1.5669 | 1.6669 | 0.28093 | 1.0638 | 0.17929 | 0.16853 | -1.4989 | -1.0498 | -0.98792 | -2.7686 | [10 empty] | 0 | 2.2159 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5234 | -1.0986 | -2.9118 | 0.42489 | 0.43903 | -1.3884 | 0.01414 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 98.77% | 14.7698 | -3.8197 | 9.2462
Lab 3 | Jul 29, 2020 | 4 | 4 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -4.695e-05; SAMPLE 49Ti40Ar: 3.083e-05; SAMPLE 90Zr: 17.1661; SAMPLE 91Zr: 3.8344; SAMPLE 92Zr: 6.0084; SAMPLE 94Zr: 6.3919; SAMPLE 96Zr: 1.0773 | 1 | -0.00017806 | 3.6358e-05 | 17.166 | 3.8344 | 6.0083 | 0 | 6.3917 | 0 | 1.0772 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2158 | 17.166 | 3.8344 | 6.0083 | 6.3917 | 1.0772 | 0.22337 | 0.35001 | 0.37235 | 0.062752 | 1.567 | 1.6669 | 0.28093 | 1.0638 | 0.17928 | 0.16853 | -1.4989 | -1.0498 | -0.98792 | -2.7686 | [10 empty] | 0 | 2.2158 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014131 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.28% | 14.2258 | 11.3455 | -21.205
Lab 3 | Jul 29, 2020 | 4 | 5 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 5.5852e-06; SAMPLE 49Ti40Ar: 4.4545e-05; SAMPLE 90Zr: 17.2561; SAMPLE 91Zr: 3.8545; SAMPLE 92Zr: 6.0399; SAMPLE 94Zr: 6.4254; SAMPLE 96Zr: 1.083 | 1 | -0.00012553 | 5.0073e-05 | 17.2559 | 3.8545 | 6.0398 | 0 | 6.4252 | 0 | 1.0829 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2158 | 17.2559 | 3.8545 | 6.0398 | 6.4252 | 1.0829 | 0.22337 | 0.35001 | 0.37235 | 0.062755 | 1.567 | 1.667 | 0.28095 | 1.0638 | 0.17929 | 0.16854 | -1.4989 | -1.0498 | -0.98792 | -2.7685 | [10 empty] | 0 | 2.2158 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014135 | -1.8132 | -1.8273 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 99.94% | -1.924 | 3.1764 | 15.1135
Lab 3 | Jul 29, 2020 | 4 | 6 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.5986e-05; SAMPLE 49Ti40Ar: -2.1949e-06; SAMPLE 90Zr: 17.3666; SAMPLE 91Zr: 3.8792; SAMPLE 92Zr: 6.0785; SAMPLE 94Zr: 6.4666; SAMPLE 96Zr: 1.0899 | 1 | -0.0001471 | 3.3326e-06 | 17.3664 | 3.8791 | 6.0784 | 0 | 6.4664 | 0 | 1.0899 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2159 | 17.3664 | 3.8791 | 6.0784 | 6.4664 | 1.0899 | 0.22337 | 0.35001 | 0.37235 | 0.062757 | 1.567 | 1.667 | 0.28095 | 1.0638 | 0.1793 | 0.16854 | -1.4989 | -1.0498 | -0.98792 | -2.7685 | [10 empty] | 0 | 2.2159 | 0.21796 | 0.33335 | 0.3381 | 0.054382 | 1.5294 | 1.5512 | 0.24951 | 1.0142 | 0.16314 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.4249 | 0.43905 | -1.3883 | 0.014144 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.9% | -10.0618 | -8.4872 | 68.7184
Lab 3 | Jul 29, 2020 | 4 | 7 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 8.477e-05; SAMPLE 49Ti40Ar: 7.0473e-06; SAMPLE 90Zr: 17.5528; SAMPLE 91Zr: 3.9208; SAMPLE 92Zr: 6.1437; SAMPLE 94Zr: 6.536; SAMPLE 96Zr: 1.1016 | 1 | -4.6342e-05 | 1.2575e-05 | 17.5527 | 3.9208 | 6.1436 | 0 | 6.5358 | 0 | 1.1015 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2162 | 17.5527 | 3.9208 | 6.1436 | 6.5358 | 1.1015 | 0.22337 | 0.35001 | 0.37236 | 0.062754 | 1.5669 | 1.667 | 0.28094 | 1.0638 | 0.17929 | 0.16853 | -1.4989 | -1.0498 | -0.9879 | -2.7685 | [10 empty] | 0 | 2.2162 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014146 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 101.4% | -2.3395 | -4.4268 | -20.5948
Lab 3 | Jul 29, 2020 | 4 | 8 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -2.9908e-05; SAMPLE 49Ti40Ar: 2.1456e-05; SAMPLE 90Zr: 17.2503; SAMPLE 91Zr: 3.8533; SAMPLE 92Zr: 6.0379; SAMPLE 94Zr: 6.4233; SAMPLE 96Zr: 1.0826 | 1 | -0.00016102 | 2.6983e-05 | 17.2502 | 3.8532 | 6.0378 | 0 | 6.4232 | 0 | 1.0825 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2161 | 17.2502 | 3.8532 | 6.0378 | 6.4232 | 1.0825 | 0.22337 | 0.35001 | 0.37236 | 0.062753 | 1.5669 | 1.667 | 0.28094 | 1.0638 | 0.17929 | 0.16853 | -1.4989 | -1.0498 | -0.98791 | -2.7685 | [10 empty] | 0 | 2.2161 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.01414 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.1% | 7.4027 | 0.8896 | -30.8248
Lab 3 | Jul 29, 2020 | 4 | 9 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -8.9664e-06; SAMPLE 49Ti40Ar: 3.0708e-05; SAMPLE 90Zr: 17.0072; SAMPLE 91Zr: 3.799; SAMPLE 92Zr: 5.9529; SAMPLE 94Zr: 6.3329; SAMPLE 96Zr: 1.0674 | 1 | -0.00014008 | 3.6235e-05 | 17.0071 | 3.799 | 5.9527 | 0 | 6.3328 | 0 | 1.0673 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2165 | 17.0071 | 3.799 | 5.9527 | 6.3328 | 1.0673 | 0.22338 | 0.35002 | 0.37236 | 0.062756 | 1.5669 | 1.667 | 0.28094 | 1.0638 | 0.17929 | 0.16853 | -1.4989 | -1.0498 | -0.98789 | -2.7685 | [10 empty] | 0 | 2.2165 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9118 | 0.42489 | 0.43903 | -1.3883 | 0.014139 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 98.98% | 15.2714 | -0.65995 | 17.4744
Lab 3 | Jul 29, 2020 | 4 | 10 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.3653e-05; SAMPLE 49Ti40Ar: 2.4369e-05; SAMPLE 90Zr: 17.1136; SAMPLE 91Zr: 3.8227; SAMPLE 92Zr: 5.9901; SAMPLE 94Zr: 6.3725; SAMPLE 96Zr: 1.074 | 1 | -0.00014476 | 2.9896e-05 | 17.1135 | 3.8227 | 5.99 | 0 | 6.3724 | 0 | 1.0739 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2165 | 17.1135 | 3.8227 | 5.99 | 6.3724 | 1.0739 | 0.22337 | 0.35002 | 0.37236 | 0.062754 | 1.567 | 1.667 | 0.28094 | 1.0638 | 0.17929 | 0.16853 | -1.4989 | -1.0498 | -0.98789 | -2.7685 | [10 empty] | 0 | 2.2165 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014137 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.09% | -4.8772 | 0.51574 | -13.31
Lab 3 | Jul 29, 2020 | 4 | 11 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 1.3328e-05; SAMPLE 49Ti40Ar: 1.0392e-05; SAMPLE 90Zr: 17.1951; SAMPLE 91Zr: 3.841; SAMPLE 92Zr: 6.0187; SAMPLE 94Zr: 6.403; SAMPLE 96Zr: 1.0792 | 1 | -0.00011778 | 1.592e-05 | 17.1949 | 3.8409 | 6.0186 | 0 | 6.4028 | 0 | 1.0791 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2169 | 17.1949 | 3.8409 | 6.0186 | 6.4028 | 1.0791 | 0.22337 | 0.35002 | 0.37237 | 0.062757 | 1.567 | 1.667 | 0.28095 | 1.0638 | 0.17929 | 0.16854 | -1.4989 | -1.0498 | -0.98788 | -2.7685 | [10 empty] | 0 | 2.2169 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5234 | -1.0985 | -2.9118 | 0.42491 | 0.43903 | -1.3883 | 0.014129 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.96% | 8.2102 | 7.1171 | 7.2745
Lab 3 | Jul 29, 2020 | 4 | 12 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -2.8261e-06; SAMPLE 49Ti40Ar: 1.6869e-05; SAMPLE 90Zr: 17.2917; SAMPLE 91Zr: 3.8625; SAMPLE 92Zr: 6.0525; SAMPLE 94Zr: 6.4389; SAMPLE 96Zr: 1.0852 | 1 | -0.00013394 | 2.2396e-05 | 17.2916 | 3.8625 | 6.0524 | 0 | 6.4388 | 0 | 1.0852 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2167 | 17.2916 | 3.8625 | 6.0524 | 6.4388 | 1.0852 | 0.22337 | 0.35002 | 0.37236 | 0.062756 | 1.567 | 1.667 | 0.28095 | 1.0638 | 0.17929 | 0.16853 | -1.4989 | -1.0498 | -0.98788 | -2.7685 | [10 empty] | 0 | 2.2167 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014136 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 101.3% | 4.0171 | 1.3975 | 13.9083
Lab 3 | Jul 29, 2020 | 4 | 13 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 4.5928e-05; SAMPLE 49Ti40Ar: 2.1062e-06; SAMPLE 90Zr: 17.4057; SAMPLE 91Zr: 3.888; SAMPLE 92Zr: 6.0924; SAMPLE 94Zr: 6.4814; SAMPLE 96Zr: 1.0924 | 1 | -8.5184e-05 | 7.6338e-06 | 17.4056 | 3.8879 | 6.0923 | 0 | 6.4813 | 0 | 1.0923 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.217 | 17.4056 | 3.8879 | 6.0923 | 6.4813 | 1.0923 | 0.22337 | 0.35002 | 0.37237 | 0.062757 | 1.567 | 1.667 | 0.28095 | 1.0638 | 0.1793 | 0.16853 | -1.4989 | -1.0498 | -0.98787 | -2.7685 | [10 empty] | 0 | 2.217 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014134 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 101.5% | 0.27042 | 2.5772 | -2.1929
Lab 3 | Jul 29, 2020 | 4 | 14 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -7.7519e-06; SAMPLE 49Ti40Ar: 2.3608e-05; SAMPLE 90Zr: 17.0129; SAMPLE 91Zr: 3.8003; SAMPLE 92Zr: 5.955; SAMPLE 94Zr: 6.3353; SAMPLE 96Zr: 1.0678 | 1 | -0.00013886 | 2.9136e-05 | 17.0127 | 3.8002 | 5.9549 | 0 | 6.3352 | 0 | 1.0677 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2175 | 17.0127 | 3.8002 | 5.9549 | 6.3352 | 1.0677 | 0.22337 | 0.35002 | 0.37238 | 0.062759 | 1.567 | 1.6671 | 0.28096 | 1.0639 | 0.1793 | 0.16854 | -1.4989 | -1.0498 | -0.98785 | -2.7684 | [10 empty] | 0 | 2.2175 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.33% | -5.7183 | 1.4522 | -13.9916
Lab 3 | Jul 29, 2020 | 4 | 15 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -7.7026e-06; SAMPLE 49Ti40Ar: 8.7804e-06; SAMPLE 90Zr: 16.848; SAMPLE 91Zr: 3.7634; SAMPLE 92Zr: 5.8973; SAMPLE 94Zr: 6.274; SAMPLE 96Zr: 1.0575 | 1 | -0.00013881 | 1.4308e-05 | 16.8479 | 3.7634 | 5.8972 | 0 | 6.2738 | 0 | 1.0574 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2176 | 16.8479 | 3.7634 | 5.8972 | 6.2738 | 1.0574 | 0.22338 | 0.35003 | 0.37238 | 0.062761 | 1.567 | 1.6671 | 0.28097 | 1.0639 | 0.1793 | 0.16854 | -1.4989 | -1.0497 | -0.98784 | -2.7684 | [10 empty] | 0 | 2.2176 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014134 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.17% | 2.0473 | 5.6897 | 8.0355
Lab 3 | Jul 29, 2020 | 4 | 16 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -3.9631e-05; SAMPLE 49Ti40Ar: 1.073e-05; SAMPLE 90Zr: 16.9638; SAMPLE 91Zr: 3.7893; SAMPLE 92Zr: 5.9378; SAMPLE 94Zr: 6.3171; SAMPLE 96Zr: 1.0647 | 1 | -0.00017074 | 1.6258e-05 | 16.9637 | 3.7893 | 5.9377 | 0 | 6.317 | 0 | 1.0647 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2177 | 16.9637 | 3.7893 | 5.9377 | 6.317 | 1.0647 | 0.22338 | 0.35002 | 0.37238 | 0.062761 | 1.567 | 1.6671 | 0.28097 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0498 | -0.98784 | -2.7684 | [10 empty] | 0 | 2.2177 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014141 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.48% | -4.6024 | 2.5527 | 18.8966
Lab 3 | Jul 29, 2020 | 4 | 17 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -3.2385e-05; SAMPLE 49Ti40Ar: 5.1511e-05; SAMPLE 90Zr: 16.9852; SAMPLE 91Zr: 3.7941; SAMPLE 92Zr: 5.9453; SAMPLE 94Zr: 6.3252; SAMPLE 96Zr: 1.0661 | 1 | -0.0001635 | 5.7039e-05 | 16.9851 | 3.7941 | 5.9452 | 0 | 6.325 | 0 | 1.066 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.218 | 16.9851 | 3.7941 | 5.9452 | 6.325 | 1.066 | 0.22338 | 0.35003 | 0.37239 | 0.062763 | 1.567 | 1.6671 | 0.28097 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0497 | -0.98783 | -2.7684 | [10 empty] | 0 | 2.218 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9118 | 0.42489 | 0.43904 | -1.3883 | 0.014142 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 99.75% | -6.1204 | 3.963 | 18.7942
Lab 3 | Jul 29, 2020 | 4 | 18 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 3.2193e-06; SAMPLE 49Ti40Ar: 3.7052e-05; SAMPLE 90Zr: 17.0935; SAMPLE 91Zr: 3.8184; SAMPLE 92Zr: 5.9832; SAMPLE 94Zr: 6.3656; SAMPLE 96Zr: 1.0729 | 1 | -0.00012789 | 4.258e-05 | 17.0933 | 3.8183 | 5.9831 | 0 | 6.3654 | 0 | 1.0728 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2183 | 17.0933 | 3.8183 | 5.9831 | 6.3654 | 1.0728 | 0.22338 | 0.35003 | 0.37239 | 0.062763 | 1.5669 | 1.6671 | 0.28097 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0497 | -0.98781 | -2.7684 | [10 empty] | 0 | 2.2183 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24948 | 1.0143 | 0.16313 | 0.16083 | -1.5234 | -1.0986 | -2.9118 | 0.42487 | 0.43902 | -1.3884 | 0.014151 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.8% | 23.5674 | -11.2699 | -3.254
Lab 3 | Jul 29, 2020 | 4 | 19 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 4.2252e-05; SAMPLE 49Ti40Ar: 3.9418e-06; SAMPLE 90Zr: 17.3014; SAMPLE 91Zr: 3.8648; SAMPLE 92Zr: 6.0561; SAMPLE 94Zr: 6.4431; SAMPLE 96Zr: 1.086 | 1 | -8.886e-05 | 9.4693e-06 | 17.3013 | 3.8647 | 6.056 | 0 | 6.443 | 0 | 1.0859 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2188 | 17.3013 | 3.8647 | 6.056 | 6.443 | 1.0859 | 0.22338 | 0.35003 | 0.3724 | 0.062765 | 1.567 | 1.6671 | 0.28098 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0497 | -0.98779 | -2.7684 | [10 empty] | 0 | 2.2188 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014145 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 101.6% | -5.0379 | -0.87205 | 9.4707
Lab 3 | Jul 29, 2020 | 4 | 20 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 1.2429e-05; SAMPLE 49Ti40Ar: 4.177e-06; SAMPLE 90Zr: 16.9664; SAMPLE 91Zr: 3.7899; SAMPLE 92Zr: 5.9389; SAMPLE 94Zr: 6.3184; SAMPLE 96Zr: 1.065 | 1 | -0.00011868 | 9.7046e-06 | 16.9663 | 3.7899 | 5.9388 | 0 | 6.3182 | 0 | 1.0649 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2187 | 16.9663 | 3.7899 | 5.9388 | 6.3182 | 1.0649 | 0.22338 | 0.35003 | 0.3724 | 0.062764 | 1.567 | 1.6671 | 0.28098 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0497 | -0.98779 | -2.7684 | [10 empty] | 0 | 2.2187 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014137 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.83% | -19.3185 | 6.3704 | -21.9874
Lab 3 | Jul 29, 2020 | 4 | 21 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -2.5785e-05; SAMPLE 49Ti40Ar: 2.6558e-05; SAMPLE 90Zr: 16.7502; SAMPLE 91Zr: 3.7417; SAMPLE 92Zr: 5.8633; SAMPLE 94Zr: 6.238; SAMPLE 96Zr: 1.0514 | 1 | -0.0001569 | 3.2086e-05 | 16.7501 | 3.7416 | 5.8631 | 0 | 6.2378 | 0 | 1.0513 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2193 | 16.7501 | 3.7416 | 5.8631 | 6.2378 | 1.0513 | 0.22338 | 0.35004 | 0.37241 | 0.062766 | 1.567 | 1.6671 | 0.28099 | 1.0639 | 0.17931 | 0.16854 | -1.4989 | -1.0497 | -0.98777 | -2.7683 | [10 empty] | 0 | 2.2193 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014139 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 98.93% | 8.4657 | -1.6651 | -4.4758
Lab 3 | Jul 29, 2020 | 4 | 22 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.9958e-05; SAMPLE 49Ti40Ar: 3.2338e-05; SAMPLE 90Zr: 16.8961; SAMPLE 91Zr: 3.7742; SAMPLE 92Zr: 5.9143; SAMPLE 94Zr: 6.2923; SAMPLE 96Zr: 1.0606 | 1 | -0.00015107 | 3.7865e-05 | 16.896 | 3.7742 | 5.9142 | 0 | 6.2922 | 0 | 1.0605 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2193 | 16.896 | 3.7742 | 5.9142 | 6.2922 | 1.0605 | 0.22338 | 0.35004 | 0.37241 | 0.062768 | 1.567 | 1.6672 | 0.28099 | 1.0639 | 0.17932 | 0.16855 | -1.4989 | -1.0497 | -0.98777 | -2.7683 | [10 empty] | 0 | 2.2193 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.33% | 4.1333 | -2.806 | 25.4785
Lab 3 | Jul 29, 2020 | 4 | 23 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.7091e-05; SAMPLE 49Ti40Ar: 2.1575e-05; SAMPLE 90Zr: 16.9438; SAMPLE 91Zr: 3.7849; SAMPLE 92Zr: 5.931; SAMPLE 94Zr: 6.3101; SAMPLE 96Zr: 1.0636 | 1 | -0.0001482 | 2.7102e-05 | 16.9436 | 3.7848 | 5.9309 | 0 | 6.31 | 0 | 1.0635 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2195 | 16.9436 | 3.7848 | 5.9309 | 6.31 | 1.0635 | 0.22338 | 0.35004 | 0.37241 | 0.062769 | 1.567 | 1.6672 | 0.281 | 1.0639 | 0.17932 | 0.16855 | -1.4989 | -1.0497 | -0.98776 | -2.7683 | [10 empty] | 0 | 2.2195 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014139 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.81% | -2.978 | -4.8659 | 7.0344
Lab 3 | Jul 29, 2020 | 4 | 24 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.504e-05; SAMPLE 49Ti40Ar: 5.0954e-05; SAMPLE 90Zr: 17.0543; SAMPLE 91Zr: 3.8096; SAMPLE 92Zr: 5.9698; SAMPLE 94Zr: 6.3514; SAMPLE 96Zr: 1.0706 | 1 | -0.00014615 | 5.6481e-05 | 17.0542 | 3.8096 | 5.9697 | 0 | 6.3512 | 0 | 1.0705 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2198 | 17.0542 | 3.8096 | 5.9697 | 6.3512 | 1.0705 | 0.22338 | 0.35004 | 0.37241 | 0.062769 | 1.567 | 1.6672 | 0.281 | 1.0639 | 0.17932 | 0.16855 | -1.4989 | -1.0497 | -0.98775 | -2.7683 | [10 empty] | 0 | 2.2198 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014131 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 101.1% | 3.8899 | 10.5944 | -15.0945
Lab 3 | Jul 29, 2020 | 4 | 25 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 9.108e-05; SAMPLE 49Ti40Ar: 1.6537e-05; SAMPLE 90Zr: 17.2108; SAMPLE 91Zr: 3.8446; SAMPLE 92Zr: 6.0246; SAMPLE 94Zr: 6.4098; SAMPLE 96Zr: 1.0804 | 1 | -4.0032e-05 | 2.2064e-05 | 17.2107 | 3.8445 | 6.0245 | 0 | 6.4096 | 0 | 1.0803 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2201 | 17.2107 | 3.8445 | 6.0245 | 6.4096 | 1.0803 | 0.22338 | 0.35004 | 0.37242 | 0.062772 | 1.567 | 1.6672 | 0.28101 | 1.0639 | 0.17933 | 0.16855 | -1.4989 | -1.0497 | -0.98773 | -2.7682 | [10 empty] | 0 | 2.2201 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014145 | -1.8132 | -1.8273 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 101.5% | 1.8244 | -6.2862 | 22.0284
Lab 3 | Jul 29, 2020 | 4 | 26 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.2724e-05; SAMPLE 49Ti40Ar: -2.849e-06; SAMPLE 90Zr: 16.8431; SAMPLE 91Zr: 3.7624; SAMPLE 92Zr: 5.8959; SAMPLE 94Zr: 6.2729; SAMPLE 96Zr: 1.0574 | 1 | -0.00014384 | 2.6786e-06 | 16.8429 | 3.7624 | 5.8958 | 0 | 6.2727 | 0 | 1.0573 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2204 | 16.8429 | 3.7624 | 5.8958 | 6.2727 | 1.0573 | 0.22338 | 0.35004 | 0.37243 | 0.062772 | 1.567 | 1.6672 | 0.28101 | 1.0639 | 0.17933 | 0.16855 | -1.4989 | -1.0497 | -0.98772 | -2.7682 | [10 empty] | 0 | 2.2204 | 0.21796 | 0.33335 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014146 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.5% | -12.9156 | -11.1492 | 11.599
Lab 3 | Jul 29, 2020 | 4 | 27 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -4.331e-06; SAMPLE 49Ti40Ar: 6.1056e-06; SAMPLE 90Zr: 16.6206; SAMPLE 91Zr: 3.7128; SAMPLE 92Zr: 5.818; SAMPLE 94Zr: 6.1902; SAMPLE 96Zr: 1.0434 | 1 | -0.00013544 | 1.1633e-05 | 16.6204 | 3.7127 | 5.8179 | 0 | 6.19 | 0 | 1.0433 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2209 | 16.6204 | 3.7127 | 5.8179 | 6.19 | 1.0433 | 0.22338 | 0.35004 | 0.37243 | 0.062774 | 1.567 | 1.6672 | 0.28101 | 1.064 | 0.17933 | 0.16855 | -1.4989 | -1.0497 | -0.9877 | -2.7682 | [10 empty] | 0 | 2.2209 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0143 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42489 | 0.43905 | -1.3883 | 0.014155 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 98.8% | -1.51 | -12.3758 | -3.3747
Lab 3 | Jul 29, 2020 | 4 | 28 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -2.048e-05; SAMPLE 49Ti40Ar: 1.3929e-05; SAMPLE 90Zr: 16.8004; SAMPLE 91Zr: 3.753; SAMPLE 92Zr: 5.8811; SAMPLE 94Zr: 6.2572; SAMPLE 96Zr: 1.0547 | 1 | -0.00015159 | 1.9457e-05 | 16.8003 | 3.7529 | 5.881 | 0 | 6.257 | 0 | 1.0546 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2211 | 16.8003 | 3.7529 | 5.881 | 6.257 | 1.0546 | 0.22339 | 0.35005 | 0.37244 | 0.062774 | 1.567 | 1.6672 | 0.28101 | 1.0639 | 0.17933 | 0.16855 | -1.4989 | -1.0497 | -0.98769 | -2.7682 | [10 empty] | 0 | 2.2211 | 0.21796 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0986 | -2.9118 | 0.4249 | 0.43903 | -1.3883 | 0.014139 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.67% | 19.0927 | 5.524 | -5.1049
Lab 3 | Jul 29, 2020 | 4 | 29 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -2.4939e-05; SAMPLE 49Ti40Ar: 5.0296e-05; SAMPLE 90Zr: 16.8163; SAMPLE 91Zr: 3.7565; SAMPLE 92Zr: 5.8867; SAMPLE 94Zr: 6.2633; SAMPLE 96Zr: 1.0557 | 1 | -0.00015605 | 5.5824e-05 | 16.8162 | 3.7565 | 5.8866 | 0 | 6.2631 | 0 | 1.0556 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2216 | 16.8162 | 3.7565 | 5.8866 | 6.2631 | 1.0556 | 0.22338 | 0.35005 | 0.37244 | 0.062775 | 1.567 | 1.6673 | 0.28102 | 1.064 | 0.17933 | 0.16855 | -1.4989 | -1.0497 | -0.98767 | -2.7682 | [10 empty] | 0 | 2.2216 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3884 | 0.014146 | -1.8133 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 99.89% | -4.4182 | -5.9165 | -24.9231
Lab 3 | Jul 29, 2020 | 4 | 30 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.8806e-05; SAMPLE 49Ti40Ar: 1.5187e-05; SAMPLE 90Zr: 16.8697; SAMPLE 91Zr: 3.7684; SAMPLE 92Zr: 5.9054; SAMPLE 94Zr: 6.2832; SAMPLE 96Zr: 1.0591 | 1 | -0.00014992 | 2.0715e-05 | 16.8695 | 3.7684 | 5.9053 | 0 | 6.283 | 0 | 1.059 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2218 | 16.8695 | 3.7684 | 5.9053 | 6.283 | 1.059 | 0.22338 | 0.35006 | 0.37245 | 0.062777 | 1.5671 | 1.6673 | 0.28103 | 1.064 | 0.17933 | 0.16855 | -1.4989 | -1.0497 | -0.98766 | -2.7682 | [10 empty] | 0 | 2.2218 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.014142 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.9% | -10.9878 | 1.2578 | -14.4083
Lab 3 | Jul 29, 2020 | 4 | 31 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 3.5704e-05; SAMPLE 49Ti40Ar: 1.9561e-05; SAMPLE 90Zr: 17.0551; SAMPLE 91Zr: 3.8098; SAMPLE 92Zr: 5.9703; SAMPLE 94Zr: 6.3523; SAMPLE 96Zr: 1.0708 | 1 | -9.5408e-05 | 2.5089e-05 | 17.0549 | 3.8098 | 5.9702 | 0 | 6.3522 | 0 | 1.0707 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2221 | 17.0549 | 3.8098 | 5.9702 | 6.3522 | 1.0707 | 0.22338 | 0.35006 | 0.37245 | 0.062778 | 1.5671 | 1.6673 | 0.28103 | 1.064 | 0.17934 | 0.16855 | -1.4989 | -1.0497 | -0.98765 | -2.7681 | [10 empty] | 0 | 2.2221 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43905 | -1.3883 | 0.014142 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 101.8% | 0.93657 | 3.9495 | -13.3058
Lab 3 | Jul 29, 2020 | 4 | 32 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 8.1773e-06; SAMPLE 49Ti40Ar: -3.3507e-06; SAMPLE 90Zr: 16.6342; SAMPLE 91Zr: 3.7158; SAMPLE 92Zr: 5.823; SAMPLE 94Zr: 6.1957; SAMPLE 96Zr: 1.0444 | 1 | -0.00012293 | 2.1769e-06 | 16.6341 | 3.7158 | 5.8229 | 0 | 6.1955 | 0 | 1.0443 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2225 | 16.6341 | 3.7158 | 5.8229 | 6.1955 | 1.0443 | 0.22338 | 0.35006 | 0.37246 | 0.062781 | 1.5671 | 1.6673 | 0.28105 | 1.064 | 0.17935 | 0.16856 | -1.4989 | -1.0497 | -0.98763 | -2.7681 | [10 empty] | 0 | 2.2225 | 0.21795 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014149 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.53% | 0.20491 | -7.0959 | 4.4461
Lab 3 | Jul 29, 2020 | 4 | 33 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -4.1349e-05; SAMPLE 49Ti40Ar: 2.1592e-05; SAMPLE 90Zr: 16.3827; SAMPLE 91Zr: 3.6597; SAMPLE 92Zr: 5.7351; SAMPLE 94Zr: 6.1022; SAMPLE 96Zr: 1.0286 | 1 | -0.00017246 | 2.7119e-05 | 16.3826 | 3.6597 | 5.735 | 0 | 6.102 | 0 | 1.0286 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2232 | 16.3826 | 3.6597 | 5.735 | 6.102 | 1.0286 | 0.22339 | 0.35007 | 0.37247 | 0.062784 | 1.5671 | 1.6674 | 0.28105 | 1.064 | 0.17935 | 0.16856 | -1.4988 | -1.0496 | -0.9876 | -2.7681 | [10 empty] | 0 | 2.2232 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43904 | -1.3883 | 0.014137 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 98.72% | 15.1341 | 7.9647 | -14.0923
Lab 3 | Jul 29, 2020 | 4 | 34 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -1.7707e-05; SAMPLE 49Ti40Ar: 7.8456e-06; SAMPLE 90Zr: 16.5566; SAMPLE 91Zr: 3.6985; SAMPLE 92Zr: 5.796; SAMPLE 94Zr: 6.1669; SAMPLE 96Zr: 1.0396 | 1 | -0.00014882 | 1.3373e-05 | 16.5564 | 3.6985 | 5.7958 | 0 | 6.1667 | 0 | 1.0395 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2231 | 16.5564 | 3.6985 | 5.7958 | 6.1667 | 1.0395 | 0.22339 | 0.35007 | 0.37247 | 0.062785 | 1.5671 | 1.6674 | 0.28106 | 1.064 | 0.17935 | 0.16856 | -1.4989 | -1.0496 | -0.9876 | -2.768 | [10 empty] | 0 | 2.2231 | 0.21796 | 0.33335 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9117 | 0.42491 | 0.43905 | -1.3883 | 0.014141 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 99.19% | -5.208 | -0.58526 | 29.6219
Lab 3 | Jul 29, 2020 | 4 | 35 | iRM6 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 3.2272e-08; SAMPLE 49Ti40Ar: 2.1734e-05; SAMPLE 90Zr: 16.6272; SAMPLE 91Zr: 3.7143; SAMPLE 92Zr: 5.8207; SAMPLE 94Zr: 6.1933; SAMPLE 96Zr: 1.044 | 1 | -0.00013108 | 2.7262e-05 | 16.627 | 3.7143 | 5.8206 | 0 | 6.1931 | 0 | 1.0439 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2234 | 16.627 | 3.7143 | 5.8206 | 6.1931 | 1.0439 | 0.22339 | 0.35007 | 0.37247 | 0.062782 | 1.5671 | 1.6674 | 0.28105 | 1.064 | 0.17934 | 0.16856 | -1.4989 | -1.0496 | -0.98759 | -2.7681 | [10 empty] | 0 | 2.2234 | 0.21795 | 0.33335 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43906 | -1.3883 | 0.014148 | -1.8132 | -1.8274 | [14 empty] | iRM6 | Lab 3 | Jul 29, 2020 | 4 | 99.84% | -8.8498 | -11.6158 | -37.6179
Lab 3 | Jul 29, 2020 | 4 | 36 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: -2.6108e-05; SAMPLE 49Ti40Ar: 2.0104e-05; SAMPLE 90Zr: 16.7497; SAMPLE 91Zr: 3.7417; SAMPLE 92Zr: 5.8637; SAMPLE 94Zr: 6.2389; SAMPLE 96Zr: 1.0517 | 1 | -0.00015722 | 2.5631e-05 | 16.7496 | 3.7417 | 5.8636 | 0 | 6.2388 | 0 | 1.0516 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2235 | 16.7496 | 3.7417 | 5.8636 | 6.2388 | 1.0516 | 0.22339 | 0.35007 | 0.37247 | 0.062784 | 1.5671 | 1.6674 | 0.28105 | 1.064 | 0.17935 | 0.16856 | -1.4988 | -1.0496 | -0.98759 | -2.7681 | [10 empty] | 0 | 2.2235 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014132 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 29, 2020 | 4 | 100.9% | 7.479 | 8.2588 | -36.2426
Lab 3 | Jul 29, 2020 | 4 | 37 | iRM11 | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 8.04e-05; SAMPLE 49Ti40Ar: 3.0657e-05; SAMPLE 90Zr: 16.9997; SAMPLE 91Zr: 3.7976; SAMPLE 92Zr: 5.9511; SAMPLE 94Zr: 6.3321; SAMPLE 96Zr: 1.0674 | 1 | -5.0712e-05 | 3.6184e-05 | 16.9996 | 3.7975 | 5.951 | 0 | 6.3319 | 0 | 1.0673 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2236 | 16.9996 | 3.7975 | 5.951 | 6.3319 | 1.0673 | 0.22339 | 0.35007 | 0.37248 | 0.062786 | 1.5671 | 1.6674 | 0.28106 | 1.064 | 0.17935 | 0.16856 | -1.4988 | -1.0496 | -0.98758 | -2.768 | [10 empty] | 0 | 2.2236 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.4249 | 0.43905 | -1.3883 | 0.014146 | -1.8132 | -1.8274 | [14 empty] | iRM11 | Lab 3 | Jul 29, 2020 | 4 | 101.8% | -5.1086 | -10.6181 | 4.1433
Lab 3 | Jul 29, 2020 | 4 | 38 | STD | OPZ 177Hf++: 0.00013111; OPZ 49Ti40Ar: -5.5276e-06; OPZ 90Zr: 0.0001332; OPZ 91Zr: 4.7487e-05; OPZ 92Zr: 0.00011459; OPZ 94Zr: 0.0001565; OPZ 96Zr: 8.4959e-05; SAMPLE 177Hf++: 8.4812e-06; SAMPLE 49Ti40Ar: 4.2996e-05; SAMPLE 90Zr: 16.6345; SAMPLE 91Zr: 3.716; SAMPLE 92Zr: 5.8233; SAMPLE 94Zr: 6.1961; SAMPLE 96Zr: 1.0445 | 1 | -0.00012263 | 4.8523e-05 | 16.6343 | 3.7159 | 5.8232 | 0 | 6.1959 | 0 | 1.0444 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2236 | 16.6343 | 3.7159 | 5.8232 | 6.1959 | 1.0444 | 0.22339 | 0.35007 | 0.37248 | 0.062787 | 1.5671 | 1.6674 | 0.28106 | 1.064 | 0.17935 | 0.16857 | -1.4988 | -1.0496 | -0.98758 | -2.768 | [10 empty] | 0 | 2.2236 | 0.21796 | 0.33335 | 0.3381 | 0.05438 | 1.5294 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0986 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014139 | -1.8132 | -1.8273
Lab 3 | Jul 30, 2020 | 5 | 2 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 5.2276e-05; SAMPLE 49Ti40Ar: 2.3158e-06; SAMPLE 90Zr: 15.4744; SAMPLE 91Zr: 3.4578; SAMPLE 92Zr: 5.4203; SAMPLE 94Zr: 5.7702; SAMPLE 96Zr: 0.97315 | 1 | -4.3583e-05 | 4.9331e-06 | 15.4743 | 3.4577 | 5.4202 | 0 | 5.77 | 0 | 0.97308 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2483 | 15.4743 | 3.4577 | 5.4202 | 5.77 | 0.97308 | 0.22345 | 0.35027 | 0.37288 | 0.062884 | 1.5676 | 1.6687 | 0.28142 | 1.0645 | 0.17953 | 0.16864 | -1.4986 | -1.049 | -0.9865 | -2.7665 | [9 empty] | L3 JulS5 | 0 | 2.2483 | 0.21796 | 0.33336 | 0.3381 | 0.054377 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42494 | 0.43905 | -1.3883 | 0.014112 | -1.8133 | -1.8274
Lab 3 | Jul 30, 2020 | 5 | 3 | iRM4 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 9.4299e-06; SAMPLE 49Ti40Ar: -2.9697e-07; SAMPLE 90Zr: 15.7933; SAMPLE 91Zr: 3.5291; SAMPLE 92Zr: 5.532; SAMPLE 94Zr: 5.8892; SAMPLE 96Zr: 0.99327 | 1 | -8.6429e-05 | 2.3203e-06 | 15.7932 | 3.529 | 5.532 | 0 | 5.8891 | 0 | 0.9932 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.249 | 15.7932 | 3.529 | 5.532 | 5.8891 | 0.9932 | 0.22345 | 0.35027 | 0.37289 | 0.062888 | 1.5676 | 1.6688 | 0.28144 | 1.0646 | 0.17954 | 0.16865 | -1.4986 | -1.049 | -0.98648 | -2.7664 | [10 empty] | 0 | 2.249 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43904 | -1.3883 | 0.014115 | -1.8132 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.9% | 4.4623 | 0.22932 | 12.3283
Lab 3 | Jul 30, 2020 | 5 | 4 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 5.5806e-05; SAMPLE 49Ti40Ar: -1.4092e-05; SAMPLE 90Zr: 15.5298; SAMPLE 91Zr: 3.4702; SAMPLE 92Zr: 5.4397; SAMPLE 94Zr: 5.7911; SAMPLE 96Zr: 0.97673 | 1 | -4.0053e-05 | -1.1475e-05 | 15.5297 | 3.4702 | 5.4397 | 0 | 5.7909 | 0 | 0.97666 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2493 | 15.5297 | 3.4702 | 5.4397 | 5.7909 | 0.97666 | 0.22345 | 0.35028 | 0.37289 | 0.06289 | 1.5676 | 1.6688 | 0.28144 | 1.0646 | 0.17954 | 0.16865 | -1.4986 | -1.049 | -0.98646 | -2.7664 | [10 empty] | 0 | 2.2493 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014119 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.3% | 5.8025 | -3.3423 | 19.6871
Lab 3 | Jul 30, 2020 | 5 | 5 | iRM5 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: -8.856e-06; SAMPLE 49Ti40Ar: 1.1503e-07; SAMPLE 90Zr: 15.4313; SAMPLE 91Zr: 3.4482; SAMPLE 92Zr: 5.4052; SAMPLE 94Zr: 5.7543; SAMPLE 96Zr: 0.97055 | 1 | -0.00010472 | 2.7323e-06 | 15.4312 | 3.4482 | 5.4052 | 0 | 5.7542 | 0 | 0.97048 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2494 | 15.4312 | 3.4482 | 5.4052 | 5.7542 | 0.97048 | 0.22346 | 0.35028 | 0.37289 | 0.062891 | 1.5675 | 1.6688 | 0.28145 | 1.0646 | 0.17955 | 0.16866 | -1.4985 | -1.049 | -0.98646 | -2.7664 | [10 empty] | 0 | 2.2494 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9118 | 0.42492 | 0.43903 | -1.3883 | 0.014119 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.45% | 10.2236 | 0.31302 | 22.4311
Lab 3 | Jul 30, 2020 | 5 | 6 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 5.2587e-05; SAMPLE 49Ti40Ar: 3.8108e-05; SAMPLE 90Zr: 15.5022; SAMPLE 91Zr: 3.464; SAMPLE 92Zr: 5.4301; SAMPLE 94Zr: 5.7809; SAMPLE 96Zr: 0.97501 | 1 | -4.3272e-05 | 4.0725e-05 | 15.5021 | 3.464 | 5.4301 | 0 | 5.7807 | 0 | 0.97494 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2498 | 15.5021 | 3.464 | 5.4301 | 5.7807 | 0.97494 | 0.22345 | 0.35028 | 0.3729 | 0.062891 | 1.5676 | 1.6688 | 0.28145 | 1.0646 | 0.17955 | 0.16865 | -1.4986 | -1.049 | -0.98644 | -2.7664 | [10 empty] | 0 | 2.2498 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.014119 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -8.4781 | -0.67832 | -12.308
Lab 3 | Jul 30, 2020 | 5 | 7 | iRM10 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 9.1822e-06; SAMPLE 49Ti40Ar: -1.6891e-05; SAMPLE 90Zr: 15.667; SAMPLE 91Zr: 3.5009; SAMPLE 92Zr: 5.4879; SAMPLE 94Zr: 5.8425; SAMPLE 96Zr: 0.98541 | 1 | -8.6677e-05 | -1.4273e-05 | 15.6669 | 3.5009 | 5.4879 | 0 | 5.8424 | 0 | 0.98534 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2504 | 15.6669 | 3.5009 | 5.4879 | 5.8424 | 0.98534 | 0.22346 | 0.35028 | 0.37291 | 0.062893 | 1.5676 | 1.6688 | 0.28145 | 1.0646 | 0.17955 | 0.16865 | -1.4985 | -1.049 | -0.98641 | -2.7663 | [10 empty] | 0 | 2.2504 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5234 | -1.0985 | -2.9118 | 0.42491 | 0.43903 | -1.3883 | 0.014118 | -1.8133 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | 9.4989 | 0.2013 | -14.6993
Lab 3 | Jul 30, 2020 | 5 | 8 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.3405e-05; SAMPLE 49Ti40Ar: 6.078e-06; SAMPLE 90Zr: 15.359; SAMPLE 91Zr: 3.4321; SAMPLE 92Zr: 5.3801; SAMPLE 94Zr: 5.7277; SAMPLE 96Zr: 0.96608 | 1 | -7.2454e-05 | 8.6952e-06 | 15.3589 | 3.4321 | 5.38 | 0 | 5.7275 | 0 | 0.966 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2506 | 15.3589 | 3.4321 | 5.38 | 5.7275 | 0.966 | 0.22346 | 0.35029 | 0.37291 | 0.062895 | 1.5676 | 1.6688 | 0.28146 | 1.0646 | 0.17955 | 0.16866 | -1.4985 | -1.049 | -0.98641 | -2.7663 | [10 empty] | 0 | 2.2506 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9118 | 0.42492 | 0.43903 | -1.3883 | 0.014117 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.1% | 7.9021 | -1.896 | -4.5155
Lab 3 | Jul 30, 2020 | 5 | 9 | iRM10#2 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 5.8468e-05; SAMPLE 49Ti40Ar: -5.1881e-07; SAMPLE 90Zr: 15.4938; SAMPLE 91Zr: 3.4622; SAMPLE 92Zr: 5.4274; SAMPLE 94Zr: 5.7781; SAMPLE 96Zr: 0.97456 | 1 | -3.7391e-05 | 2.0985e-06 | 15.4937 | 3.4622 | 5.4273 | 0 | 5.7779 | 0 | 0.97448 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2511 | 15.4937 | 3.4622 | 5.4273 | 5.7779 | 0.97448 | 0.22346 | 0.35029 | 0.37292 | 0.062896 | 1.5676 | 1.6689 | 0.28146 | 1.0646 | 0.17955 | 0.16866 | -1.4985 | -1.049 | -0.98638 | -2.7663 | [10 empty] | 0 | 2.2511 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014117 | -1.8133 | -1.8274 | [14 empty] | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -2.785 | -2.1439 | -44.2701
Lab 3 | Jul 30, 2020 | 5 | 10 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 5.3135e-05; SAMPLE 49Ti40Ar: 2.3721e-05; SAMPLE 90Zr: 15.179; SAMPLE 91Zr: 3.3919; SAMPLE 92Zr: 5.3171; SAMPLE 94Zr: 5.6607; SAMPLE 96Zr: 0.95481 | 1 | -4.2724e-05 | 2.6338e-05 | 15.1789 | 3.3919 | 5.317 | 0 | 5.6605 | 0 | 0.95474 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.251 | 15.1789 | 3.3919 | 5.317 | 5.6605 | 0.95474 | 0.22346 | 0.35029 | 0.37292 | 0.062899 | 1.5676 | 1.6689 | 0.28148 | 1.0646 | 0.17956 | 0.16867 | -1.4985 | -1.049 | -0.98639 | -2.7662 | [10 empty] | 0 | 2.251 | 0.21796 | 0.33336 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42493 | 0.43904 | -1.3883 | 0.014112 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 99.01% | 5.5028 | 6.6802 | 34.6036
Lab 3 | Jul 30, 2020 | 5 | 11 | iRM4 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 1.7022e-05; SAMPLE 49Ti40Ar: 1.6097e-05; SAMPLE 90Zr: 15.4996; SAMPLE 91Zr: 3.4635; SAMPLE 92Zr: 5.4295; SAMPLE 94Zr: 5.7805; SAMPLE 96Zr: 0.97498 | 1 | -7.8837e-05 | 1.8714e-05 | 15.4995 | 3.4635 | 5.4294 | 0 | 5.7804 | 0 | 0.97491 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2521 | 15.4995 | 3.4635 | 5.4294 | 5.7804 | 0.97491 | 0.22346 | 0.3503 | 0.37294 | 0.062899 | 1.5676 | 1.6689 | 0.28148 | 1.0646 | 0.17956 | 0.16866 | -1.4985 | -1.049 | -0.98634 | -2.7662 | [10 empty] | 0 | 2.2521 | 0.21795 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42494 | 0.43905 | -1.3883 | 0.014118 | -1.8133 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.7% | -8.2372 | -2.1133 | -43.6591
Lab 3 | Jul 30, 2020 | 5 | 12 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 1.535e-05; SAMPLE 49Ti40Ar: 2.3172e-05; SAMPLE 90Zr: 15.3017; SAMPLE 91Zr: 3.4193; SAMPLE 92Zr: 5.3602; SAMPLE 94Zr: 5.7067; SAMPLE 96Zr: 0.96255 | 1 | -8.0509e-05 | 2.579e-05 | 15.3016 | 3.4193 | 5.3601 | 0 | 5.7065 | 0 | 0.96248 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.252 | 15.3016 | 3.4193 | 5.3601 | 5.7065 | 0.96248 | 0.22346 | 0.3503 | 0.37294 | 0.0629 | 1.5676 | 1.6689 | 0.28149 | 1.0646 | 0.17956 | 0.16866 | -1.4985 | -1.049 | -0.98634 | -2.7662 | [10 empty] | 0 | 2.252 | 0.21795 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.01412 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.8% | -13.1029 | 2.8269 | -7.7566
Lab 3 | Jul 30, 2020 | 5 | 13 | iRM5 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 4.8952e-06; SAMPLE 49Ti40Ar: 2.2435e-05; SAMPLE 90Zr: 15.1222; SAMPLE 91Zr: 3.3792; SAMPLE 92Zr: 5.2973; SAMPLE 94Zr: 5.6398; SAMPLE 96Zr: 0.95125 | 1 | -9.0964e-05 | 2.5053e-05 | 15.1221 | 3.3792 | 5.2972 | 0 | 5.6397 | 0 | 0.95117 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2522 | 15.1221 | 3.3792 | 5.2972 | 5.6397 | 0.95117 | 0.22346 | 0.3503 | 0.37294 | 0.0629 | 1.5676 | 1.6689 | 0.28148 | 1.0646 | 0.17956 | 0.16866 | -1.4985 | -1.049 | -0.98634 | -2.7662 | [10 empty] | 0 | 2.2522 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.014122 | -1.8133 | -1.8274 | [14 empty] | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.2% | 0.50305 | 5.3597 | -11.2572
Lab 3 | Jul 30, 2020 | 5 | 14 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: -1.6858e-06; SAMPLE 49Ti40Ar: 2.2335e-05; SAMPLE 90Zr: 15.1872; SAMPLE 91Zr: 3.3938; SAMPLE 92Zr: 5.3201; SAMPLE 94Zr: 5.6642; SAMPLE 96Zr: 0.95538 | 1 | -9.7545e-05 | 2.4953e-05 | 15.1871 | 3.3937 | 5.32 | 0 | 5.6641 | 0 | 0.95531 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.253 | 15.1871 | 3.3937 | 5.32 | 5.6641 | 0.95531 | 0.22346 | 0.3503 | 0.37295 | 0.062903 | 1.5676 | 1.669 | 0.28149 | 1.0647 | 0.17957 | 0.16866 | -1.4985 | -1.049 | -0.9863 | -2.7662 | [10 empty] | 0 | 2.253 | 0.21796 | 0.33335 | 0.3381 | 0.054377 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3884 | 0.014134 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100% | 13.4545 | -16.3101 | -47.1196
Lab 3 | Jul 30, 2020 | 5 | 15 | iRM10 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: -1.7799e-06; SAMPLE 49Ti40Ar: 3.19e-06; SAMPLE 90Zr: 15.3425; SAMPLE 91Zr: 3.4285; SAMPLE 92Zr: 5.3746; SAMPLE 94Zr: 5.7222; SAMPLE 96Zr: 0.96521 | 1 | -9.7639e-05 | 5.8073e-06 | 15.3424 | 3.4285 | 5.3745 | 0 | 5.7221 | 0 | 0.96514 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2533 | 15.3424 | 3.4285 | 5.3745 | 5.7221 | 0.96514 | 0.22346 | 0.3503 | 0.37296 | 0.062906 | 1.5676 | 1.669 | 0.28151 | 1.0647 | 0.17958 | 0.16867 | -1.4985 | -1.049 | -0.98629 | -2.7661 | [10 empty] | 0 | 2.2533 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42491 | 0.43904 | -1.3883 | 0.014129 | -1.8132 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.4% | 6.8667 | -4.3113 | 8.0067
Lab 3 | Jul 30, 2020 | 5 | 16 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 6.1753e-05; SAMPLE 49Ti40Ar: 2.5997e-05; SAMPLE 90Zr: 15.0683; SAMPLE 91Zr: 3.3672; SAMPLE 92Zr: 5.2786; SAMPLE 94Zr: 5.62; SAMPLE 96Zr: 0.94799 | 1 | -3.4106e-05 | 2.8615e-05 | 15.0682 | 3.3671 | 5.2785 | 0 | 5.6198 | 0 | 0.94792 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2534 | 15.0682 | 3.3671 | 5.2785 | 5.6198 | 0.94792 | 0.22346 | 0.35031 | 0.37296 | 0.062908 | 1.5677 | 1.669 | 0.28152 | 1.0647 | 0.17958 | 0.16867 | -1.4985 | -1.0489 | -0.98628 | -2.7661 | [10 empty] | 0 | 2.2534 | 0.21795 | 0.33336 | 0.3381 | 0.054381 | 1.5295 | 1.5512 | 0.24951 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9117 | 0.42494 | 0.43906 | -1.3883 | 0.014115 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 99.96% | -17.29 | 11.4632 | 78.321
Lab 3 | Jul 30, 2020 | 5 | 17 | iRM10#2 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 6.8904e-05; SAMPLE 49Ti40Ar: 8.2227e-06; SAMPLE 90Zr: 15.1991; SAMPLE 91Zr: 3.3964; SAMPLE 92Zr: 5.3244; SAMPLE 94Zr: 5.6688; SAMPLE 96Zr: 0.95619 | 1 | -2.6955e-05 | 1.084e-05 | 15.1989 | 3.3964 | 5.3243 | 0 | 5.6687 | 0 | 0.95612 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2537 | 15.1989 | 3.3964 | 5.3243 | 5.6687 | 0.95612 | 0.22346 | 0.35031 | 0.37296 | 0.062907 | 1.5676 | 1.669 | 0.28151 | 1.0647 | 0.17958 | 0.16867 | -1.4985 | -1.0489 | -0.98627 | -2.7661 | [10 empty] | 0 | 2.2537 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.014123 | -1.8132 | -1.8274 | [14 empty] | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.2% | -3.7305 | -5.5205 | -0.38315
Lab 3 | Jul 30, 2020 | 5 | 18 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 4.0649e-05; SAMPLE 49Ti40Ar: 3.3557e-05; SAMPLE 90Zr: 14.9614; SAMPLE 91Zr: 3.3434; SAMPLE 92Zr: 5.2412; SAMPLE 94Zr: 5.5803; SAMPLE 96Zr: 0.94122 | 1 | -5.521e-05 | 3.6175e-05 | 14.9613 | 3.3433 | 5.2411 | 0 | 5.5801 | 0 | 0.94115 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2541 | 14.9613 | 3.3433 | 5.2411 | 5.5801 | 0.94115 | 0.22347 | 0.35031 | 0.37297 | 0.062906 | 1.5676 | 1.669 | 0.2815 | 1.0647 | 0.17957 | 0.16866 | -1.4985 | -1.0489 | -0.98625 | -2.7661 | [10 empty] | 0 | 2.2541 | 0.21796 | 0.33336 | 0.3381 | 0.054376 | 1.5295 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3884 | 0.014119 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 99.41% | 6.4852 | 1.5307 | -74.3031
Lab 3 | Jul 30, 2020 | 5 | 19 | iRM4 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.7807e-06; SAMPLE 49Ti40Ar: 7.3844e-06; SAMPLE 90Zr: 15.2163; SAMPLE 91Zr: 3.4003; SAMPLE 92Zr: 5.3305; SAMPLE 94Zr: 5.6754; SAMPLE 96Zr: 0.95732 | 1 | -9.3078e-05 | 1.0002e-05 | 15.2162 | 3.4003 | 5.3305 | 0 | 5.6753 | 0 | 0.95725 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2545 | 15.2162 | 3.4003 | 5.3305 | 5.6753 | 0.95725 | 0.22347 | 0.35031 | 0.37298 | 0.06291 | 1.5676 | 1.6691 | 0.28152 | 1.0647 | 0.17958 | 0.16867 | -1.4985 | -1.0489 | -0.98623 | -2.766 | [10 empty] | 0 | 2.2545 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014121 | -1.8133 | -1.8274 | [14 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -7.3172 | 1.2003 | 12.4733
Lab 3 | Jul 30, 2020 | 5 | 20 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 1.7321e-05; SAMPLE 49Ti40Ar: -3.1048e-05; SAMPLE 90Zr: 15.0335; SAMPLE 91Zr: 3.3595; SAMPLE 92Zr: 5.2665; SAMPLE 94Zr: 5.6073; SAMPLE 96Zr: 0.94586 | 1 | -7.8538e-05 | -2.8431e-05 | 15.0334 | 3.3595 | 5.2664 | 0 | 5.6072 | 0 | 0.94579 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2548 | 15.0334 | 3.3595 | 5.2664 | 5.6072 | 0.94579 | 0.22347 | 0.35032 | 0.37298 | 0.062912 | 1.5676 | 1.6691 | 0.28153 | 1.0647 | 0.17959 | 0.16867 | -1.4985 | -1.0489 | -0.98622 | -2.766 | [10 empty] | 0 | 2.2548 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16313 | 0.16084 | -1.5234 | -1.0985 | -2.9118 | 0.42491 | 0.43903 | -1.3883 | 0.014126 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.3% | 12.1667 | -6.7645 | 39.1375
Lab 3 | Jul 30, 2020 | 5 | 21 | iRM5 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 7.42e-06; SAMPLE 49Ti40Ar: 2.4377e-05; SAMPLE 90Zr: 14.9044; SAMPLE 91Zr: 3.3307; SAMPLE 92Zr: 5.2213; SAMPLE 94Zr: 5.5592; SAMPLE 96Zr: 0.93772 | 1 | -8.8439e-05 | 2.6995e-05 | 14.9043 | 3.3306 | 5.2212 | 0 | 5.5591 | 0 | 0.93765 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2549 | 14.9043 | 3.3306 | 5.2212 | 5.5591 | 0.93765 | 0.22347 | 0.35032 | 0.37298 | 0.062911 | 1.5676 | 1.6691 | 0.28152 | 1.0647 | 0.17958 | 0.16867 | -1.4985 | -1.0489 | -0.98622 | -2.766 | [10 empty] | 0 | 2.2549 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014126 | -1.8133 | -1.8274 | [14 empty] | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.23% | -0.53785 | -3.4229 | -13.3367
Lab 3 | Jul 30, 2020 | 5 | 22 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.4547e-05; SAMPLE 49Ti40Ar: -8.2076e-06; SAMPLE 90Zr: 15.0076; SAMPLE 91Zr: 3.3537; SAMPLE 92Zr: 5.2575; SAMPLE 94Zr: 5.5979; SAMPLE 96Zr: 0.94425 | 1 | -7.1312e-05 | -5.5903e-06 | 15.0075 | 3.3537 | 5.2575 | 0 | 5.5977 | 0 | 0.94418 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2555 | 15.0075 | 3.3537 | 5.2575 | 5.5977 | 0.94418 | 0.22347 | 0.35032 | 0.37299 | 0.062914 | 1.5677 | 1.6691 | 0.28154 | 1.0647 | 0.17959 | 0.16867 | -1.4985 | -1.0489 | -0.98619 | -2.766 | [10 empty] | 0 | 2.2555 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.01412 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -12.1159 | 5.4901 | -3.8008
Lab 3 | Jul 30, 2020 | 5 | 23 | iRM10 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 1.6277e-05; SAMPLE 49Ti40Ar: 4.6135e-06; SAMPLE 90Zr: 15.1501; SAMPLE 91Zr: 3.3856; SAMPLE 92Zr: 5.3075; SAMPLE 94Zr: 5.651; SAMPLE 96Zr: 0.95326 | 1 | -7.9582e-05 | 7.2307e-06 | 15.15 | 3.3856 | 5.3074 | 0 | 5.6509 | 0 | 0.95318 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2556 | 15.15 | 3.3856 | 5.3074 | 5.6509 | 0.95318 | 0.22347 | 0.35032 | 0.373 | 0.062917 | 1.5677 | 1.6691 | 0.28154 | 1.0647 | 0.1796 | 0.16868 | -1.4985 | -1.0489 | -0.98619 | -2.7659 | [10 empty] | 0 | 2.2556 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014118 | -1.8132 | -1.8273 | [14 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | 9.2559 | 3.9726 | 39.4407
Lab 3 | Jul 30, 2020 | 5 | 24 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: -2.7489e-06; SAMPLE 49Ti40Ar: 1.236e-05; SAMPLE 90Zr: 14.8523; SAMPLE 91Zr: 3.3191; SAMPLE 92Zr: 5.2032; SAMPLE 94Zr: 5.54; SAMPLE 96Zr: 0.9345 | 1 | -9.8608e-05 | 1.4978e-05 | 14.8522 | 3.319 | 5.2031 | 0 | 5.5399 | 0 | 0.93443 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.256 | 14.8522 | 3.319 | 5.2031 | 5.5399 | 0.93443 | 0.22347 | 0.35033 | 0.373 | 0.062915 | 1.5677 | 1.6691 | 0.28154 | 1.0647 | 0.17959 | 0.16867 | -1.4985 | -1.0489 | -0.98617 | -2.766 | [10 empty] | 0 | 2.256 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014124 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.2% | -5.4377 | -1.6996 | -9.5743
Lab 3 | Jul 30, 2020 | 5 | 25 | iRM10#2 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 8.9474e-05; SAMPLE 49Ti40Ar: 1.5613e-05; SAMPLE 90Zr: 14.9663; SAMPLE 91Zr: 3.3445; SAMPLE 92Zr: 5.2432; SAMPLE 94Zr: 5.5827; SAMPLE 96Zr: 0.94172 | 1 | -6.3846e-06 | 1.8231e-05 | 14.9662 | 3.3445 | 5.2432 | 0 | 5.5825 | 0 | 0.94165 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2564 | 14.9662 | 3.3445 | 5.2432 | 5.5825 | 0.94165 | 0.22347 | 0.35033 | 0.37301 | 0.062918 | 1.5677 | 1.6692 | 0.28155 | 1.0647 | 0.1796 | 0.16868 | -1.4985 | -1.0489 | -0.98615 | -2.7659 | [10 empty] | 0 | 2.2564 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42494 | 0.43905 | -1.3883 | 0.014112 | -1.8133 | -1.8274 | [14 empty] | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -12.6969 | 13.2978 | 13.0571
Lab 3 | Jul 30, 2020 | 5 | 26 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 3.1643e-05; SAMPLE 49Ti40Ar: 1.1025e-05; SAMPLE 90Zr: 14.6421; SAMPLE 91Zr: 3.2722; SAMPLE 92Zr: 5.1296; SAMPLE 94Zr: 5.4618; SAMPLE 96Zr: 0.92132 | 1 | -6.4216e-05 | 1.3642e-05 | 14.642 | 3.2721 | 5.1295 | 0 | 5.4616 | 0 | 0.92125 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2566 | 14.642 | 3.2721 | 5.1295 | 5.4616 | 0.92125 | 0.22347 | 0.35033 | 0.37301 | 0.062918 | 1.5676 | 1.6691 | 0.28155 | 1.0647 | 0.1796 | 0.16868 | -1.4985 | -1.0489 | -0.98614 | -2.7659 | [10 empty] | 0 | 2.2566 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5294 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5234 | -1.0985 | -2.9118 | 0.4249 | 0.43903 | -1.3883 | 0.014126 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 98.84% | 14.0945 | -3.4966 | -11.9222
Lab 3 | Jul 30, 2020 | 5 | 27 | iRM4 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 3.0736e-05; SAMPLE 49Ti40Ar: 2.3713e-05; SAMPLE 90Zr: 14.9188; SAMPLE 91Zr: 3.334; SAMPLE 92Zr: 5.2266; SAMPLE 94Zr: 5.5651; SAMPLE 96Zr: 0.93879 | 1 | -6.5123e-05 | 2.6331e-05 | 14.9187 | 3.3339 | 5.2265 | 0 | 5.5649 | 0 | 0.93872 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.257 | 14.9187 | 3.3339 | 5.2265 | 5.5649 | 0.93872 | 0.22347 | 0.35034 | 0.37302 | 0.062922 | 1.5677 | 1.6692 | 0.28156 | 1.0647 | 0.17961 | 0.16868 | -1.4985 | -1.0489 | -0.98613 | -2.7659 | [10 empty] | 0 | 2.257 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014116 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 101.4% | -3.3261 | 6.3048 | 20.4419
Lab 3 | Jul 30, 2020 | 5 | 28 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.881e-05; SAMPLE 49Ti40Ar: -1.0402e-05; SAMPLE 90Zr: 14.7753; SAMPLE 91Zr: 3.3019; SAMPLE 92Zr: 5.1763; SAMPLE 94Zr: 5.5115; SAMPLE 96Zr: 0.92975 | 1 | -6.7049e-05 | -7.7852e-06 | 14.7752 | 3.3018 | 5.1762 | 0 | 5.5114 | 0 | 0.92968 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2569 | 14.7752 | 3.3018 | 5.1762 | 5.5114 | 0.92968 | 0.22347 | 0.35033 | 0.37302 | 0.062922 | 1.5677 | 1.6692 | 0.28156 | 1.0647 | 0.1796 | 0.16868 | -1.4985 | -1.0489 | -0.98613 | -2.7659 | [10 empty] | 0 | 2.2569 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.01412 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -2.3071 | 0.22484 | 17.2229
Lab 3 | Jul 30, 2020 | 5 | 29 | iRM5 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 1.3309e-05; SAMPLE 49Ti40Ar: 2.8047e-05; SAMPLE 90Zr: 14.6295; SAMPLE 91Zr: 3.2693; SAMPLE 92Zr: 5.1253; SAMPLE 94Zr: 5.4573; SAMPLE 96Zr: 0.92059 | 1 | -8.255e-05 | 3.0664e-05 | 14.6294 | 3.2693 | 5.1252 | 0 | 5.4571 | 0 | 0.92052 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2575 | 14.6294 | 3.2693 | 5.1252 | 5.4571 | 0.92052 | 0.22347 | 0.35034 | 0.37303 | 0.062922 | 1.5677 | 1.6692 | 0.28156 | 1.0648 | 0.1796 | 0.16868 | -1.4985 | -1.0489 | -0.98611 | -2.7659 | [10 empty] | 0 | 2.2575 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014118 | -1.8132 | -1.8274 | [14 empty] | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 98.98% | 4.3124 | -1.4287 | -19.4449
Lab 3 | Jul 30, 2020 | 5 | 30 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: -5.605e-07; SAMPLE 49Ti40Ar: 4.5361e-06; SAMPLE 90Zr: 14.7845; SAMPLE 91Zr: 3.304; SAMPLE 92Zr: 5.1797; SAMPLE 94Zr: 5.5152; SAMPLE 96Zr: 0.93038 | 1 | -9.6419e-05 | 7.1534e-06 | 14.7844 | 3.3039 | 5.1796 | 0 | 5.515 | 0 | 0.93031 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2578 | 14.7844 | 3.3039 | 5.1796 | 5.515 | 0.93031 | 0.22347 | 0.35034 | 0.37303 | 0.062925 | 1.5677 | 1.6692 | 0.28158 | 1.0648 | 0.17961 | 0.16869 | -1.4985 | -1.0488 | -0.98609 | -2.7658 | [10 empty] | 0 | 2.2578 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.014114 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.5% | -2.0475 | 6.1145 | -4.7415
Lab 3 | Jul 30, 2020 | 5 | 31 | iRM10 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.8307e-05; SAMPLE 49Ti40Ar: 5.2496e-05; SAMPLE 90Zr: 14.9432; SAMPLE 91Zr: 3.3394; SAMPLE 92Zr: 5.2353; SAMPLE 94Zr: 5.5744; SAMPLE 96Zr: 0.94038 | 1 | -6.7552e-05 | 5.5114e-05 | 14.9431 | 3.3394 | 5.2352 | 0 | 5.5743 | 0 | 0.94031 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2581 | 14.9431 | 3.3394 | 5.2352 | 5.5743 | 0.94031 | 0.22347 | 0.35034 | 0.37304 | 0.062926 | 1.5677 | 1.6693 | 0.28158 | 1.0648 | 0.17961 | 0.16869 | -1.4985 | -1.0488 | -0.98608 | -2.7658 | [10 empty] | 0 | 2.2581 | 0.21795 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42494 | 0.43906 | -1.3883 | 0.014121 | -1.8132 | -1.8273 | [14 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 101.5% | -9.093 | -3.5471 | -3.1532
Lab 3 | Jul 30, 2020 | 5 | 32 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.8057e-05; SAMPLE 49Ti40Ar: -1.4098e-05; SAMPLE 90Zr: 14.6545; SAMPLE 91Zr: 3.2749; SAMPLE 92Zr: 5.1342; SAMPLE 94Zr: 5.4668; SAMPLE 96Zr: 0.92224 | 1 | -6.7802e-05 | -1.148e-05 | 14.6544 | 3.2749 | 5.1341 | 0 | 5.4667 | 0 | 0.92217 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2583 | 14.6544 | 3.2749 | 5.1341 | 5.4667 | 0.92217 | 0.22347 | 0.35034 | 0.37304 | 0.062927 | 1.5677 | 1.6693 | 0.28159 | 1.0648 | 0.17962 | 0.16869 | -1.4985 | -1.0488 | -0.98607 | -2.7658 | [10 empty] | 0 | 2.2583 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.014121 | -1.8132 | -1.8273 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 100.1% | -5.6146 | -6.1394 | 8.1469
Lab 3 | Jul 30, 2020 | 5 | 33 | iRM10#2 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 5.8268e-05; SAMPLE 49Ti40Ar: -7.1836e-06; SAMPLE 90Zr: 14.7694; SAMPLE 91Zr: 3.3006; SAMPLE 92Zr: 5.1745; SAMPLE 94Zr: 5.5097; SAMPLE 96Zr: 0.92945 | 1 | -3.7591e-05 | -4.5663e-06 | 14.7693 | 3.3006 | 5.1744 | 0 | 5.5096 | 0 | 0.92938 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2584 | 14.7693 | 3.3006 | 5.1744 | 5.5096 | 0.92938 | 0.22348 | 0.35035 | 0.37304 | 0.062926 | 1.5677 | 1.6693 | 0.28158 | 1.0648 | 0.17961 | 0.16868 | -1.4985 | -1.0488 | -0.98606 | -2.7658 | [10 empty] | 0 | 2.2584 | 0.21796 | 0.33336 | 0.3381 | 0.054379 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43904 | -1.3883 | 0.014117 | -1.8133 | -1.8274 | [14 empty] | iRM10#2 | Lab 3 | Jul 30, 2020 | 5 | 101.4% | 1.2449 | 1.7627 | -24.6316
Lab 3 | Jul 30, 2020 | 5 | 34 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 4.4587e-05; SAMPLE 49Ti40Ar: 1.6778e-05; SAMPLE 90Zr: 14.4875; SAMPLE 91Zr: 3.2376; SAMPLE 92Zr: 5.0757; SAMPLE 94Zr: 5.4046; SAMPLE 96Zr: 0.91174 | 1 | -5.1272e-05 | 1.9395e-05 | 14.4874 | 3.2376 | 5.0756 | 0 | 5.4045 | 0 | 0.91167 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2588 | 14.4874 | 3.2376 | 5.0756 | 5.4045 | 0.91167 | 0.22348 | 0.35035 | 0.37305 | 0.062929 | 1.5677 | 1.6693 | 0.28159 | 1.0648 | 0.17962 | 0.16869 | -1.4984 | -1.0488 | -0.98605 | -2.7658 | [10 empty] | 0 | 2.2588 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43904 | -1.3883 | 0.014116 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 99.05% | 7.3317 | 6.3134 | -4.173
Lab 3 | Jul 30, 2020 | 5 | 35 | iRM4 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 1.6857e-05; SAMPLE 49Ti40Ar: 1.9294e-05; SAMPLE 90Zr: 14.6004; SAMPLE 91Zr: 3.2629; SAMPLE 92Zr: 5.1153; SAMPLE 94Zr: 5.4469; SAMPLE 96Zr: 0.91889 | 1 | -7.9002e-05 | 2.1912e-05 | 14.6003 | 3.2628 | 5.1153 | 0 | 5.4467 | 0 | 0.91882 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2594 | 14.6003 | 3.2628 | 5.1153 | 5.4467 | 0.91882 | 0.22348 | 0.35035 | 0.37306 | 0.062931 | 1.5677 | 1.6693 | 0.2816 | 1.0648 | 0.17962 | 0.16869 | -1.4984 | -1.0488 | -0.98602 | -2.7657 | [10 empty] | 0 | 2.2594 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014121 | -1.8132 | -1.8273 | [14 empty] | iRM4 | Lab 3 | Jul 30, 2020 | 5 | 100.4% | -0.46873 | -1.234 | 2.0367
Lab 3 | Jul 30, 2020 | 5 | 36 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 7.0964e-05; SAMPLE 49Ti40Ar: 3.9791e-05; SAMPLE 90Zr: 14.5979; SAMPLE 91Zr: 3.2623; SAMPLE 92Zr: 5.1145; SAMPLE 94Zr: 5.4461; SAMPLE 96Zr: 0.91876 | 1 | -2.4895e-05 | 4.2409e-05 | 14.5978 | 3.2623 | 5.1144 | 0 | 5.4459 | 0 | 0.91868 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2599 | 14.5978 | 3.2623 | 5.1144 | 5.4459 | 0.91868 | 0.22348 | 0.35036 | 0.37307 | 0.062933 | 1.5677 | 1.6694 | 0.28161 | 1.0648 | 0.17963 | 0.16869 | -1.4984 | -1.0488 | -0.986 | -2.7657 | [10 empty] | 0 | 2.2599 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014123 | -1.8132 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 101.1% | -6.4061 | -4.0978 | 17.5595
Lab 3 | Jul 30, 2020 | 5 | 37 | iRM5 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: -1.5287e-05; SAMPLE 49Ti40Ar: 8.128e-06; SAMPLE 90Zr: 14.4779; SAMPLE 91Zr: 3.2355; SAMPLE 92Zr: 5.0724; SAMPLE 94Zr: 5.4013; SAMPLE 96Zr: 0.91121 | 1 | -0.00011115 | 1.0745e-05 | 14.4778 | 3.2355 | 5.0724 | 0 | 5.4012 | 0 | 0.91114 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2599 | 14.4778 | 3.2355 | 5.0724 | 5.4012 | 0.91114 | 0.22348 | 0.35035 | 0.37307 | 0.062934 | 1.5677 | 1.6694 | 0.28161 | 1.0648 | 0.17963 | 0.16869 | -1.4984 | -1.0488 | -0.986 | -2.7657 | [10 empty] | 0 | 2.2599 | 0.21796 | 0.33336 | 0.3381 | 0.05438 | 1.5295 | 1.5512 | 0.2495 | 1.0142 | 0.16313 | 0.16084 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43905 | -1.3883 | 0.014127 | -1.8132 | -1.8273 | [14 empty] | iRM5 | Lab 3 | Jul 30, 2020 | 5 | 99.91% | -1.3804 | -4.709 | 26.5393
Lab 3 | Jul 30, 2020 | 5 | 38 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 4.6267e-06; SAMPLE 49Ti40Ar: 9.4822e-06; SAMPLE 90Zr: 14.3835; SAMPLE 91Zr: 3.2145; SAMPLE 92Zr: 5.0394; SAMPLE 94Zr: 5.3662; SAMPLE 96Zr: 0.90525 | 1 | -9.1232e-05 | 1.2099e-05 | 14.3834 | 3.2144 | 5.0394 | 0 | 5.366 | 0 | 0.90518 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2603 | 14.3834 | 3.2144 | 5.0394 | 5.366 | 0.90518 | 0.22348 | 0.35036 | 0.37307 | 0.062932 | 1.5677 | 1.6694 | 0.2816 | 1.0648 | 0.17962 | 0.16869 | -1.4984 | -1.0488 | -0.98599 | -2.7657 | [10 empty] | 0 | 2.2603 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42492 | 0.43904 | -1.3883 | 0.014122 | -1.8133 | -1.8274 | [6 empty] | STD | Lab 3 | Jul 30, 2020 | 5 | 99.06% | 7.9592 | 0.089571 | -19.2518
Lab 3 | Jul 30, 2020 | 5 | 39 | iRM10 | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.6113e-05; SAMPLE 49Ti40Ar: 3.4298e-05; SAMPLE 90Zr: 14.7275; SAMPLE 91Zr: 3.2913; SAMPLE 92Zr: 5.16; SAMPLE 94Zr: 5.4946; SAMPLE 96Zr: 0.92692 | 1 | -6.9746e-05 | 3.6915e-05 | 14.7274 | 3.2913 | 5.1599 | 0 | 5.4945 | 0 | 0.92685 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2608 | 14.7274 | 3.2913 | 5.1599 | 5.4945 | 0.92685 | 0.22348 | 0.35036 | 0.37308 | 0.062934 | 1.5678 | 1.6694 | 0.28161 | 1.0648 | 0.17962 | 0.16869 | -1.4984 | -1.0488 | -0.98596 | -2.7657 | [10 empty] | 0 | 2.2608 | 0.21796 | 0.33336 | 0.3381 | 0.054377 | 1.5295 | 1.5512 | 0.24948 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3884 | 0.014119 | -1.8133 | -1.8274 | [14 empty] | iRM10 | Lab 3 | Jul 30, 2020 | 5 | 102.2% | 0.21775 | 2.5404 | -15.9936
Lab 3 | Jul 30, 2020 | 5 | 40 | STD | OPZ 177Hf++: 9.5859e-05; OPZ 49Ti40Ar: -2.6173e-06; OPZ 90Zr: 0.00010713; OPZ 91Zr: 4.5621e-05; OPZ 92Zr: 6.9842e-05; OPZ 94Zr: 0.00014589; OPZ 96Zr: 7.2235e-05; SAMPLE 177Hf++: 2.0084e-05; SAMPLE 49Ti40Ar: 7.8007e-06; SAMPLE 90Zr: 14.4427; SAMPLE 91Zr: 3.2277; SAMPLE 92Zr: 5.0602; SAMPLE 94Zr: 5.3884; SAMPLE 96Zr: 0.90902 | 1 | -7.5775e-05 | 1.0418e-05 | 14.4426 | 3.2276 | 5.0602 | 0 | 5.3883 | 0 | 0.90894 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |  | 1 | 1 | 2.2609 | 14.4426 | 3.2276 | 5.0602 | 5.3883 | 0.90894 | 0.22348 | 0.35036 | 0.37308 | 0.062935 | 1.5678 | 1.6694 | 0.28161 | 1.0648 | 0.17963 | 0.16869 | -1.4984 | -1.0488 | -0.98596 | -2.7657 | [10 empty] | 0 | 2.2609 | 0.21796 | 0.33336 | 0.3381 | 0.054378 | 1.5295 | 1.5512 | 0.24949 | 1.0142 | 0.16312 | 0.16083 | -1.5235 | -1.0985 | -2.9118 | 0.42493 | 0.43905 | -1.3883 | 0.014121 | -1.8133 | -1.8274
= -0.7 ± 2.3 (2SD)
= 0.0 ± 2.6 (2SD)
= 0.2 ± 7.2 (2SD)
Bottle #1
Bottle #2
Bottle #3
Bottle #4
Bottle #5
Bottle #6
Bottle #10
Bottle #11
Bottle #12
SRM 3169