## 1-Uncontr_Temp_Vol
Uncontrolled temperature of extraction volumes
T (°C) | ρ1 (g mL-1)
60 | 0.9832 | 0.2053% | 0.2904%
70 | 0.97776 | 0.2042% | 0.2888%
80 | 0.97179 | 0.2029% | 0.287%
90 | 0.96531 | 0.2016% | 0.2851%
100 | 0.95835 | 0.2001% | 0.283%
| Min | 0.2001% | 0.283%
| Max | 0.2053% | 0.2904%
Val. 2024/04/10

## 2-Mass_of_Analytical_Portion
Evaluation of the uncertainty of the mass of the analytical portion
Type of weighing: | weighing by difference
Analytical portion mass (mg) *: | 2.5
Repeatability standard deviation (mg): | 0.015
Balance bias observed from the calibration (mg): | 0.06
* - from sheet: "6-Recovery"
Standard uncertainty of the mass, u(m) (mg): | 0.053385
Relative standard uncertainty of the mass, u'(m): | 2.135%
Val. 2024/04/10

## 3-Calibrators_Values
Evaluation of the uncertainty of calibrators values
| Value | u | Units | u'
Pure melatonin mass | 2 | 0.053385 | mg | 0.026693
Purity of the reference substance | 0.98 | 0.011547 | - | 0.011783
V1 | 100 | 0.077969 | mL (with temperature effect) | 0.00077969
V2 | 10 | 0.13013 | µL (without temperature effect) | 0.013013
V3 | 10 | 0.02053 | mL (without temperature effect) | 0.002053
|  |  | u'_comb | 3.202%
Val. 2024/04/10
Concentration of the calibrator with lower analyte concentration:
being V1 the volume of the stock solution (100 mL), V2 the lower melatonin spiked volume (10 µL) and V3 the flask volume (10 mL).
For the uncertainty of V2 and V3, the temperature effect cancels because it's the same in both cases.
Concentration of the calibrator with lower analyte concentration:
𝒄_𝐒= (𝒎 𝑷)/𝑽_𝟏 𝑽_𝟐/𝑽_𝟑
being V1 the volume of the stock solution (100 mL), V2 the lower melatonin spiked volume (10 µL) and V3 the flask volume (10 mL).
For the uncertainty of V2 and V3, the temperature effect cancels because it's the same in both cases.

## 4-HPLC_Sensitivity_Variation
Evaluation of the uncertainty of calibrators values
Daily curve and specific sample signal*
[13 empty] | s_a | 17610.1888 |  | a | -30589.7027 |  | c_s | 90.7618
Max | 235 | [11 empty] | s_b | 319.8077 |  | b | 5103.5523
Min | 20.8 | [14 empty] | y_s | 432617.92
[6 empty] | y^ - Area predicted by the curve | [9 empty] | * - from sheet: "6-Recovery"
Calibration 1 | [5 empty] | Calibration 1 |  |  |  |  | Start | End | SUMSQ | Count | SQRT(SUMSQ/(n-2)) |  |  | Term 1 | Term 2 | Term 3 | Term 4 |  | s_c_S | s'_c_S
c (µg L-1) | Area, y | a (intercept) | b (slope) | r (correl) | r_ab | c | y^ | (y-y^)/y | ((y-y^)/y)^2 | 1 | 20.8 | 44.6 | 0.068934 | 10 | 9.283% |  |  | 61.9166 | 11.9065 | 32.3474 | -28.0368 |  | - | -
23.5 | 103141.5 | -30589.7027 | 5103.5523 | 0.99912 | -0.71431 | 23.5 | 89343.7765 | 13.38% | 0.017896 | 2 | 44.6 | 68.4 | 0.011954 | 6 | 5.467% |  |  | 21.4733 | 11.9065 | 32.3474 | -28.0368 |  | - | -
47 | 216792.5 |  |  |  |  | 47 | 209277.2558 | 3.467% | 0.0012017 | 3 | 68.4 | 92.2 | 0.0063301 | 6 | 3.978% |  |  | 11.3715 | 11.9065 | 32.3474 | -28.0368 |  | 5.2525 | 5.787%
70.5 | 337709 |  |  |  |  | 70.5 | 329210.735 | 2.516% | 0.00063325 | 4 | 92.2 | 116 | 0.0043468 | 6 | 3.297% |  |  | 7.8085 | 11.9065 | 32.3474 | -28.0368 |  | - | -
94 | 432617.92 |  |  |  |  | 94 | 449144.2142 | -3.82% | 0.0014593 | 5 | 116 | 139.8 | 0.0059184 | 6 | 3.847% |  |  | 10.6319 | 11.9065 | 32.3474 | -28.0368 |  | - | -
117.5 | 562659 |  |  |  |  | 117.5 | 569077.6934 | -1.141% | 0.00013014 | 6 | 139.8 | 163.6 | 0.0071784 | 6 | 4.236% |  |  | 12.8953 | 11.9065 | 32.3474 | -28.0368 |  | - | -
141 | 654113.54 |  |  |  |  | 141 | 689011.1726 | -5.335% | 0.0028463 | 7 | 163.6 | 187.4 | 0.0034923 | 8 | 2.413% |  |  | 4.1824 | 11.9065 | 32.3474 | -28.0368 |  | - | -
164.5 | 810180.94 |  |  |  |  | 164.5 | 808944.6518 | 0.1526% | 2.3285e-06 | 8 | 187.4 | 211.2 | 0.0031499 | 6 | 2.806% |  |  | 5.6585 | 11.9065 | 32.3474 | -28.0368 |  | - | -
188 | 936443.33 |  |  |  |  | 188 | 928878.131 | 0.8079% | 6.5265e-05 | 9 | 211.2 | 235 | 0.0021919 | 6 | 2.341% |  |  | 3.9375 | 11.9065 | 32.3474 | -28.0368 |  | - | -
211.5 | 1057024.79 |  |  |  |  | 211.5 | 1048811.6102 | 0.777% | 6.0374e-05 | [12 empty] | FINAL | 5.2525 | 5.787%
235 | 1179761.81 |  |  |  |  | 235 | 1168745.0895 | 0.9338% | 8.72e-05
Calibration 2 | [5 empty] | Calibration 2
c (µg L-1) | Area, y | a (intercept) | b (slope) | r (correl) |  | c | y^ | (y-y^)/y | ((y-y^)/y)^2 | Instrumental response linearity
23.5 | 126130.46 | 7763.758 | 4901.3714 | 0.9974 |  | 23.5 | 122945.9862 | 2.525% | 0.00063743
47 | 259603.34 |  |  |  |  | 47 | 238128.2144 | 8.272% | 0.0068431
70.5 | 339941.43 |  |  |  |  | 70.5 | 353310.4425 | -3.933% | 0.0015466
94 | 461532.75 |  |  |  |  | 94 | 468492.6707 | -1.508% | 0.00022741
117.5 | 552839.6 |  |  |  |  | 117.5 | 583674.8989 | -5.578% | 0.003111
141 | 689697.87 |  |  |  |  | 141 | 698857.1271 | -1.328% | 0.00017636
164.5 | 846405.34 |  |  |  |  | 164.5 | 814039.3553 | 3.824% | 0.0014622
188 | 963865.99 |  |  |  |  | 188 | 929221.5835 | 3.594% | 0.0012919
211.5 | 1003711.91 |  |  |  |  | 211.5 | 1044403.8116 | -4.054% | 0.0016436
235 | 1168931.44 |  |  |  |  | 235 | 1159586.0398 | 0.7995% | 6.3917e-05
Calibration 3 | [5 empty] | Calibration 3
c (µg L-1) | Area, y | a (intercept) | b (slope) | r (correl) |  | c | y^ | (y-y^)/y | ((y-y^)/y)^2
20.8 | 97066 | 483.776 | 4313.9365 | 0.99925 |  | 20.8 | 90213.6551 | 7.059% | 0.0049836
41.6 | 186967.48 |  |  |  |  | 41.6 | 179943.5342 | 3.757% | 0.0014113
62.4 | 262480.1 |  |  |  |  | 62.4 | 269673.4133 | -2.741% | 0.00075104
83.2 | 345235.52 |  |  |  |  | 83.2 | 359403.2924 | -4.104% | 0.0016841
104 | 458281.79 |  |  |  |  | 104 | 449133.1715 | 1.996% | 0.00039852
124.8 | 530827.15 |  |  |  |  | 124.8 | 538863.0505 | -1.514% | 0.00022917
145.6 | 639462.69 |  |  |  |  | 145.6 | 628592.9296 | 1.7% | 0.00028894
166.4 | 701118 |  |  |  |  | 166.4 | 718322.8087 | -2.454% | 0.00060217
187.2 | 813750.69 |  |  |  |  | 187.2 | 808052.6878 | 0.7002% | 4.903e-05
208 | 904791.69 |  |  |  |  | 208 | 897782.5669 | 0.7747% | 6.0011e-05
Calibration 4 | [5 empty] | Calibration 4
c (µg L-1) | Area, y | a (intercept) | b (slope) | r (correl) |  | c | y^ | (y-y^)/y | ((y-y^)/y)^2
20.8 | 128970.69 | 22644.196 | 4352.9549 | 0.99822 |  | 20.8 | 113185.658 | 12.24% | 0.01498
41.6 | 199743.21 |  |  |  |  | 41.6 | 203727.12 | -1.995% | 0.00039781
62.4 | 279987.38 |  |  |  |  | 62.4 | 294268.582 | -5.101% | 0.0026017
83.2 | 369796.44 |  |  |  |  | 83.2 | 384810.044 | -4.06% | 0.0016483
104 | 462131.06 |  |  |  |  | 104 | 475351.506 | -2.861% | 0.00081839
124.8 | 577361.4 |  |  |  |  | 124.8 | 565892.968 | 1.986% | 0.00039456
145.6 | 687123.39 |  |  |  |  | 145.6 | 656434.43 | 4.466% | 0.0019948
166.4 | 759131.15 |  |  |  |  | 166.4 | 746975.892 | 1.601% | 0.00025639
187.2 | 821224.49 |  |  |  |  | 187.2 | 837517.354 | -1.984% | 0.00039361
208 | 920753.16 |  |  |  |  | 208 | 928058.816 | -0.7934% | 6.2955e-05
Calibration 5 | [5 empty] | Calibration 5
c (µg L-1) | Area, y | a (intercept) | b (slope) | r (correl) |  | c | y^ | (y-y^)/y | ((y-y^)/y)^2
22.1 | 120959.22 | 4793.3927 | 4481.7603 | 0.998 |  | 22.1 | 103840.2958 | 14.15% | 0.02003
44.2 | 191380.99 |  |  |  |  | 44.2 | 202887.199 | -6.012% | 0.0036147
66.3 | 295694.15 |  |  |  |  | 66.3 | 301934.1021 | -2.11% | 0.00044532
88.4 | 394354.59 |  |  |  |  | 88.4 | 400981.0053 | -1.68% | 0.00028235
110.5 | 494173.64 |  |  |  |  | 110.5 | 500027.9084 | -1.185% | 0.00014034
132.6 | 588440.97 |  |  |  |  | 132.6 | 599074.8116 | -1.807% | 0.00032657
154.7 | 728039.1 |  |  |  |  | 154.7 | 698121.7147 | 4.109% | 0.0016886
176.8 | 816421.81 |  |  |  |  | 176.8 | 797168.6179 | 2.358% | 0.00055613
198.9 | 861181.2 |  |  |  |  | 198.9 | 896215.521 | -4.068% | 0.001655
221 | 1004867.93 |  |  |  |  | 221 | 995262.4242 | 0.9559% | 9.1374e-05
Calibration 6 | [5 empty] | Calibration 6
c (µg L-1) | Area, y | a (intercept) | b (slope) | r (correl) |  | c | y^ | (y-y^)/y | ((y-y^)/y)^2
22.1 | 95258.11 | -4326.2273 | 4764.288 | 0.99904 |  | 22.1 | 100964.5367 | -5.99% | 0.0035886
44.2 | 214258.5 |  |  |  |  | 44.2 | 206255.3008 | 3.735% | 0.0013952
66.3 | 308302.43 |  |  |  |  | 66.3 | 311546.0648 | -1.052% | 0.00011069
88.4 | 407409.6 |  |  |  |  | 88.4 | 416836.8289 | -2.314% | 0.00053543
110.5 | 503938.16 |  |  |  |  | 110.5 | 522127.593 | -3.609% | 0.0013028
132.6 | 654622.67 |  |  |  |  | 132.6 | 627418.357 | 4.156% | 0.001727
154.7 | 742766.96 |  |  |  |  | 154.7 | 732709.1211 | 1.354% | 0.00018336
176.8 | 849083.3 |  |  |  |  | 176.8 | 837999.8852 | 1.305% | 0.00017039
198.9 | 939681.28 |  |  |  |  | 198.9 | 943290.6492 | -0.3841% | 1.4754e-05
221 | 1032408.74 |  |  |  |  | 221 | 1048581.4133 | -1.566% | 0.00024539
Val. 2024/04/10

## 5-Precision
Evaluation of precision uncertainty
Precision estimated under intermediate precision conditions
LOQ (µg L-1): | 20.88 | (only instrumental quantification) | [5 empty] | LOQ (mg kg-1): | 6065.6516 | (complete process)
2*LOQ (µg L-1) (42 µg L-1): | Day 01 |  | Day 02 |  | Day 03 |  | Day 04 |  | Day 05 |  | Day 06 |  | Day 07 |  | Day 08 |  | Day 09 |  | Day 10
| Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2
| 4083 | 3946 | 4783 | 6559 | 3607 | 3832 | 3917 | 3790 | 4018 | 4147 | 4219 | 4238 | 4129 | 4251 | 4107 | 4196 | 4173 | 4129 | 4352 | 4124
sr (mg kg-1): | 96.8736 |  | 1255.8216 |  | 159.099 |  | 89.8026 |  | 91.2168 |  | 13.435 |  | 86.267 |  | 62.9325 |  | 31.1127 |  | 161.2203
d.f. | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1
sr 2*(n-1) | 9384.5 |  | 1577088 |  | 25312.5 |  | 8064.5 |  | 8320.5 |  | 180.5 |  | 7442 |  | 3960.5 |  | 968 |  | 25992
| Intermediate precision: |  |  |  |  | Repeatability:
|  | sI (mg/kg) | d.f. (m-1) | sI2*(n-1) |  |  | sr (mg/kg) | d.f.
| Estimate 1 | 301.638 | 9 | 818869.6 |  | Pooling, sr<I> | 408.254 | 10
| Estimate 2 | 803.0299 | 9 | 5803713.6 |  | Pooling, s'r<II> | 9.651%
| Pooling, sI<I> | 606.5652 | 18
| Pooling, s'I<II> | 14.34% |  |  |  | Mean | 4230 | mg/kg
10*LOQ (µg L-1) (209 µg L-1): | Day 01 |  | Day 02 |  | Day 03 |  | Day 04 |  | Day 05 |  | Day 06 |  | Day 07 |  | Day 08 |  | Day 09 |  | Day 10
| Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2
| 20134 | 19879 | 19073 | 18793 | 19020 | 19540 | 18203 | 18089 | 18891 | 20198 | 20257 | 20101 | 20662 | 19568 | 19812 | 19947 | 20403 | 19916 | 19516 | 19973
sr (mg kg-1): | 180.3122 |  | 197.9899 |  | 367.6955 |  | 80.6102 |  | 924.1886 |  | 110.3087 |  | 773.5748 |  | 95.4594 |  | 344.361 |  | 323.1478
d.f. | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1
sr 2*(n-1) | 32512.5 |  | 39200 |  | 135200 |  | 6498 |  | 854124.5 |  | 12168 |  | 598418 |  | 9112.5 |  | 118584.5 |  | 104424.5
| Intermediate precision: |  |  |  |  | Repeatability:
|  | sI (mg kg-1) | d.f. (m-1) | sI2*(n-1) |  |  | sr (mg kg-1) | d.f.
| Estimate 1 | 789.5 | 9 | 5609792.9 |  | Pooling, sr | 437.0632 | 10
| Estimate 2 | 665.4015 | 9 | 3984832.4 |  | Pooling, s'r<III> | 2.23%
| Pooling, sI | 730.0923 | 18
| Pooling, s'I<III> | 3.725% |  |  |  | Mean | 19598.75 | mg kg-1
|  | Mass fraction (mg kg-1) |  | Pooling, sr<I> | Pooling, sI<I>
|  | Initial | Final | (mg kg-1) | (mg kg-1) | Pooling, s'r<II> | Pooling, s'I<II> | Pooling, s'r<III> | Pooling, s'I<III>
| Interval I | 2115 | 4230 | 408.254 | 606.5652
| Interval II | 4230 | 19598.75 | - | - | 9.651% | 14.34%
| Interval III | 19598.75 | (…) | - | - | - | - | 2.23% | 3.725%
Val. 2024/04/10

## 6-Recovery
Evaluation of precision uncertainty
Spiking procedure:
Description |  |  |  | ---> 0.5 mL, V2, to tea portion
m (2 mg)/ 100 mL of water, V1 ---> | [7 empty] | Normal procedure where m_s is the tea sample mass (2.5 mg without uncertainty since spiking was performed after taking the analytical portion)
Purity melatonin, P - 98% |  |  |  | ---> 2.5 mL, V3, to tea portion
[7 empty] | Value | u | units | u' | Contribut.
w_sp(1): | - spiked mass fraction ((m*P/V1)*V2/m_s |  |  |  |  | m | 2 | 0.053385 | mg | 2.669% | 73.78% | w_sp(1): | 3920 | (mg kg-1)
[6 empty] | P | 0.98 | 0.011547 | - | 1.178% | 14.38%
[6 empty] | V1 | 100 | 0.077969 | mL | 0.07797% | 0.06295% | u'(w_sp(1)): | 3.108%
[6 empty] | V2 | 0.5 | 0.0053325 | mL | 1.066% | 11.78%
[6 empty] | m_s | 2.5 | 0 | mg | 0% | 0%
[10 empty] | SUM | 100%
[7 empty] | Value | u | units | u' | Contribut.
w_sp(2): | - spiked mass fraction ((m*P/V1)*V3/m_s |  |  |  |  | m | 2 | 0.053385 | mg | 2.669% | 77.05% | w_sp(2): | 19600 | (mg kg-1)
[6 empty] | P | 0.98 | 0.011547 | - | 1.178% | 15.01%
[6 empty] | V1 | 100 | 0.077969 | mL | 0.07797% | 0.06574% | u'(w_sp(2)): | 3.041%
[6 empty] | V3 | 2.5 | 0.021324 | mL | 0.8529% | 7.868%
[6 empty] | m_s | 2.5 | 0 | mg | 0% | 0%
[10 empty] | SUM | 100%
Day 01 |  | Day 02 |  | Day 03 |  | Day 04 |  | Day 05 |  | Day 06 |  | Day 07 |  | Day 08 |  | Day 09 |  | Day 10
Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2
4083 | 3946 | 4783 | 6559 | 3607 | 3832 | 3917 | 3790 | 4018 | 4147 | 4219 | 4238 | 4129 | 4251 | 4107 | 4196 | 4173 | 4129 | 4352 | 4124 | Mean | SD | n |  | Mean rec. |  | term
4014.5 |  | 5671 |  | 3719.5 |  | 3853.5 |  | 4082.5 |  | 4228.5 |  | 4190 |  | 4151.5 |  | 4151 |  | 4238 |  | 4230 | 533.4489 | 10 |  | 107.9% |  | 0.0029763
Day 01 |  | Day 02 |  | Day 03 |  | Day 04 |  | Day 05 |  | Day 06 |  | Day 07 |  | Day 08 |  | Day 09 |  | Day 10
Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2 | Rep.1 | Rep.2
20134 | 19879 | 19073 | 18793 | 19020 | 19540 | 18203 | 18089 | 18891 | 20198 | 20257 | 20101 | 20662 | 19568 | 19812 | 19947 | 20403 | 19916 | 19516 | 19973 | Mean | SD | n |  | Mean rec. |  | term
20006.5 |  | 18933 |  | 19280 |  | 18146 |  | 19544.5 |  | 20179 |  | 20115 |  | 19879.5 |  | 20159.5 |  | 19744.5 |  | 19598.75 | 653.3859 | 10 |  | 99.99% |  | 0.0010357
[23 empty] | overall mean: | 104%
[23 empty] | uRec | 3.167%
[23 empty] | t-test | 1.2475 | < 2
[23 empty] | u'Rec | 0.03167
Val. 2024/04/10

## 7-Comb&Expansion
Combination and expansion of the uncertainty
Daily data:
m (mg) | 2.5 | (tea sample mass) | Uncertainty components
|  |  | Precision uncertainty | [5 empty] | EXPANDED UNCERTAINTY
a (intercept) | -30589.7027 |  | Mass fraction (g kg-1)
b (slope) | 5103.5523 |  | Initial | Final
y_s (mean) | 432617.92 |  | 2.115 | 4.23 | sr<I> (g kg-1) | 0.40825 | sI<I> (g kg-1) | 0.60657
|  |  | 4.23 | 19.5987 | s'r<II> | 9.651% | s'I<II> | 14.34%
c_s (µg L-1) | 90.7618 |  | 19.5987 | 195.9875 | s'r<III> | 2.23% | s'I<III> | 3.725%
Within day replicates | 1 |  |  | s'r | s'I | Comb.
|  |  | s'I (final) | 9.651% | 14.34% | 14.34%
w (g kg-1) | 9.0762
u'w | 16.25% |  | u'm | 2.135%
U'w (k = 2, c.l. 95%) | 32.49% |  | u'c | 3.202%
Uw (k = 2, c.l. 95%) | 2.9492 |  | u'c_s | 5.787%
|  |  | u'R | 3.167%
Val. 2024/04/10
Being k = 2 for a level of confidence of 95% or k = 3 for 99% l.o.c.
𝑼=𝒌 ∗ 𝒖
Being k = 2 for a level of confidence of 95% or k = 3 for 99% l.o.c.