## Overview
Flow IDMS Calculation Overview
Date of Calculation | 2022-09-02
Date of Analysis | 2022-09-01
Analyst(s) | SP
Generator | PSA 10.534 (high-level Hg0)
Detector | ICP-MS06
CV | Cetac HGX-200
Peripump | ICP-MS06 peripump
Mass Flow Meter | Drycal Pro Flexcal (S/N 175689)
Set-up | Panayot's Flow IDMS set-up
Gen Flow Settings | 1.5 L/min dilution flow + 1 mL/min reservoir flow
Temp (°C) | 35.2
Press(hPa) | 1027.1
Dumarey Conc (ng/L) measured flow rate | 19.96
Dumarey Conc (ng/L) flow setting | 30.38
Measured Conc (ng/L) | 34.0337
Measured/Calc Ratio | 112%

## Calculation template
| Type (measured/calc) | Term | Anotation | Unit | Value | Uncertainty
[7 empty] | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
ü | calculated | Spike Mass Flow 199Hg | MFt_e | [pg min-1] | 1053.9121 | 30.6862 | 1084.5983 | 1053.9121 | 1053.9121 | 1053.9121 | 1053.9121 | 1053.9121 | 1053.9121 | 1053.9121 | 1053.9121
ü | calculated | Volumetric flow (split-flow) | Vf | [sl min-1] | 0.22678 | 0.0058292 | 0.22678 | 0.23261 | 0.22678 | 0.22678 | 0.22678 | 0.22678 | 0.22678 | 0.22678 | 0.22678
ü | cert (IUPAC website) | natHg Atomic weight | Aw(Hg_n) | g mol-1 | 200.592 | 0.0015 | 200.592 | 200.592 | 200.5935 | 200.592 | 200.592 | 200.592 | 200.592 | 200.592 | 200.592
ü | calculated | 199Hg Atomic weight | Aw(Hg_e) | g mol-1 | 199.081 | 0.13352 | 199.081 | 199.081 | 199.081 | 199.2145 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081
ü | cert | enrHg Atom fraction199Hg | A(199Hg)_e | % | 91.95 | 0.05 | 91.95 | 91.95 | 91.95 | 91.95 | 92 | 91.95 | 91.95 | 91.95 | 91.95
ü | cert (IUPAC website) | natHg Atom fraction202Hg | A(202Hg)_n | % | 29.74 | 0.13 | 29.74 | 29.74 | 29.74 | 29.74 | 29.74 | 29.87 | 29.74 | 29.74 | 29.74
ext file checked ü | measured | Blend IR (202/199) | Rb(202Hg/199Hg) | mol mol-1 | 1.0084 | 0.0071001 | 1.0084 | 1.0084 | 1.0084 | 1.0084 | 1.0084 | 1.0084 | 1.0155 | 1.0084 | 1.0084
ü | calculated | 199Hg IR (202/199) | Re(202Hg/199Hg) | mol mol-1 | 0.0079391 | 0.00010884 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0080479 | 0.0079391
ü | calculated | natHg IR (199/202) | Rn(199Hg/202Hg) | mol mol-1 | 0.5696 | 0.0047413 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.57434
| measured by JSI (Dec 2022) | Cold Vapour Efficiency | CVE |  | 0.99652 | 0.00023108 | 0.99652 | 0.99652 | 0.99652 | 0.99652 | 0.99652 | 0.99652 | 0.99652 | 0.99652 | 0.99652 | 0.99675
ü | calculated | Generator mass flow (natHg) | MFgen | ng m-3 (=pg/L) | 34033.6514 | 1485.6114 | 35024.5918 | 33180.7503 | 34033.9059 | 34010.841 | 34052.158 | 33885.5304 | 34603.9979 | 34029.949 | 34420.3379
[5 empty] | RSU | 4.365%
[7 empty] | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
[6 empty] | Squared deviations | 981962.7346 | 727440.4171 | 0.06477 | 520.3158 | 342.4945 | 21939.8348 | 325295.0705 | 13.7081 | 149526.4593
[6 empty] | Uncertainty Budget | 44.49% | 32.96% | 2.935e-06% | 0.02358% | 0.01552% | 0.9941% | 14.74% | 0.0006211% | 6.775%
|  | Term | Anotation | Unit | Value | Uncertainty
ü | calculated (external) | enrHg mass fraction | W(199Hg) | pg g-1 | 4182.1911 | 41.793
ü | cert | enrHg Atom fraction 196Hg | A(196Hg)_e | % | 0.02 | 0.02
ü | cert | enrHg Atom fraction 198Hg | A(198Hg)_e | % | 1.63 | 0.02
ü | cert | enrHg Atom fraction 200Hg | A(200Hg)_e | % | 4.92 | 0.03
ü | cert | enrHg Atom fraction 201Hg | A(201Hg)_e | % | 0.66 | 0.01
ü | cert | enrHg Atom fraction 202Hg | A(202Hg)_e | % | 0.73 | 0.01
ü | cert | enrHg Atom fraction 204Hg | A(204Hg)_e | % | 0.11 | 0.01
ü | cert (IUPAC website) | natHg Atom fraction199Hg | A(199Hg)_n | % | 16.94 | 0.12
ü | measured | Total Gen flow | F(Gen_tot) | [sl min-1] | 2.2653 | 0.0045153
ü | measured | Waste Gen Flow | F(Gen_exh) | [sl min-1] | 2.0385 | 0.0036868
ü | measured | Peri Pump flow | F(PP) | [g min-1] | 0.25288 | 0.0069155
|  |  |  | ü | (difference)2 |  | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
|  |  |  | ü |  | 199Hg Atomic weight | 0 | 0 | 0 | 0.017828 | 0 | 0 | 0 | 0 | 0
|  |  |  | ü |  | natHg IR (199/202) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.248e-05
|  |  |  | ü |  | 199Hg IR (202/199) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.1846e-08 | 0
|  |  |  | ü |  | Spike Mass Flow 199Hg | 941.6438 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | ü |  | Valumetric flow (split-flow) | 0 | 3.398e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | ü |  | Generator mass flow (natHg) | 981962.7346 | 727440.4171 | 0.06477 | 520.3158 | 342.4945 | 21939.8348 | 325295.0705 | 13.7081 | 149526.4593
|  |  |  | ü | budget |  | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
|  |  |  | ü |  | 199Hg Atomic weight | 0% | 0% | 0% | 100% | 0% | 0% | 0% | 0% | 0%
|  |  |  | ü |  | natHg IR (199/202) | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 100%
|  |  |  | ü |  | 199Hg IR (202/199) | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0%
|  |  |  | ü |  | Spike Mass Flow 199Hg | 100% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
|  |  |  | ü |  | Valumetric flow (split-flow) | 0% | 100% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
|  |  |  | ü |  | Generator mass flow (natHg) | 44.49% | 32.96% | 2.935e-06% | 0.02358% | 0.01552% | 0.9941% | 14.74% | 0.0006211% | 6.775%
Auto updating table from the website
| Atomic mass (Da) | Isotopic abundance (amount fraction)
ü | 195.965 83(2) | 0.0015(1)
ü | 197.966 769(3) | 0.1004(3)
ü | 198.968 281(4) | 0.1694(12)
ü | 199.968 327(4) | 0.2314(9)
ü | 200.970 303(5) | 0.1317(9)
ü | 201.970 644(5) | 0.2974(13)
ü | 203.973 494(3) | 0.0682(4)
ü
ü
ü
ü
ü
ü
ü
ü
ü
〖𝑀𝐹〗_𝐺𝐸𝑁=〖𝑀𝐹〗_(𝑡_𝑒)/𝑉_𝑓 ⋅〖𝐴𝑤〗_((𝐻〖𝑔_〗_𝑛))/〖𝐴𝑤〗_((𝐻𝑔_(_𝑒))) ⋅𝐴_((199𝐻𝑔)_𝑒) )/𝐴_((202𝐻𝑔)_𝑛) ⋅((𝑅_(𝑏(202𝐻𝑔/199𝐻𝑔))−𝑅_(𝑒(202𝐻𝑔/199𝐻𝑔)))/(1−𝑅_(𝑏(202𝐻𝑔/199𝐻𝑔))⋅𝑅_(𝑛(199𝐻𝑔/202𝐻𝑔)) ))
Note on G3: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Adding the cold vapour efficiency to the calculation has a negligible effect on the overall MFgen uncertainty, and a minor effect on the spike mass flow uncertainty (increases by 0.01 pg/min). The contribution of the spike mass flow in the uncertainty budget is still 26% and does not change.
Note on G5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: https://www.ciaaw.org/mercury.htm - assumed k=2
Note on C12: Sophie Page: negligible effect on uncertainty and very minor effect on generator mass flow result
Note on G29: Sophie Page: based on uncertainty from certificate. The SD of flow measurement is of similar magnitude
Note on G30: Sophie Page: based on uncertainty from certificate. The SD of flow measurement is of similar magnitude

## Peripump MassFlow
Peripump Massflow Determination
Date | 2022-09-02
Peripump | ICP-MS06 peripump
Analysts | SP/PP
Solution | 10% v/v HCl
Time [s] | Time [min] | weight loss [g] | Momentum mass flow [g/min] | Mass flow [g/min] |  | Mass flow at sample scale
0 | 0 | -0.037 |  |  | AVG | 0.25288
10 | 0.16667 | -0.0796 | 0.2556 |  | SD | 0.0069155
20 | 0.33333 | -0.124 | 0.2664 |  | RSD | 2.735%
30 | 0.5 | -0.15 | 0.156
40 | 0.66667 | -0.2066 | 0.3396
50 | 0.83333 | -0.2532 | 0.2796
60 | 1 | -0.2966 | 0.2604 | 0.2596
70 | 1.1667 | -0.3411 | 0.267
80 | 1.3333 | -0.371 | 0.1794
90 | 1.5 | -0.4129 | 0.2514
100 | 1.6667 | -0.4548 | 0.2514
110 | 1.8333 | -0.5024 | 0.2856
120 | 2 | -0.5483 | 0.2754 | 0.2517
130 | 2.1667 | -0.5856 | 0.2238
140 | 2.3333 | -0.6226 | 0.222
150 | 2.5 | -0.6625 | 0.2394
160 | 2.6667 | -0.7132 | 0.3042
170 | 2.8333 | -0.751 | 0.2268
180 | 3 | -0.8017 | 0.3042 | 0.2534
190 | 3.1667 | -0.8458 | 0.2646
200 | 3.3333 | -0.8852 | 0.2364
210 | 3.5 | -0.9427 | 0.345
220 | 3.6667 | -0.9835 | 0.2448
230 | 3.8333 | -1.0181 | 0.2076
240 | 4 | -1.0571 | 0.234 | 0.2554
250 | 4.1667 | -1.0881 | 0.186
260 | 4.3333 | -1.1393 | 0.3072
270 | 4.5 | -1.1792 | 0.2394
280 | 4.6667 | -1.221 | 0.2508
290 | 4.8333 | -1.2647 | 0.2622
300 | 5 | -1.3027 | 0.228 | 0.2456
310 | 5.1667 | -1.3693 | 0.3996
320 | 5.3333 | -1.4004 | 0.1866
330 | 5.5 | -1.4262 | 0.1548
340 | 5.6667 | -1.4943 | 0.4086
350 | 5.8333 | -1.5216 | 0.1638
360 | 6 | -1.558 | 0.2184 | 0.2553
370 | 6.1667 | -1.6484 | 0.5424
380 | 6.3333 | -1.6556 | 0.0432
390 | 6.5 | -1.7177 | 0.3726
400 | 6.6667 | -1.7332 | 0.093
410 | 6.8333 | -1.7624 | 0.1752
420 | 7 | -1.8002 | 0.2268 | 0.2422
430 | 7.1667 | -1.8445 | 0.2658
440 | 7.3333 | -1.8888 | 0.2658
450 | 7.5 | -1.9352 | 0.2784
460 | 7.6667 | -1.9795 | 0.2658
470 | 7.8333 | -2.0166 | 0.2226
480 | 8 | -2.0651 | 0.291 | 0.2649
490 | 8.1667 | -2.0987 | 0.2016
500 | 8.3333 | -2.1472 | 0.291
510 | 8.5 | -2.1886 | 0.2484
520 | 8.6667 | -2.2305 | 0.2514
530 | 8.8333 | -2.2689 | 0.2304
540 | 9 | -2.3106 | 0.2502 | 0.2455
550 | 9.1667 | -2.3575 | 0.2814
560 | 9.3333 | -2.4034 | 0.2754
570 | 9.5 | -2.4417 | 0.2298
580 | 9.6667 | -2.4858 | 0.2646
590 | 9.8333 | -2.5184 | 0.1956
600 | 10 | -2.5658 | 0.2844 | 0.2552

## FlexCal-H value and u
Reference flow | Bias | Standard uncertainty
502.38 | -0.39% | 0.18%
5001.2 | -0.12% | 0.18%
49911 | 0.4% | 0.19%
Measured flow [ml min-1] | Corrected flow | Calibration Unceratinty (k=1) | Repeatability | Calibration Unceratinty (k=1) | Repeatability | Combined U
2.2515 | 2.2653 | 0.1794% | 0.08678% | 0.0040649 | 0.0019658 | 0.0045153
2.0265 | 2.0385 | 0.1794% | 0.02259% | 0.0036579 | 0.0004605 | 0.0036868
0.1993%
0.1809%

## Aw(Hg_e)
Type (measured/calc) | Term | Anotation | Unit | Value | Uncertainty
[6 empty] | Atomic mass 196Hg | Atomic mass 198Hg | Atomic mass 199Hg | Atomic mass 200Hg | Atomic mass 201Hg | Atomic mass 202Hg | Atomic mass 204Hg | enrHg Atom fraction 196Hg | enrHg Atom fraction 198Hg | enrHg Atom fraction199Hg | enrHg Atom fraction 200Hg | enrHg Atom fraction 201Hg | enrHg Atom fraction 202Hg | enrHg Atom fraction 204Hg
cert (IUPAC website) | Atomic mass 196Hg | Aw(196Hg) | Da | 195.9658 | 2e-05 | 195.9659 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658
cert (IUPAC website) | Atomic mass 198Hg | Aw(198Hg) | Da | 197.9668 | 3e-06 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668
cert (IUPAC website) | Atomic mass 199Hg | Aw(199Hg) | Da | 198.9683 | 4e-06 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683
cert (IUPAC website) | Atomic mass 200Hg | Aw(200Hg) | Da | 199.9683 | 4e-06 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683
cert (IUPAC website) | Atomic mass 201Hg | Aw(201Hg) | Da | 200.9703 | 5e-06 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703
cert (IUPAC website) | Atomic mass 202Hg | Aw(202Hg) | Da | 201.9706 | 5e-06 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706
cert (IUPAC website) | Atomic mass 204Hg | Aw(204Hg) | Da | 203.9735 | 3e-06 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735
cert | enrHg Atom fraction 196Hg | A(196Hg)_e | % | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
cert | enrHg Atom fraction 198Hg | A(198Hg)_e | % | 1.63 | 0.02 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.65 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63
cert | enrHg Atom fraction199Hg | A(199Hg)_e | % | 91.95 | 0.05 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 92 | 91.95 | 91.95 | 91.95 | 91.95
cert | enrHg Atom fraction 200Hg | A(200Hg)_e | % | 4.92 | 0.03 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.95 | 4.92 | 4.92 | 4.92
cert | enrHg Atom fraction 201Hg | A(201Hg)_e | % | 0.66 | 0.01 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.67 | 0.66 | 0.66
cert | enrHg Atom fraction 202Hg | A(202Hg)_e | % | 0.73 | 0.01 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.74 | 0.73
cert | enrHg Atom fraction 204Hg | A(204Hg)_e | % | 0.11 | 0.01 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12
calculated | 199Hg Atomic weight | Aw(Hg_e) | g mol-1 | 199.081 | 0.13352 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081 | 199.1202 | 199.1206 | 199.1805 | 199.141 | 199.1011 | 199.1012 | 199.1014
[6 empty] | Atomic mass 196Hg | Atomic mass 198Hg | Atomic mass 199Hg | Atomic mass 200Hg | Atomic mass 201Hg | Atomic mass 202Hg | Atomic mass 204Hg | enrHg Atom fraction 196Hg | enrHg Atom fraction 198Hg | enrHg Atom fraction199Hg | enrHg Atom fraction 200Hg | enrHg Atom fraction 201Hg | enrHg Atom fraction 202Hg | enrHg Atom fraction 204Hg
[5 empty] | Squared deviations | 1.6e-17 | 2.3912e-15 | 1.3528e-11 | 3.873e-14 | 1.089e-15 | 1.3322e-15 | 1.089e-17 | 0.0015361 | 0.0015676 | 0.0098971 | 0.0035989 | 0.00040389 | 0.00040792 | 0.00041605
[5 empty] | Budget | 8.975e-14% | 1.341e-11% | 7.588e-08% | 2.172e-10% | 6.109e-12% | 7.473e-12% | 6.109e-14% | 8.616% | 8.793% | 55.52% | 20.19% | 2.266% | 2.288% | 2.334%
Auto updating table from the website
Atomic mass (Da) | Isotopic abundance (amount fraction)
195.965 83(2) | 0.0015(1)
197.966 769(3) | 0.1004(3)
198.968 281(4) | 0.1694(12)
199.968 327(4) | 0.2314(9)
200.970 303(5) | 0.1317(9)
201.970 644(5) | 0.2974(13)
203.973 494(3) | 0.0682(4)

## Mass fraction 199Hg
Dilution with 3 weighings (bottle-analyte-solvent)
Analyte/Std ID | 199Hg Spike (ORNL) SAMD/116 in 20% HNO3 prepared 02/03/2005 |  | Analyst/Lab book |  | SP1 88-89
Solvent | 10 % v/v c.HCl Romil UpA |  | Date |  | 2022-08-31
Source/Lot no | TBC |  | Balance |  | Denver Si-234A, S/N 23408682, Weights NPL 1405
199Hg Dilution 1
Bottle | 12.658 | Data: | value | repeat | linearity | std. uncert. |  | n | derivatives | variances
| 12.6584 | m1 | 12.6582 | 0.00012019 | 9.2376e-06 | 0.00012054 | g | 3 | -14.6283 | 3.1092e-06 | 0.33259 | %
| 12.6583 | m2 | 12.8665 | 3.3333e-05 | 9.2376e-06 | 3.459e-05 | g | 3 | 14.6892 | 2.5816e-07 | 0.027615 | %
|  | m3 | 62.9111 | 3.3333e-05 | 0.00012124 | 0.00012574 | g | 3 | -0.060887 | 5.8616e-11 | 6.2701e-06 | %
|  | wstock | 738.1719 |  |  | 7.3631 | µg/g |  | 0.004145 | 0.00093148 | 99.6398 | %
Bottle +analyte | 12.8665 | [7 empty] | sum | 0.00093485
| 12.8665 | Results | mass (g) | uncert. | unit | rsu (%)
| 12.8666 | wAnalyte | 3.0598 | 0.030575 | µg/g | 0.99927 | %
|  | wAnalyte | 3059.752 | 30.5753 | ng/g | 0.99927 | %
Bottle +analyte+solvent | 62.9111 | m(Analyte) | 0.2083 | 0.0001254 | g | 0.060204 | %
| 62.9111 | m (Solution) | 50.2528 | 0.00017419 | g | 0.00034662 | %
62.911
199Hg Dilution 2
Bottle | 12.6758 | Data: | value | repeat | linearity | std. uncert. |  | n | derivatives | variances
| 12.6756 | m1 | 12.6757 | 5.7735e-05 | 9.2376e-06 | 5.8469e-05 | g | 3 | -59.7922 | 1.2222e-05 | 0.0032629 | %
| 12.6757 | m2 | 13.6795 | 5.7735e-05 | 9.2376e-06 | 5.8469e-05 | g | 3 | 61.0134 | 1.2726e-05 | 0.0033975 | %
|  | m3 | 62.8245 | 3.3333e-05 | 0.00012124 | 0.00012574 | g | 3 | -1.2213 | 2.3582e-08 | 6.2957e-06 | %
|  | wstock | 3059.752 |  |  | 30.5753 | ng/g |  | 0.020016 | 0.37455 | 99.9933 | %
Bottle +analyte | 13.6794 | [7 empty] | sum | 0.37458
| 13.6795 | Results | mass (g) | uncert. | unit | rsu (%)
| 13.6796 | wAnalyte | 61.2453 | 0.61203 | ng/g | 0.99931 | %
|  | wAnalyte | 61245.2747 | 612.0281 | pg/g | 0.99931 | %
Bottle +analyte+solvent | 62.8245 | m(Analyte) | 1.0038 | 8.2688e-05 | g | 0.0082375 | %
| 62.8245 | m (Solution) | 50.1488 | 0.00013867 | g | 0.00027652 | %
62.8246
4 ppb 199Hg
Bottle | 12.8958 | Data: | value | repeat | linearity | std. uncert. |  | n | derivatives | variances
| 12.8958 | m1 | 12.8958 | 3.3333e-05 | 9.2376e-06 | 3.459e-05 | g | 3 | -1.1483 | 1.5776e-09 | 9.0319e-05 | %
| 12.8957 | m2 | 16.2892 | 5.7735e-05 | 9.2376e-06 | 5.8469e-05 | g | 3 | 1.2324 | 5.1926e-09 | 0.00029729 | %
|  | m3 | 62.5902 | 3.3333e-05 | 0.00012124 | 0.00012574 | g | 3 | -0.084158 | 1.1198e-10 | 6.4113e-06 | %
|  | wstock | 61.2453 |  |  | 0.61203 | ng/g |  | 0.068286 | 0.0017466 | 99.9996 | %
Bottle +analyte | 16.2892 | [7 empty] | sum | 0.0017467
| 16.2893 | Results | mass (g) | uncert. | unit | rsu (%)
| 16.2891 | wAnalyte | 4.1822 | 0.041793 | ng/g | 0.99931 | %
|  | wAnalyte | 4182.1911 | 41.793 | pg/g | 0.99931 | %
Bottle +analyte+solvent | 62.5902 | m(Analyte) | 3.3934 | 6.7935e-05 | g | 0.0020019 | %
| 62.5903 | m (Solution) | 49.6945 | 0.00013041 | g | 0.00026243 | %
62.5902
Note on H12: Sophie Page: Concentration and uncertainty from reverse IDMS of 199Hg from 10/02/2022 SP/PP (T:\QT20\Inorganic Speciation Analysis Team\Projects\Current\CBM2125\WP2\220210 Reverse IDMS 199Hg)