## Overview
Flow IDMS Calculation Overview
Date of Calculation | 2023-05-04
Date of Analysis | 2023-04-26
Analyst(s) | SP
Generator | PSA 10.532 fixed flow HgCl2
Detector | ICP-MS06
CV | Cetac HGX-200
Peripump | ICP-MS06 peripump
Mass Flow Meter | Drycal Pro Flexcal (S/N 175689)
Set-up | Panayot's Flow IDMS set-up
Gen Flow Settings | 14.8 L/min dil + 9 mL/min res
Temp (°C) | 40.1
Press(hPa) | 1173

## Calculation template
| Type (measured/calc) | Term | Anotation | Unit | Value | Uncertainty
[7 empty] | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
ü | calculated | Spike Mass Flow 199Hg | MFt_e | [pg min-1] | 140.0943 | 3.3002 | 143.3944 | 140.0943 | 140.0943 | 140.0943 | 140.0943 | 140.0943 | 140.0943 | 140.0943 | 140.0943
ü | calculated | Volumetric flow (split-flow) | Vf | [sl min-1] | 0.204 | 0.00408 | 0.204 | 0.20808 | 0.204 | 0.204 | 0.204 | 0.204 | 0.204 | 0.204 | 0.204
ü | cert (IUPAC website) | natHg Atomic weight | Aw(Hg_n) | g mol-1 | 200.592 | 0.0015 | 200.592 | 200.592 | 200.5935 | 200.592 | 200.592 | 200.592 | 200.592 | 200.592 | 200.592
ü | calculated | 199Hg Atomic weight | Aw(Hg_e) | g mol-1 | 199.081 | 0.13352 | 199.081 | 199.081 | 199.081 | 199.2145 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081
ü | cert | enrHg Atom fraction199Hg | A(199Hg)_e | % | 91.95 | 0.05 | 91.95 | 91.95 | 91.95 | 91.95 | 92 | 91.95 | 91.95 | 91.95 | 91.95
ü | cert (IUPAC website) | natHg Atom fraction202Hg | A(202Hg)_n | % | 29.74 | 0.13 | 29.74 | 29.74 | 29.74 | 29.74 | 29.74 | 29.87 | 29.74 | 29.74 | 29.74
ext file checked ü | measured | Blend IR (202/199) | Rb(202Hg/199Hg) | mol mol-1 | 0.94128 | 0.0021174 | 0.94128 | 0.94128 | 0.94128 | 0.94128 | 0.94128 | 0.94128 | 0.9434 | 0.94128 | 0.94128
ü | calculated | 199Hg IR (202/199) | Re(202Hg/199Hg) | mol mol-1 | 0.0079391 | 0.00010884 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0079391 | 0.0080479 | 0.0079391
ü | calculated | natHg IR (199/202) | Rn(199Hg/202Hg) | mol mol-1 | 0.5696 | 0.0047413 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.5696 | 0.57434
ü | calculated | Generator mass flow (natHg) | MFgen | ng m-3 (=pg/L) | 4304.8011 | 141.293 | 4406.2079 | 4220.3933 | 4304.8333 | 4301.9159 | 4307.142 | 4286.0658 | 4325.8148 | 4304.2991 | 4346.6227
[5 empty] | RSU | 3.282%
[7 empty] | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
[6 empty] | Squared deviations | 10283.3413 | 7124.6877 | 0.0010362 | 8.3245 | 5.4795 | 351.0124 | 441.5755 | 0.252 | 1749.0463
[6 empty] | Uncertainty Budget | 51.51% | 35.69% | 5.191e-06% | 0.0417% | 0.02745% | 1.758% | 2.212% | 0.001262% | 8.761%
|  | Term | Anotation | Unit | Value | Uncertainty
ü | calculated (external) | enrHg mass fraction | W(199Hg) | pg g-1 | 514.1828 | 6.195
ü | cert | enrHg Atom fraction 196Hg | A(196Hg)_e | % | 0.02 | 0.02
ü | cert | enrHg Atom fraction 198Hg | A(198Hg)_e | % | 1.63 | 0.02
ü | cert | enrHg Atom fraction 200Hg | A(200Hg)_e | % | 4.92 | 0.03
ü | cert | enrHg Atom fraction 201Hg | A(201Hg)_e | % | 0.66 | 0.01
ü | cert | enrHg Atom fraction 202Hg | A(202Hg)_e | % | 0.73 | 0.01
ü | cert | enrHg Atom fraction 204Hg | A(204Hg)_e | % | 0.11 | 0.01
ü | cert (IUPAC website) | natHg Atom fraction199Hg | A(199Hg)_n | % | 16.94 | 0.12
ü | measured | Total Gen flow | F(Gen_tot) | [sl min-1] | 13.4497 | 0.024249
ü | measured | Waste Gen Flow | F(Gen_exh) | [sl min-1] | 13.1831 | 0.026712
ü | measured | Peri Pump flow | F(PP) | [g min-1] | 0.27246 | 0.0055153
|  |  |  | ü | (difference)2 |  | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
|  |  |  | ü |  | 199Hg Atomic weight | 0 | 0 | 0 | 0.017828 | 0 | 0 | 0 | 0 | 0
|  |  |  | ü |  | natHg IR (199/202) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.248e-05
|  |  |  | ü |  | 199Hg IR (202/199) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.1846e-08 | 0
|  |  |  | ü |  | Spike Mass Flow 199Hg | 10.891 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | ü |  | Valumetric flow (split-flow) | 0 | 1.6646e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | ü |  | Generator mass flow (natHg) | 10283.3413 | 7124.6877 | 0.0010362 | 8.3245 | 5.4795 | 351.0124 | 441.5755 | 0.252 | 1749.0463
|  |  |  | ü | budget |  | Spike Mass Flow 199Hg | Volumetric flow (split-flow) | natHg Atomic weight | 199Hg Atomic weight | enrHg Atom fraction199Hg | natHg Atom fraction202Hg | Blend IR (202/199) | 199Hg IR (202/199) | natHg IR (199/202)
|  |  |  | ü |  | 199Hg Atomic weight | 0% | 0% | 0% | 100% | 0% | 0% | 0% | 0% | 0%
|  |  |  | ü |  | natHg IR (199/202) | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 100%
|  |  |  | ü |  | 199Hg IR (202/199) | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0%
|  |  |  | ü |  | Spike Mass Flow 199Hg | 100% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
|  |  |  | ü |  | Valumetric flow (split-flow) | 0% | 100% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
|  |  |  | ü |  | Generator mass flow (natHg) | 51.51% | 35.69% | 5.191e-06% | 0.0417% | 0.02745% | 1.758% | 2.212% | 0.001262% | 8.761%
Auto updating table from the website
| Atomic mass (Da) | Isotopic abundance (amount fraction)
ü | 195.965 83(2) | 0.0015(1)
ü | 197.966 769(3) | 0.1004(3)
ü | 198.968 281(4) | 0.1694(12)
ü | 199.968 327(4) | 0.2314(9)
ü | 200.970 303(5) | 0.1317(9)
ü | 201.970 644(5) | 0.2974(13)
ü | 203.973 494(3) | 0.0682(4)
ü
ü
ü
ü
ü
ü
ü
ü
ü
〖𝑀𝐹〗_𝐺𝐸𝑁=〖𝑀𝐹〗_(𝑡_𝑒)/𝑉_𝑓 ⋅〖𝐴𝑤〗_((𝐻〖𝑔_〗_𝑛))/〖𝐴𝑤〗_((𝐻𝑔_(_𝑒))) ⋅𝐴_((199𝐻𝑔)_𝑒) )/𝐴_((202𝐻𝑔)_𝑛) ⋅((𝑅_(𝑏(202𝐻𝑔/199𝐻𝑔))−𝑅_(𝑒(202𝐻𝑔/199𝐻𝑔)))/(1−𝑅_(𝑏(202𝐻𝑔/199𝐻𝑔))⋅𝑅_(𝑛(199𝐻𝑔/202𝐻𝑔)) ))
Note on G4: Sophie Page: based on uncertianty from certificate since the ADM flow meter only shows the real time flow and there is no way to record the standard devaition over time.
Note on G5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: https://www.ciaaw.org/mercury.htm - assumed k=2
Note on F28: Sophie Page: FelxCal-H flow meter is broken. Instead split flow was measured by connecting the 1/8 inch tubing after the split, but before the merge with the cold vapour output, to the ADM flow meter. However this had to be done before turning on the ICP, so may not be as accurate.
Note on G28: Sophie Page: based on uncertainty from certificate. The SD of flow measurement is of similar magnitude
Note on F29: Sophie Page: FelxCal-H flow meter is broken. Instead split flow was measured by connecting the 1/8 inch tubing after the split, but before the merge with the cold vapour output, to the ADM flow meter. However this had to be done before turning on the ICP, so may not be as accurate.
Note on G29: Sophie Page: based on uncertainty from certificate. The SD of flow measurement is of similar magnitude

## Peripump MassFlow
Peripump Massflow Determination
Date | 2023-04-26
Peripump | ICP-MS06 peripump
Analysts | SP
Solution | 199Hg in 10% v/v HCl
Time [s] | Time [min] | weight loss [g] | Momentum mass flow [g/min] | Mass flow [g/min] |  | Mass flow at sample scale
0 | 0 | -0.2637 |  |  | AVG | 0.27246
10 | 0.16667 | -0.3081 | 0.2664 |  | SD | 0.0055153
20 | 0.33333 | -0.3518 | 0.2622 |  | RSD | 2.024%
30 | 0.5 | -0.3956 | 0.2628
40 | 0.66667 | -0.44 | 0.2664
50 | 0.83333 | -0.4836 | 0.2616
60 | 1 | -0.5275 | 0.2634 | 0.2638
70 | 1.1667 | -0.5733 | 0.2748
80 | 1.3333 | -0.6161 | 0.2568
90 | 1.5 | -0.6601 | 0.264
100 | 1.6667 | -0.7048 | 0.2682
110 | 1.8333 | -0.7489 | 0.2646
120 | 2 | -0.7944 | 0.273 | 0.2669
130 | 2.1667 | -0.8377 | 0.2598
140 | 2.3333 | -0.8828 | 0.2706
150 | 2.5 | -0.9292 | 0.2784
160 | 2.6667 | -0.974 | 0.2688
170 | 2.8333 | -1.0185 | 0.267
180 | 3 | -1.0609 | 0.2544 | 0.2665
190 | 3.1667 | -1.1057 | 0.2688
200 | 3.3333 | -1.1511 | 0.2724
210 | 3.5 | -1.1983 | 0.2832
220 | 3.6667 | -1.2415 | 0.2592
230 | 3.8333 | -1.2861 | 0.2676
240 | 4 | -1.3298 | 0.2622 | 0.2689
250 | 4.1667 | -1.3769 | 0.2826
260 | 4.3333 | -1.4235 | 0.2796
270 | 4.5 | -1.47 | 0.279
280 | 4.6667 | -1.5126 | 0.2556
290 | 4.8333 | -1.5576 | 0.27
300 | 5 | -1.6045 | 0.2814 | 0.2747
310 | 5.1667 | -1.6504 | 0.2754
320 | 5.3333 | -1.6948 | 0.2664
330 | 5.5 | -1.7418 | 0.282
340 | 5.6667 | -1.7864 | 0.2676
350 | 5.8333 | -1.8336 | 0.2832
360 | 6 | -1.8781 | 0.267 | 0.2736
370 | 6.1667 | -1.926 | 0.2874
380 | 6.3333 | -1.9692 | 0.2592
390 | 6.5 | -2.0152 | 0.276
400 | 6.6667 | -2.0618 | 0.2796
410 | 6.8333 | -2.1074 | 0.2736
420 | 7 | -2.1543 | 0.2814 | 0.2762
430 | 7.1667 | -2.201 | 0.2802
440 | 7.3333 | -2.2483 | 0.2838
450 | 7.5 | -2.2934 | 0.2706
460 | 7.6667 | -2.3387 | 0.2718
470 | 7.8333 | -2.3838 | 0.2706
480 | 8 | -2.4313 | 0.285 | 0.277
490 | 8.1667 | -2.476 | 0.2682
500 | 8.3333 | -2.5225 | 0.279
510 | 8.5 | -2.5679 | 0.2724
520 | 8.6667 | -2.6142 | 0.2778
530 | 8.8333 | -2.6635 | 0.2958
540 | 9 | -2.7115 | 0.288 | 0.2802
550 | 9.1667 | -2.7594 | 0.2874
560 | 9.3333 | -2.8082 | 0.2928
570 | 9.5 | -2.8508 | 0.2556
580 | 9.6667 | -2.895 | 0.2652
590 | 9.8333 | -2.9414 | 0.2784
600 | 10 | -2.9883 | 0.2814 | 0.2768

## FlexCal-H value and u
Reference flow | Bias | Standard uncertainty
502.38 | -0.39% | 0.18%
5001.2 | -0.12% | 0.18%
49911 | 0.4% | 0.19%
Measured flow [ml min-1] | Corrected flow | Calibration Unceratinty (k=1) | Repeatability | Calibration Unceratinty (k=1) | Repeatability | Combined U
13.4056 | 13.4497 | 0.1794% | 0.0175% | 0.024135 | 0.0023537 | 0.024249
13.1393 | 13.1831 | 0.1794% | 0.0941% | 0.023657 | 0.012405 | 0.026712
0.1803%
0.2026%

## Aw(Hg_e)
Type (measured/calc) | Term | Anotation | Unit | Value | Uncertainty
[6 empty] | Atomic mass 196Hg | Atomic mass 198Hg | Atomic mass 199Hg | Atomic mass 200Hg | Atomic mass 201Hg | Atomic mass 202Hg | Atomic mass 204Hg | enrHg Atom fraction 196Hg | enrHg Atom fraction 198Hg | enrHg Atom fraction199Hg | enrHg Atom fraction 200Hg | enrHg Atom fraction 201Hg | enrHg Atom fraction 202Hg | enrHg Atom fraction 204Hg
cert (IUPAC website) | Atomic mass 196Hg | Aw(196Hg) | Da | 195.9658 | 2e-05 | 195.9659 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658 | 195.9658
cert (IUPAC website) | Atomic mass 198Hg | Aw(198Hg) | Da | 197.9668 | 3e-06 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668 | 197.9668
cert (IUPAC website) | Atomic mass 199Hg | Aw(199Hg) | Da | 198.9683 | 4e-06 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683 | 198.9683
cert (IUPAC website) | Atomic mass 200Hg | Aw(200Hg) | Da | 199.9683 | 4e-06 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683 | 199.9683
cert (IUPAC website) | Atomic mass 201Hg | Aw(201Hg) | Da | 200.9703 | 5e-06 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703 | 200.9703
cert (IUPAC website) | Atomic mass 202Hg | Aw(202Hg) | Da | 201.9706 | 5e-06 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706 | 201.9706
cert (IUPAC website) | Atomic mass 204Hg | Aw(204Hg) | Da | 203.9735 | 3e-06 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735 | 203.9735
cert | enrHg Atom fraction 196Hg | A(196Hg)_e | % | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
cert | enrHg Atom fraction 198Hg | A(198Hg)_e | % | 1.63 | 0.02 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63 | 1.65 | 1.63 | 1.63 | 1.63 | 1.63 | 1.63
cert | enrHg Atom fraction199Hg | A(199Hg)_e | % | 91.95 | 0.05 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 91.95 | 92 | 91.95 | 91.95 | 91.95 | 91.95
cert | enrHg Atom fraction 200Hg | A(200Hg)_e | % | 4.92 | 0.03 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.92 | 4.95 | 4.92 | 4.92 | 4.92
cert | enrHg Atom fraction 201Hg | A(201Hg)_e | % | 0.66 | 0.01 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.66 | 0.67 | 0.66 | 0.66
cert | enrHg Atom fraction 202Hg | A(202Hg)_e | % | 0.73 | 0.01 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.73 | 0.74 | 0.73
cert | enrHg Atom fraction 204Hg | A(204Hg)_e | % | 0.11 | 0.01 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12
calculated | 199Hg Atomic weight | Aw(Hg_e) | g mol-1 | 199.081 | 0.13352 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081 | 199.081 | 199.1202 | 199.1206 | 199.1805 | 199.141 | 199.1011 | 199.1012 | 199.1014
[6 empty] | Atomic mass 196Hg | Atomic mass 198Hg | Atomic mass 199Hg | Atomic mass 200Hg | Atomic mass 201Hg | Atomic mass 202Hg | Atomic mass 204Hg | enrHg Atom fraction 196Hg | enrHg Atom fraction 198Hg | enrHg Atom fraction199Hg | enrHg Atom fraction 200Hg | enrHg Atom fraction 201Hg | enrHg Atom fraction 202Hg | enrHg Atom fraction 204Hg
[5 empty] | Squared deviations | 1.6e-17 | 2.3912e-15 | 1.3528e-11 | 3.873e-14 | 1.089e-15 | 1.3322e-15 | 1.089e-17 | 0.0015361 | 0.0015676 | 0.0098971 | 0.0035989 | 0.00040389 | 0.00040792 | 0.00041605
[5 empty] | Budget | 8.975e-14% | 1.341e-11% | 7.588e-08% | 2.172e-10% | 6.109e-12% | 7.473e-12% | 6.109e-14% | 8.616% | 8.793% | 55.52% | 20.19% | 2.266% | 2.288% | 2.334%
Auto updating table from the website
Atomic mass (Da) | Isotopic abundance (amount fraction)
195.965 83(2) | 0.0015(1)
197.966 769(3) | 0.1004(3)
198.968 281(4) | 0.1694(12)
199.968 327(4) | 0.2314(9)
200.970 303(5) | 0.1317(9)
201.970 644(5) | 0.2974(13)
203.973 494(3) | 0.0682(4)

## Mass fraction 199Hg
Dilution with 3 weighings (bottle-analyte-solvent)
Analyte/Std ID | 199Hg Spike (ORNL) SAMD/116 in 20% HNO3 prepared 02/03/2005 |  | Analyst/Lab book |  | SP1 88, 92
Solvent | 10 % v/v c.HCl Romil UpA |  | Date |  | 2022-09-14
Source/Lot no | TBC |  | Balance |  | Ohaus Explorer EX224 S/N C212744832
199Hg Dilution 1 - 25/04/2023 SP1 pg. 122
Bottle | 12.5139 | Data: | value | repeat | linearity | std. uncert. |  | n | derivatives | variances
| 12.5141 | m1 | 12.5138 | 0.00020817 | 9.2376e-06 | 0.00020837 | g | 3 | -14.6172 | 9.2769e-06 | 0.97953 | %
| 12.5134 | m2 | 12.7224 | 0.00015275 | 9.2376e-06 | 0.00015303 | g | 3 | 14.6781 | 5.0455e-06 | 0.53274 | %
|  | m3 | 62.8046 | 0.00020817 | 0.00012124 | 0.0002409 | g | 3 | -0.060883 | 2.1511e-10 | 2.2713e-05 | %
|  | wstock | 738.1719 |  |  | 7.3631 | µg/g |  | 0.0041479 | 0.00093276 | 98.4877 | %
Bottle +analyte | 12.7226 | [7 empty] | sum | 0.00094708
| 12.7225 | Results | mass (g) | uncert. | unit | rsu (%)
| 12.7221 | wAnalyte | 3.0618 | 0.030775 | µg/g | 1.0051 | %
|  | wAnalyte | 3061.8455 | 30.7746 | ng/g | 1.0051 | %
Bottle +analyte+solvent | 62.8049 | m(Analyte) | 0.2086 | 0.00025853 | g | 0.12394 | %
| 62.8047 | m (Solution) | 50.2908 | 0.00031852 | g | 0.00063335 | %
62.8042
199Hg Dilution 2 - 25/04/2023 SP1 pg. 122
Bottle | 12.7236 | Data: | value | repeat | linearity | std. uncert. |  | n | derivatives | variances
| 12.7236 | m1 | 12.7237 | 6.6667e-05 | 9.2376e-06 | 6.7304e-05 | g | 3 | -59.6858 | 1.6137e-05 | 0.0028243 | %
| 12.7238 | m2 | 13.7537 | 0 | 9.2376e-06 | 9.2376e-06 | g | 3 | 60.9349 | 3.1685e-07 | 5.5456e-05 | %
|  | m3 | 62.9715 | 0.33333 | 0.00012124 | 0.33333 | g | 3 | -1.2491 | 0.17336 | 30.3426 | %
|  | wstock | 3061.8455 |  |  | 30.7746 | ng/g |  | 0.020499 | 0.39797 | 69.6545 | %
Bottle +analyte | 13.7537 | [7 empty] | sum | 0.57135
| 13.7537 | Results | mass (g) | uncert. | unit | rsu (%)
| 13.7537 | wAnalyte | 62.765 | 0.75588 | ng/g | 1.2043 | %
|  | wAnalyte | 62764.9952 | 755.8796 | pg/g | 1.2043 | %
Bottle +analyte+solvent | 62.3048 | m(Analyte) | 1.03 | 6.7935e-05 | g | 0.0065954 | %
| 63.3048 | m (Solution) | 50.2478 | 0.33333 | g | 0.66338 | %
63.3048
0.5 ppb 199Hg - 25/04/2023 SP1 pg. 123
Bottle | 12.552 | Data: | value | repeat | linearity | std. uncert. |  | n | derivatives | variances
| 12.5518 | m1 | 12.552 | 0.00011547 | 9.2376e-06 | 0.00011584 | g | 3 | -1.2376 | 2.0552e-08 | 0.053553 | %
| 12.5522 | m2 | 12.9641 | 8.8192e-05 | 9.2376e-06 | 8.8674e-05 | g | 3 | 1.2478 | 1.2243e-08 | 0.031902 | %
|  | m3 | 62.8519 | 0.00013333 | 0.00012124 | 0.00018022 | g | 3 | -0.010222 | 3.3938e-12 | 8.8432e-06 | %
|  | wstock | 62.765 |  |  | 0.75588 | ng/g |  | 0.0081922 | 3.8345e-05 | 99.9145 | %
Bottle +analyte | 12.9639 | [7 empty] | sum | 3.8378e-05
| 12.9641 | Results | mass (g) | uncert. | unit | rsu (%)
| 12.9642 | wAnalyte | 0.51418 | 0.006195 | ng/g | 1.2048 | %
|  | wAnalyte | 514.1828 | 6.195 | pg/g | 1.2048 | %
Bottle +analyte+solvent | 62.8522 | m(Analyte) | 0.41207 | 0.00014588 | g | 0.035403 | %
| 62.8518 | m (Solution) | 50.2999 | 0.00021423 | g | 0.00042591 | %
62.8518
Note on H12: Sophie Page: Concentration and uncertainty from reverse IDMS of 199Hg from 10/02/2022 SP/PP (T:\QT20\Inorganic Speciation Analysis Team\Projects\Current\CBM2125\WP2\220210 Reverse IDMS 199Hg)