## Sheet1
Paper | Sample Types | Processing Volume | Elution Volume | Application Domain | Automation Level | Duration of the extraction | DNA Quantitative Characterisation | cfDNA Recovery characterisation method | cfDNA Recovery (Caution with interpretration: Not for comparison unless same method used) | Benchmark choice and method | Comparison to benchmark (Caution with interpretration: Not for comparison unless same method used)
Yang et al. 2015 | Plasma, Whole blood (patients with severe sepsis and healthy controls) | <10 μL droplet | Not specified | Sepsis, ICU mortality prognosis | Semi-automated | 5 min | PicoGreen fluorescence labeling, Electrophoresis | Not available | n/a | Not mentioned | n/a
10.1039/C5LC00681C
Lab on a Chip
Kim et al. 2018 | Whole blood (>3 mL), Plasma from NSCLC patients | >3 mL whole blood | Not specified | Cancer (non-small cell lung cancer), Cancer therapy monitoring | Fully automated | 30 min | qPCR, Droplet digital PCR | Synthetic short DNA (300 base pairs) spiked into the human serum to mimic cfDNA, with the recovery rate quantified by real-time PCR. The concentration/amount spiked was not available | 50-75% | Qiagen QIAamp DNA Blood Mini Kit (Q-BM) and QIAamp cfDNA Kit (Q-cfNDA) Synthetic short DNA (300 base pairs) spiked into the human serum to mimic cfDNA, with the recovery rate quantified by real-time PCR. The concentration/amount spiked was not available. Characterised with real-time qPCR | Similar performance
10.1039/C8LC00165K
Lab on a Chip
Kye et al. 2020 | Synthetic cfDNA at a concentration of 50 ng/mL and 250 ng/mL in DEPC-DW with 4,000 ppm PEO (Sigma Aldrich, USA) at a concentration of 50 ng/mL and 250 ng/mL. | 1 mL sample volume | 50 μL | Cancer diagnostics, Cancer prognosis | Semi-automated | At max flowrate (150 μL/min) would equate to 6-7 min | PCR amplification, Gel electrophoresis | Not available | n/a | Not specified | n/a
10.1007/s13206-020-4208-1
BioChip Journal
Campos et al. 2018 | Plasma from healthy donors, NSCLC patients, Colorectal cancer patients | Not clearly stated. Paper states ambition to process > 5mL | Not explicitly stated | Cancer (NSCLC, colorectal cancer), Oncology | Manual / Semi-automated (pre-processing of the plasma, which included digestion with proteinase K, the sample was mixed with the IB (1:3)) | At max flowrate (2 μL/min), 1mL would take 8 hours 20 min | Qubit quantification, Ligase detection reaction (LDR) for KRAS mutations, qPCR | Method 1: Using a DNA ladder consisting of 50–700 bp double-stranded DNA fragments. DNA ladder was spiked into PBS (concentration not available) ( presumed then DNA measured with Tape station) Method 2: Spiked plasma from healthy donors with the 122 bp fragment bearing KRAS mutations. qPCR detection (spiked concentration/amount not available) | Method 1: >90% for 100-700bp; >70% for 50bp Method 2: 92% recovered | Benchmarks: Norgen and Qiagen cfDNA kits. Method: Using a DNA ladder consisting of 50–700 bp double-stranded DNA fragments. DNA ladder was spiked into PBS (concentration/amount not available) | Similar performance
10.1039/C8LC00716K
Lab on a Chip
Gwak et al. 2019 | cfDNA-spiked PBS at 1 ng/ml, Serum from breast cancer patients and healthy controls, Whole blood | 500 μL serum | 50 μL | Cancer (breast cancer), Cancer diagnostics | Semi-automated | 19 min | Qubit quantification, Real-time PCR, qPCR | cfDNA-spiked (IL2) PBS at 1 ng/ml characterised by qPCR | 51.8% | Benchmark: MagMAX Cell-Free DNA Isolation Kit (Applied Biosystems) - Should be noted that the figures on the paper state "Accugene" which is a different kit. Method: cfDNA-spiked (IL2) PBS at 1 ng/ml characterised by qPCR | Similar performance
10.1063/1.5100009
Biomicrofluidics
Lee et al. 2020 | Plasma; Synthetic plasma with cfDNA reference standards | 0.5-1 mL plasma | 50-150 μL (adjustable) | Cancer (HER-2 breast cancer with liver metastasis), Cancer monitoring | Fully automated | 15 min | ddPCR, NGS, qPCR | DNA extraction efficiency was calculated based on specific DNA recovery by comparing the absolute copy number of DNA detected in qPCR to the theoretically expected value assuming 100% DNA extraction efficiency. To measure the DNA recovery rate, the 180bp lambda DNA PCR product was spiked into the plasma of a healthy control. For comparison, the volumes of eluent and the plasma were fixed at 50μL and 1mL. Concentration of 10 to 100 ng/mL of input DNA | 70-80% | Benchmark: Qiagen QIAamp cfDNA. Method: The comparative extraction efficiency between QIAamp and PIBEX chip was investigated by Ct value calculation. In addition, the levels of the four genes in cfDNA healthy plasma samples (N=8) extracted by either QIAamp or PIBEX chip were compared. Real-time PCR characterisation. | Similar performance
10.1038/s41698-019-0107-0
npj Precision Oncology
Schneider et al. 2022 | cfDNA in AcroMetrix EDTA Plasma Dilution Matrix | 1 ml of AcroMetrix EDTA Plasma Dilution Matrix | 10-50 μL typical | Prenatal diagnosis (NIPT), Prenatal screening | Semi-automated | 20 min | Agilent Bioanalyzer | cfDNA eluate sample yield was performed using the Agilent Bioanalyzer 2100 Machine (Agilent Technologies). The Agilent DNA 1000 Kit and DNA Chip (Agilent Technologies) were used for On-Chip Electrophoresis, and the resulting electropherogram traces were analyzed to identify the DNA concentration (ng/μl) of 149-bp cfDNA input sample and the eluate to calculate the cfDNA yield. As small as 16.56 ng/ml and as large as 58.92 ng/ml was used for the input sample. | 60-70% | Not mentioned | n/a
10.1002/pd.6092
Prenatal Diagnosis
Hu et al. 2019 | Whole blood (4 mL) | 4 mL whole blood 1.6 mL plasma | Not specified | Not specified (liquid biopsy applications) | Semi-automated | 15 min | qPCR; Droplet digital PCR | T790 M mutation (varying degrees of mutation including 5%, 20%, and 50%) spiked in the whole blood in ctDNA isolated from whole blood samples. low concentration samples (101–104 copies). High concentration samples (104–107). Quantified by ddPCR. | 30-65% | Benchmark: MagMAX Cell-Free Total Nucleic Acid Kit. Method: T790 M mutation (varying degrees of mutation including 5%, 20%, and 50%) spiked in the whole blood in ctDNA isolated from whole blood samples. low concentration samples (101–104 copies). High concentration samples (104–107). Quantified by ddPCR. | Similar performance
10.1039/C9AN00493A
Analyst
Jin et al. 2018 | Spiked buffer, 14 samples from prospective CRC patients, and 10 from healthy controls | 500 μL plasma samples | Not specified | Cancer (colorectal cancer), KRAS and BRAF mutation detection | Semi-automated | 20 min | qPCR | 200 µL of the ALU gene (247 bp amplicon) was mixed with 100 µL of DTBP (100 mg/mL) and then was injected into the platform. PCR (only Cycle Numbers provided) for Actin Gene | 1 Ct cycle difference between input and chip eluate could be cautiously interpreted as 50% recovery | Benchmark: Qiagen QIAamp cfDNA Kit . Method: Characterised on plasma samples with non calibrated non quantitative PCR (only Cycle Numbers provided) for Actin Gene. Unclear wether on same volumes and biological replicates | Information in the paper is conflicting and no statistical analysis provided
10.1002/advs.201800614
Advanced Science
This paper | Plasma from healthy donors (400 samples) and hospital patients with suspected sepsis (9 patients) | 1-4 mL plasma | 50 μL | Sepsis, Pathogen identification | Fully automated | 40 min | qPCR, mNGS | Identical healthy human plasma samples processed in same volume on the benchmark manual extraction and microfluidic device. qPCR on LINE sequence. Recovery expressed as percentage relative to benchmark amount | Relative recovery to benchmark is 100% ( averaged over 333 samples) | Benchmark: Qiagen QIAamp cfDNA kit. Method:Characterised with quantitative PCR on LINE sequence on matched 1-4mL plasma samples (biological replicates N=458, considered for statistics N=333). | Similar performance