## Table S1 (Immunassays)
Table S1 Recent examples of µPADs (2018-2025) for medical diagnostic applications (immunoassays)
This is an extended version of Table 1 in the main text
Columns not shown in the main text tables are indicated with light blue colour column headings
(Blood separation) in parentheses indicates that no specific information is provided in the cited work. Since serrum is mentioned as the sample, a prior serum separation from whole blood is assumed
Analyte | Sample matrix | Real sample test | Sample type category | Concentration of measured clinical or spiked sample | Detection technique | Detection range | Detection limit | Storage stability | Assay time | Necessity of time control | Off-device pretreatment | Operation steps | Reagent handling | External equipment | Sample volume control | Article title | Ref
Tumor necrosis factor (TNF)-α | Rat tissue sample | Clinical sample (rat) | 5 | 231.6 ± 103.09 pM (4 weeks after surgery) 69.9 ± 10.4 pM (2 days after surgery) | Colorimetric | N.A. | 22 pM (PBS) | N.A. | 31 min | No | No | Sample application Continuous flow of PBS | No Microcontroller driven | RGB color sensor Box with heater Pump | 3 µL | A paper-based microfluidic platform with shape-memory-polymer-actuated fluid valves for automated multi-step immunoassays | H. Fu, P. Song, Q. Wu, C. Zhao, P. Pan, X. Li, N. Y. K. Li-Jessen and X. Liu, Microsyst. Nanoeng., 2019, 5, 50
Human interleukin-6 (IL-6) | Human saliva | Clinical sample | 5 | 3.2–22.5 pg/mL | Colored length | 0.05–25.0 pg/mL | 0.05 pg/mL | Sealed bag 4°C ~7 weeks (no decrease) ~12 weeks (93.6%) | 35 min | Yes Multiple operations | No | Sample application Reagent application Removal and insertion of paper bridge Reagent application | Yes HRP–antibody H2O2 | No | 15 µL | Distance-Based All-In-One Immunodevice for Point-of-Care Monitoring of Cytokine Interleukin-6 | K. Khachornsakkul, W. Dungchai and N. Pamme, ACS Sens., 2022, 7, 2410-2419
Human interleukin-6 (IL-6) | Human urine | Clinical sample | 5 | 3.88–18.10 pg/mL | Colored length | 0.05–25.0 pg/mL | 0.05 pg/mL | Sealed bag 4°C ~7 weeks (no decrease) ~12 weeks (93.6%) | 35 min | Yes Multiple operations | No | Sample application Reagent application Removal and insertion of paper bridge Reagent application | Yes HRP–antibody H2O2 | No | 15 µL | Distance-Based All-In-One Immunodevice for Point-of-Care Monitoring of Cytokine Interleukin-6 | K. Khachornsakkul, W. Dungchai and N. Pamme, ACS Sens., 2022, 7, 2410-2419
SARS-CoV-2 | Exhaled breath | Clinical sample | 4 | N.A. | Echem (DPV) | 3.7–10,000 ng/mL (Exhaled Breath) | 1 ng/mL (Exhaled Breath) | 4°C >33 days >90% | 60 min | Yes Multiple operations | Off–chip post washing and detection | Reagent application Sample application Off–chip washing and echem measurement | Yes Antibody 1–Naphthyl phosphate Disodium salt | Microtube Potentiostat | >50 µL | Functionalized N95 Face Mask with a Chemical-Free Paper-Based Collector for Exhaled Breath Analysis: SARS-CoV-2 Detection with a Printed Immunosensor as a Case Study | L. Gutiérrez-Gálvez, N. Seddaoui, L. Fiore, L. Fabiani, T. García-Mendiola, E. Lorenzo and F. Arduini, ACS Sens., 2024, 9, 4047-4057
Hepatitis B virus (HBV) | Human serum | Clinical sample | 4 | 44.61–420 pg/mL | Echem (CA) | 0.1–250 ng/mL (PBS) | 8.2 pg/mL (PBS) | –20°C >30 days 95.3–106.6 % | <500 s | No Immediate measurement | Blood separation | Sample application Buffer application Echem measurement | No | Potentiostat | 15 µL | An automated fast-flow/delayed paper-based platform for the simultaneous electrochemical detection of hepatitis B virus and hepatitis C virus core antigen | S. Boonkaew, A. Yakoh, N. Chuaypen, P. Tangkijvanich, S. Rengpipat, W. Siangproh and O. Chailapakul, Biosens. Bioelectron., 2021, 193, 113543
Hepatitis C virus (HCV) | Human serum | Clinical sample | 4 | 84.84–205.85 pg/mL | Echem (CA) | 0.001–250 ng/mL (PBS) | 1.19 pg/mL (PBS) | –20°C >30 days 95.3–106.6 % | <500 s | No Immediate measurement | Blood separation | Sample application Buffer application Echem measurement | No | Potentiostat | 15 µL | An automated fast-flow/delayed paper-based platform for the simultaneous electrochemical detection of hepatitis B virus and hepatitis C virus core antigen | S. Boonkaew, A. Yakoh, N. Chuaypen, P. Tangkijvanich, S. Rengpipat, W. Siangproh and O. Chailapakul, Biosens. Bioelectron., 2021, 193, 113543
SARS-CoV-2 | Human serum | Clinical sample | 4 | N.A. | Colorimetric | N.A. | 5 nM (serum) | N.A. | 15 min | Yes Multiple operations | (Blood separation) Treated with surfactant (biosafety reason) Pre–incubation | Mixing with reagent Sample application Buffer application Reagent application Imaging | Yes HRP–protein Immobilization protein TMB | Microtube Imager | 25 µL (off–chip) 50 µL (on–chip) | Vertical Flow Cellulose-Based Assays for SARS-CoV-2 Antibody Detection in Human Serum | S. Kim, Y. Hao, E. A. Miller, D. M. Y. Tay, E. Yee, P. Kongsuphol, H. Jia, M. McBee, P. R. Preiser and H. D. Sikes, ACS Sens., 2021, 6, 1891-1898
SARS-CoV-2 | Human serum | Clinical sample | 4 | N.A. | Colorimetric | N.A. | N.A. 89.5% accuracy | N.A. | <20 min | Yes Multiple operations | (Blood separation) | Buffer application Sample application Buffer application Changing top case Buffer and reagent application Detachement and imaging | Yes AuNP–antibody | Plastic box Smartphone | 40 µL | A Paper-Based Multiplexed Serological Test to Monitor Immunity against SARS-COV-2 Using Machine Learning | M. Eryilmaz, A. Goncharov, G.-R. Han, H.-A. Joung, Z. S. Ballard, R. Ghosh, Y. Zhang, D. Di Carlo and A. Ozcan, ACS Nano, 2024, 18, 16819-16831
Cardiac troponin I (cTnI) | Human serum | Clinical sample | 4 | <40–4290 pg/mL | Colorimetric | –100 ng/mL (serum) | 0.2 pg/mL (serum) | N.A. | <15 min | Yes Multiple operations | (Blood separation) | Buffer application Sample and reagent application Buffer application Changing top case Reagent application Detachement and imaging | Yes AuNP–antibody HAuCl4 NH2OH | Plastic box Smartphone | 50 µL | Deep Learning-Enhanced Paper-Based Vertical Flow Assay for High-Sensitivity Troponin Detection Using Nanoparticle Amplification | G.-R. Han, A. Goncharov, M. Eryilmaz, H.-A. Joung, R. Ghosh, G. Yim, N. Chang, M. Kim, K. Ngo, M. Veszpremi, K. Liao, O. B. Garner, D. Di Carlo and A. Ozcan, ACS Nano, 2024, 18, 27933–27948
Cardiac troponin I (cTnI) | Human serum | Clinical sample | 4 | few–18 000 pg/mL | CL | 0.001–100 ng/mL (serum) | 0.16 pg/mL (serum) | N.A. | 25 min | Yes Multiple operations | (Blood separation) | Buffer application Sample and reagent application Buffer application Changing top case Buffer application Changing top case Reagent application CL imaging | Yes AuNP-HEP-antibody Luminol H2O2 | Plastic box Portable reader | 50 µL | Deep Learning-Enhanced Chemiluminescence Vertical Flow Assay for High-Sensitivity Cardiac Troponin I Testing | G.-R. Han, A. Goncharov, M. Eryilmaz, S. Ye, H.-A. Joung, R. Ghosh, E. Ngo, A. Tomoeda, Y. Lee, K. Ngo, E. Melton, O. B. Garner, D. Di Carlo and A. Ozcan, Small, 2025, 21, 2411585
Anti-Borrelia burgdorferi antibody | Human serum | Clinical sample | 4 | N.A. | Colorimetric | N.A. | N.A. Sensitivity 90.5% Specificity 87.0% (n = 96) | N.A. | <15 min | Yes Multiple operations | (Blood separation) | Buffer application Sample application Change of top case Buffer application Disassemble Imaging | No | 3D–printed cartridge Smartphone | 20 µL | Point-of-Care Serodiagnostic Test for Early-Stage Lyme Disease Using a Multiplexed Paper-Based Immunoassay and Machine Learning | H.-A. Joung, Z. S. Ballard, J. Wu, D. K. Tseng, H. Teshome, L. Zhang, E. J. Horn, P. M. Arnaboldi, R. J. Dattwyler, O. B. Garner, D. Di Carlo and A. Ozcan, ACS Nano, 2020, 14, 229-240
Anti-Borrelia burgdorferi antibody | Human serum | Clinical sample | 4 | N.A. | Colorimetric | N.A. | N.A. Sensitivity 95.5% Specificity 100.0% | Room temperature >60 days 95% | <20 min | Yes Multiple operations | (Blood separation) | Buffer and sample application Changing top case Reagent application Detachement and imaging | Yes AuNP–antibody | Plastic box Smartphone | 20 µL | Rapid single-tier serodiagnosis of Lyme disease | R. Ghosh, H.-A. Joung, A. Goncharov, B. Palanisamy, K. Ngo, K. Pejcinovic, N. Krockenberger, E. J. Horn, O. B. Garner, E. Ghazal, A. O’Kula, P. M. Arnaboldi, R. J. Dattwyler, A. Ozcan and D. Di Carlo, Nat. Commun., 2024, 15, 7124
Myoglobin | Human serum | Clinical sample | 4 | 8.6–75 ng/mL | FL | N.A. | 0.52 ng/mL (serum) | N.A. | <15 min | Yes Multiple operations | Blood separation Pre–mix with antibody | Buffer application Pre–mixing sample and reagent Sample application Buffer application Removing top case Inserting to phone reader Imaging | Yes AuNP-antibody | Plastic box Smartphone Light source and filter | 50 µL | Deep Learning-Enabled Multiplexed Point-of-Care Sensor using a Paper-Based Fluorescence Vertical Flow Assay | A. Goncharov, H.-A. Joung, R. Ghosh, G.-R. Han, Z. S. Ballard, Q. Maloney, A. Bell, C. T. Z. Aung, O. B. Garner, D. D. Carlo and A. Ozcan, Small, 2023, 19, 2300617
Creatine kinase-MB (CK-MB) | Human serum | Clinical sample | 4 | 1.1–23.4 ng/mL | FL | N.A. | 0.3 ng/mL (serum) | N.A. | <15 min | Yes Multiple operations | Blood separation Pre–mix with antibody | Buffer application Pre–mixing sample and reagent Sample application Buffer application Removing top case Inserting to phone reader Imaging | Yes AuNP-antibody | Plastic box Smartphone Light source and filter | 50 µL | Deep Learning-Enabled Multiplexed Point-of-Care Sensor using a Paper-Based Fluorescence Vertical Flow Assay | A. Goncharov, H.-A. Joung, R. Ghosh, G.-R. Han, Z. S. Ballard, Q. Maloney, A. Bell, C. T. Z. Aung, O. B. Garner, D. D. Carlo and A. Ozcan, Small, 2023, 19, 2300617
Heart-type fatty acid binding protein (FABP) | Human serum | Clinical sample | 4 | 1.3–45.7 ng/mL | FL | N.A. | 0.49 ng/mL (serum) | N.A. | <15 min | Yes Multiple operations | Blood separation Pre–mix with antibody | Buffer application Pre–mixing sample and reagent Sample application Buffer application Removing top case Inserting to phone reader Imaging | Yes AuNP-antibody | Plastic box Smartphone Light source and filter | 50 µL | Deep Learning-Enabled Multiplexed Point-of-Care Sensor using a Paper-Based Fluorescence Vertical Flow Assay | A. Goncharov, H.-A. Joung, R. Ghosh, G.-R. Han, Z. S. Ballard, Q. Maloney, A. Bell, C. T. Z. Aung, O. B. Garner, D. D. Carlo and A. Ozcan, Small, 2023, 19, 2300617
C-Reactive Protein (CRP) | Human serum | Diluted clinical sample | 3 | 0–10, 83.6 mg/L | Colorimetric | 0–10 mg/L (serum) | N.A. | N.A. | <12 min | Yes Multiple operations | (Blood separation) 10 times dilution with buffer | Buffer application Sample and reagent application Running buffer Disassemble Imaging | Yes AuNP-antibody | 3D–printed cartridge Smartphone | 50 µL | Deep learning-enabled point-of-care sensing using multiplexed paper-based sensors | Z. S. Ballard, H.-A. Joung, A. Goncharov, J. Liang, K. Nugroho, D. Di Carlo, O. B. Garner and A. Ozcan, NPJ Digit. Med., 2020, 3, 66
Anti-interferon-γ autoantibodies (anti-IFN-γ Abs) | Human serum | Diluted clinical sample | 3 | 111.09–1937.44 µg/mL | Colorimetric | 0.003–3 μg/mL (PBS) | 0.001 μg/mL (smartphone) 0.03 µg/mL (naked eye) (PBS) | N.A. | <10 min | No >10 min | (Blood separation) 100-fold dilution with PBS | Sample and buffer application Imaging | No | (Smartphone) | 20 µL | One-Step Gold Nanoparticle Signal Amplification in a 3D Paper-Based Device for Rapid Detection of Anti-Interferon-γ Autoantibodies | Y.-S. Chien, U.-I. Wu, Y.-H. Pan and C.-F. Chen, Biosens. Bioelectron., 2025, 287, 117724
SARS-CoV-2 | Human saliva | Spiked into biofluid sample | 2 | 10–160 pg/mL | Colorimetric | 10–160 pg/mL (saliva) | 10 pg/mL (saliva) | N.A. | 8 min | Yes Multiple operations | No | Reagent application Sample application Centrifugation for 5 times Imaging | Yes AuNP-antibody | Centrifuge Black box Smartphone | 25 µL | μPADs on Centrifugal Microfluidic Discs for Rapid Sample-to-Answer Salivary Diagnostics | S. Liu, Y. Hou, Z. Li, C. Yang and G. Liu, ACS Sens., 2023, 8, 3520-3529
Cardiac troponin I (cTnI) | Human whole blood | Spiked into biofluid sample | 2 | 90 pg/mL–90 ng/mL | Echem (EIS) | 90 pg/mL to 90 ng/mL (artificial blood plasma) | 4.6 pg/mL (190 fM) (artificial blood plasma) | N.A. | 46 min | Yes Multiple operations | No | Folding and sample application Floding and buffer application Detaching Peeling off and folding paper Reagent application Echem meausrement | Yes Fe(CN)63-/4- | Potentiostat | 20 µL | Paper-Based All-in-One Origami Nanobiosensor for Point-of-Care Detection of Cardiac Protein Markers in Whole Blood | H. Fu, Z. Qin, X. Li, Y. Pan, H. Xu, P. Pan, P. Song and X. Liu, ACS Sens., 2023, 8, 3574-3584
Brain natriuretic peptide (BNP) | Human whole blood | Spiked into biofluid sample | 2 | 6 pg/mL–6 ng/mL | Echem (EIS) | 6 pg/mL to 6 ng/mL (artificial blood plasma) | 1.2 pg/mL (40 fM) (artificial blood plasma) | N.A. | 46 min | Yes Multiple operations | No | Folding and sample application Floding and buffer application Detaching Peeling off and folding paper Reagent application Echem meausrement | Yes Fe(CN)63-/4- | Potentiostat | 20 µL | Paper-Based All-in-One Origami Nanobiosensor for Point-of-Care Detection of Cardiac Protein Markers in Whole Blood | H. Fu, Z. Qin, X. Li, Y. Pan, H. Xu, P. Pan, P. Song and X. Liu, ACS Sens., 2023, 8, 3574-3584
D-Dimer | Human whole blood | Spiked into biofluid sample | 2 | 91 ng/mL–91 µg/mL | Echem (EIS) | 91 ng/mL to 91 μg/mL (artificial blood plasma) | 146 pg/mL (730 fM) (artificial blood plasma) | N.A. | 46 min | Yes Multiple operations | No | Folding and sample application Floding and buffer application Detaching Peeling off and folding paper Reagent application Echem meausrement | Yes Fe(CN)63-/4- | Potentiostat | 20 µL | Paper-Based All-in-One Origami Nanobiosensor for Point-of-Care Detection of Cardiac Protein Markers in Whole Blood | H. Fu, Z. Qin, X. Li, Y. Pan, H. Xu, P. Pan, P. Song and X. Liu, ACS Sens., 2023, 8, 3574-3584
HIV1 | Porcine whole blood | Spiked into biofluid sample | 2 | 50–150 nM | BL | N.A. | 2.8 nM (porcine serum) | Capped vials with argon, silica gel -20°C >2 months | 21 min | No 21–30 min* | No | Sample application Imaging | No | Camera Box | 30–60 µL | Paper-Based Antibody Detection Devices Using Bioluminescent BRET-Switching Sensor Proteins | K. Tenda, B. Van Gerven, R. Arts, Y. Hiruta, M. Merkx and D. Citterio, Angew. Chem. Int. Ed., 2018, 57, 15369-15373
Hemagglutinin (HA) | Porcine whole blood | Spiked into biofluid sample | 2 | 50–150 nM | BL | N.A. | 7.1 nM (porcine serum) | Capped vials with argon, silica gel -20°C >2 months | 21 min | No 21–30 min* | No | Sample application Imaging | No | Camera Box | 30–60 µL | Paper-Based Antibody Detection Devices Using Bioluminescent BRET-Switching Sensor Proteins | K. Tenda, B. Van Gerven, R. Arts, Y. Hiruta, M. Merkx and D. Citterio, Angew. Chem. Int. Ed., 2018, 57, 15369-15373
Dengue virus type 1 (DEN1) | Porcine whole blood | Spiked into biofluid sample | 2 | 50–150 nM | BL | N.A. | 19.3 nM (porcine serum) | Capped vials with argon, silica gel -20°C >2 months | 21 min | No 21–30 min* | No | Sample application Imaging | No | Camera Box | 30–60 µL | Paper-Based Antibody Detection Devices Using Bioluminescent BRET-Switching Sensor Proteins | K. Tenda, B. Van Gerven, R. Arts, Y. Hiruta, M. Merkx and D. Citterio, Angew. Chem. Int. Ed., 2018, 57, 15369-15373
Hepatitis B virus (HBV) | Porcine whole blood | Spiked into biofluid sample | 2 | 0.005–200 ng/mL | FL | N.A. | 2.8 ng/mL (porcine whole blood) | Wrapped with reagent paper in aluminum bag with a silica gel pack -20°C >28 days >80% | 45 min | Yes Multiple operations | No | Sample application Folding and buffer application Folding and buffer application Discarding and folding Buffer application Discarding FL meausrement | No | Smarthone Fluorescence readout platform | 5 µL | Origami Paper-Based Immunoassay Device with CRISPR/Cas12a Signal Amplification | H. Suzuki, G. Tong, P. Nath, Y. Hiruta and D. Citterio, ACS Sens., 2025, 10, 1811-1821
HIV-1 | Human plasma | Spiked into biofluid sample | 2 | 0–30 ng/mL | Colorimetric | 0.03–3 ng/mL (plasma) | 0.03 ng/mL (plasma) | N.A. | <10 min | Yes (color timer on device) | Blood separation | Sample application Pulling device and buffer application Pulling device and reagent application Imaging | Yes TMB | (Smartphone) | 3 µL | An electricity- and instrument-free infectious disease sensor based on a 3D origami paper-based analytical device | C.-A. Chen, H. Yuan, C.-W. Chen, Y.-S. Chien, W.-H. Sheng and C.-F. Chen, Lab Chip, 2021, 21, 1908-1915
α-Fetoprotein | Human serum | Spiked into biofluid sample | 2 | 0.1–100 ng/mL | Echem (EIS) | 0.01–500 ng/mL (9 min) 0.5 pg/mL to 10 ng/mL (30 min) | 1.65 pg/mL (9 min) 3.54 fg/mL (30 min) | 4°C <15 days 105% | 9 min 30 min | Yes Multiple operations | (Blood separation) | Sample and buffer application Device rotation for 4 times Echem measurement | No | Potentiostat | 20 µL | Hand-Operated, Paper-Based Rotational Vertical-Flow Immunosensor for the Impedimetric Detection of α-Fetoprotein | A. Yakoh, E. Mehmeti, K. Kalcher and S. Chaiyo, Anal. Chem., 2022, 94, 5893-5900
Procalcitonin | Synthetic serum | Spiked into artificial biofluid sample | 1 | 2–10 µg/mL | AR | 2–10 µg/mL (serum) | 2 µg/mL (serum) | N.A. | 84 min | Yes Multiple operations | (Blood separation) Pre–mix with anti–target antibody | Pre-incubation Sample application Washing Reagent application Washing for three times Imaging | Yes Biotin-antibody Streptavidin-AuNP | Smartphone | 1 µL | A Robust and User-Friendly Alternative to Densitometry Using Origami Biosensors and Digital Logic | S. M. Russell, A. Alba-Patiño, M. Borges and R. de la Rica, ACS Sens., 2018, 3, 1712-1718
α-fetoprotein | Buffer | N.A. | - | N.A. | Echem (CA) | 10–100 ng/mL (KNO3) | 0.63 ng/mL (KNO3) | N.A. | 33 min | Yes Multiple operations | N.A. | Sample application Washing Reagent application Moving PAD Echem measurement | Yes Fe(CN)63-/4- | Potentiostat | 1 µL | 3D Capillary-Driven Paper-Based Sequential Microfluidic Device for Electrochemical Sensing Applications | A. Yakoh, S. Chaiyo, W. Siangproh and O. Chailapakul, ACS Sens., 2019, 4, 1211-1221
Plasmodium falciparum histidine-rich protein 2 (PfHRP2) | PBS | N.A. | - | N.A. | MS | N.A. | 5 nM (PBS) | N.A. | >10 min | Yes Multiple operations | N.A. | Sample and buffer application Removal of capture layer Device assembly NH4OH application MS analysis | Yes NH4OH | MS spectrometer | 20 µL | Automated Immunoassay Performed on a 3D Microfluidic Paper-Based Device for Malaria Detection by Ambient Mass Spectrometry | S. Jackson, S. Lee and A. K. Badu-Tawiah, Anal. Chem., 2022, 94, 5132-5139.

## Table S2 (Nucleic acid tests)
Table S2 Recent examples of µPADs (2018-2025) for medical diagnostic applications (nucleic acid testing)
This is an extended version of Table 2 in the main text
Columns not shown in the main text tables are indicated with light blue colour column headings
Analyte | Sample matrix | Real sample test | Sample type category | Concentration of measured clinical or spiked sample | Detection technique | Detection range | Detection limit | Storage stability | Assay time | Necessity of time control | Off-device pretreatment | Operation steps | Reagent handling | External equipment | Sample volume control | Article title | Ref
Circulating tumor DNA (ctDNA) | Mice serum | Clinical sample (mice) | 4 | N.A. | Colorimetric | 200–1500pM (serum) | 95.81 pM (serum) | 4°C >15 days ~90% | 63 min (on chip) | Yes Multiple operations | Blood separation DNA extraction Annealing Pre–mix with probe | Pre-mixing of sample and reagent Sample application Buffer application Cutting device Reagent application Imaging | Yes Nanozyme–DNA TMB H2O2 | Culture dish Incubator Smartphone 3D–printed dark box | 2.5 µL | Pop-Up Paper-Based Biosensor for a Dual-Mode Lung Cancer ctDNA Assay Based on Novel CoB Nanosheets with Dual-Enzyme Activities and a Portable Smartphone/Barometer for Readout | S. Yang, J. Zhu, L. Yang, H. Fa, Y. Wang, D. Huo, C. Hou, D. Zhong and M. Yang, ACS Sens., 2025, 10, 133-147
Circulating tumor DNA (ctDNA) | Mice serum | Clinical sample (mice) | 4 | N.A. | 63 min (on chip) | 200–1500pM (serum) | 79.88 pM (serum) | 4°C >15 days ~90% | 63 min (on chip) | Yes Multiple operations | Blood separation DNA extraction Annealing Pre–mix with probe | Pre-mixing of sample and reagent Sample application Buffer application Cutting device Reagent application Imaging | Yes Nanozyme–DNA TMB H2O2 | Culture dish Incubator Microtube Gas pressure meter | ~2.5 µL (no details) | Pop-Up Paper-Based Biosensor for a Dual-Mode Lung Cancer ctDNA Assay Based on Novel CoB Nanosheets with Dual-Enzyme Activities and a Portable Smartphone/Barometer for Readout | S. Yang, J. Zhu, L. Yang, H. Fa, Y. Wang, D. Huo, C. Hou, D. Zhong and M. Yang, ACS Sens., 2025, 10, 133-147
SARS-CoV-2 | Human saliva | Clinical sample | 5 | N.A. | Colorimetric | N.A. | 50 copies/μL (saliva) | N.A. | 60 min | Yes Single reaction Automated sequential reagent addition | No | Sample, reagent, and buffer application Sealing Water pouring | Yes Proteinase K Amplification-related reagent (without primers) | IH module (no power supply) | 20 µL | Point-of-Care Toolkit for Multiplex Molecular Diagnosis of SARS-CoV-2 and Influenza A and B Viruses | D. Lee, C.-H. Chu and A. F. Sarioglu, ACS Sens., 2021, 6, 3204-3213
SARS-CoV-2 | Swab sample | Clinical sample | 4 | N.A. | FL | 102–104 copies | 102 copies | Room temperature >30 days | 20 min (off-chip) 40 min (on-chip) | Yes Single reaction (without clear kinetic data) | RNA extraction | Sample and reagent application Sealing FL measurement | Yes Amplification-related reagent | Heating plate FL meter | 1 µL | Autonomous lab-on-paper for multiplexed, CRISPR-based diagnostics of SARS-CoV-2 | K. Yin, X. Ding, Z. Li, M. M. Sfeir, E. Ballesteros and C. Liu, Lab Chip, 2021, 21, 2730-2737
SARS-CoV-2 | Swab sample | Clinical sample | 4 | N.A. | Echem (amperometry) | 0.05–100 nM (PBS) | 0.12 pM (PBS) | 4°C >4 weeks 90–110 % | <25 min (extraction) 52 min (on-chip) | Yes Multiple operations | RNA extraction | Sample application Buffer application Device assebmle Reagent application Echem meausrement | Yes Fe(CN)63-/4- | Potentiostat | 3 µL | Peptide nucleic acid probe-assisted paper-based electrochemical biosensor for multiplexed detection of respiratory viruses | A. Lomae, K. Teekayupak, P. Preechakasedkit, E. Pasomsub, T. Ozer, C. S. Henry, D. Citterio, T. Vilaivan, O. Chailapakul and N. Ruecha, Talanta, 2024, 279, 126613
SARS-CoV-2 | Swab sample | Clinical sample | 4 | 348–156,036 copies/μL | Colorimetric | N.A. | 400 copies/μL (smartphone) 2400 copies/μL (naked–eye) | N.A. | <30 min | Yes Multiple operations | RNA extraction | Sample application Folding for three times Imaging | No | 3D–printed attachment (Smartphone) | 5 µL | A paper-based assay for the colorimetric detection of SARS-CoV-2 variants at single-nucleotide resolution | T. Zhang, R. Deng, Y. Wang, C. Wu, K. Zhang, C. Wang, N. Gong, R. Ledesma-Amaro, X. Teng, C. Yang, T. Xue, Y. Zhang, Y. Hu, Q. He, W. Li and J. Li, Nat. Biomed. Eng., 2022, 6, 957-967
Human papillomavirus (HPV) | Swab sample | Clinical sample | 4 | N.A. | FL | 1–105 fM | 1 fM | N.A. | 35 min | Yes Multiple operations | RNA extraction Reverse transcription and digestion | Sample and MgOAc application Opening valve Folding and reagent addition FL imaging | Yes MgOAc FAM-ssDNA-quencher | Incubator Black box Portable fluorescence detector | 5 µL | CRISPR/Cas on Microfluidic Paper-Based Analytical Devices for Point-of-Care Screening of Cervical Cancer | S. Liu, T. Yu, L. Song, K. Kalantar-Zadeh and G. Liu, ACS Sens., 2025, 10, 4569–4579
Human papillomavirus (HPV) subtypes | Swab sample | Clinical sample | 4 | N.A. | FL | N.A. | 1–100 aM | N.A. | 40 min | Yes Multiple operations | DNA extraction Pre–amplification | Sample application Sealing FL imaging | Yes Amplification-related reagent | Heater FL meter | 7.5 µL | Programmable Multiplexed Nucleic Acid Detection by Harnessing Specificity Defect of CRISPR-Cas12a | X. Guan, R. Yang, J. Zhang, J. Moon, C. Hou, C. Guo, L. Avery, D. Scarola, D. S. Roberts, R. LaSala and C. Liu, Adv. Sci., 2025, 12, 2411021
Influenza (H1N1) | Swab sample | Clinical sample | 4 | N.A. | Echem (amperometry) | 0.05–1000 nM (PBS) | 0.35 pM (PBS) | 4°C >4 weeks 90–110 % | <25 min (extraction) 52 min (on–chip) | Yes Multiple operations | RNA extraction | Sample application Buffer application Device assembly Reagent application Echem meausrement | Yes Fe(CN)63-/4- | Potentiostat | 3 µL | Peptide nucleic acid probe-assisted paper-based electrochemical biosensor for multiplexed detection of respiratory viruses | A. Lomae, K. Teekayupak, P. Preechakasedkit, E. Pasomsub, T. Ozer, C. S. Henry, D. Citterio, T. Vilaivan, O. Chailapakul and N. Ruecha, Talanta, 2024, 279, 126613
Respiratory syncytial virus (RSV) | Swab sample | Clinical sample | 4 | N.A. | Echem (amperometry) | 10–1000 nM (PBS) | 0.36 pM (PBS) | 4°C >4 weeks 90–110 % | <25 min (extraction) 52 min (on–chip) | Yes Multiple operations | RNA extraction | Sample application Buffer application Device assembly Reagent application Echem meausrement | Yes Fe(CN)63-/4- | Potentiostat | 3 µL | Peptide nucleic acid probe-assisted paper-based electrochemical biosensor for multiplexed detection of respiratory viruses | A. Lomae, K. Teekayupak, P. Preechakasedkit, E. Pasomsub, T. Ozer, C. S. Henry, D. Citterio, T. Vilaivan, O. Chailapakul and N. Ruecha, Talanta, 2024, 279, 126613
Hepatitis B virus | Human serum | Clinical sample | 4 | 5.27–5.39 log copies/µL | Echem (SWV) | 10 pM to 2 μM (PBS) | 7.23 pM (PBS) | 4°C <1 month 110% | 7 min | Yes Single reaction Automated sequential reagent addition | (Blood separation) DNA extraction | Sample application Echem measurement | No | Potentiostat | 200 µL (off chip) 50 µL (on chip) | Amplification-free DNA Sensor for the One-Step Detection of the Hepatitis B Virus Using an Automated Paper-Based Lateral Flow Electrochemical Device | C. Srisomwat, A. Yakoh, N. Chuaypen, P. Tangkijvanich, T. Vilaivan and O. Chailapakul, Anal. Chem., 2021, 93, 2879-2887
Dengue virus (DEN) | Human serum | Clinical sample | 4 | N.A. | Colorimetric | N.A. | 5 copies (serum) | <30°C >30 hr | ~30 min | Yes Multiple operations | Blood separation | Pre-mixning sample and lysis buffer Samle application Washing for three times Membrane placement Buffer application Reagent application Sealing | Yes Lysis buffer Ethanol Amplification-related reagent | Waste container Portable heater | ~50 µL | All-Serotype Dengue Virus Detection through Multilayered Origami-Based Paper/Polymer Microfluidics | P. Biswas, G. N. Mukunthan Sulochana, T. N. Banuprasad, P. Goyal, D. Modak, A. K. Ghosh and S. Chakraborty, ACS Sens., 2022, 7, 3720-3729
Dengue virus | Human serum | Clinical sample | 4 | N.A. | FL | N.A. | 10 copies (serum) | N.A. | <1 hr | Yes Multiple operations | (Blood separation) DNA extraction on different device | Sample and buffer application Reagent application Sample collection Mix with buffer Sample transfer Sample applicatoin FL measurement | Yes NaOH | Humid petri dish Heater Fluorescence meter | 20 µL | Paper-based nucleic acid testing system for simple and early diagnosis of mosquito-borne RNA viruses from human serum | B. S. Batule, Y. Seok and M.-G. Kim, Biosens. Bioelectron., 2020, 151, 111998
Zika virus | Human serum | Clinical sample | 4 | N.A. | FL | N.A. | 10 copies (serum) | N.A. | <1 hr | Yes Multiple operations | (Blood separation) DNA extraction on different device | Sample and buffer application Reagent application Sample collection Mix with buffer Sample transfer Sample applicatoin FL measurement | Yes NaOH | Humid petri dish Heater Fluorescence meter | 20 µL | Paper-based nucleic acid testing system for simple and early diagnosis of mosquito-borne RNA viruses from human serum | B. S. Batule, Y. Seok and M.-G. Kim, Biosens. Bioelectron., 2020, 151, 111998
Let-7a (miRNA) | Human serum | Diluted clinical sample | 3 | N.A. | BL | 2 fM to 50 pM | 1.7 fM | Silica gel desiccation Room temperature >3 months | 15 min (off-chip) 150 min (on-chip) | Yes Multiple operations | RNA extraction (Blood separation) 2.5 times dilution with water | Sample application Overlapping paper disc Reagent application Imaging | Yes Luciferin | Incubator 3D–printed dark box Smartphone | 20 µL | Portable and Field-Ready Detection of Circulating MicroRNAs with Paper-Based Bioluminescent Sensing and Isothermal Amplification | Y. Li, L. Zhou, W. Ni, Q. Luo, C. Zhu and Y. Wu, Anal. Chem., 2019, 91, 14838-14841
SARS-CoV-2 | Human saliva | Spiked into biofluid sample | 2 | 6 µM | BL | N.A. | 40 nM (30 % (v/v) saliva) | Sealed bags Room temperature <7 weeks | 7–12 min | Yes Single reaction | (RNA extraction) | Sample application Sealing BL imaging | No | 3D printed cassette Heater Camera | 15 µL | Towards detection of SARS-CoV-2 RNA in human saliva: A paper-based cell-free toehold switch biosensor with a visual bioluminescent output | J. P. Hunt, E. L. Zhao, T. J. Free, M. Soltani, C. A. Warr, A. B. Benedict, M. K. Takahashi, J. S. Griffitts, W. G. Pitt and B. C. Bundy, N. Biotechnol., 2022, 66, 53-60
Influenza A viruses | Saliva | Spiked into biofluid sample | 2 | 0.5-5000 copies/µL | Colorimetric | N.A. | 50 copies/μL (saliva) | N.A. | 60 min | Yes Single reaction Automated sequential reagent addition | No | Sample, reagent, and buffer application Sealing Water puring | Yes Proteinase K Amplification-related reagent (without primers) | IH module (no power supply) | 20 µL | Point-of-Care Toolkit for Multiplex Molecular Diagnosis of SARS-CoV-2 and Influenza A and B Viruses | D. Lee, C.-H. Chu and A. F. Sarioglu, ACS Sens., 2021, 6, 3204-3213
Influenza B viruses | Saliva | Spiked into biofluid sample | 2 | 0.5-5000 copies/µL | Colorimetric | N.A. | 50 copies/μL (saliva) | N.A. | 60 min | Yes Single reaction Automated sequential reagent addition | No | Sample, reagent, and buffer application Sealing Water puring | Yes Proteinase K Amplification-related reagent (without primers) | IH module (no power supply) | 20 µL | Point-of-Care Toolkit for Multiplex Molecular Diagnosis of SARS-CoV-2 and Influenza A and B Viruses | D. Lee, C.-H. Chu and A. F. Sarioglu, ACS Sens., 2021, 6, 3204-3213
HIV-1 | Human whole blood | Spiked into biofluid sample | 2 | 3–3x105 copies/mL | FL | N.A. | 3 × 103 copies/mL (whole blood) | N.A. | 30 min (extraction) >10 min (post amplification) | Yes Multiple operations | Post DNA isothermal amplification | Sample application Filter membrane removal Sealing Buffer application Cutting membrane Transfer to RPA tube FL measurement | Yes Amplification-related reagent | Cissors Microtube Heater FL meter | 33 µL | Nucleic acid sample preparation from whole blood in a paper microfluidic device using isotachophoresis | B. P. Sullivan, A. T. Bender, D. N. Ngyuen, J. Y. Zhang and J. D. Posner, J. Chromatogr. B, 2021, 1163, 122494
MRSA | Blood | Spiked into biofluid sample | 2 | 50 copies | FL | 1–104 copies (35 min) 105–108 copies (42 min) | 1 copy (42 min) | Room temperature >35 days 81.25% | 36 or 43 min | Yes Single reaction (without clear kinetic data) | DNA extraction | Sample and water application Sealing Real–time FL measurement | No | Heater FL meter | 1 µL | Fluorometric Paper-Based, Loop-Mediated Isothermal Amplification Devices for Quantitative Point-of-Care Detection of Methicillin-Resistant Staphylococcus aureus (MRSA) | I. Choopara, A. Suea-Ngam, Y. Teethaisong, P. D. Howes, M. Schmelcher, A. Leelahavanichkul, S. Thunyaharn, D. Wongsawaeng, A. J. Demello, D. Dean and N. Somboonna, ACS Sens., 2021, 6, 742-751
Chikungunya | Human serum | Spiked into biofluid sample | 2 | N.A. | FL | N.A. | 10 copies (serum) | N.A. | <1 hr | Yes Multiple operations | (Blood separation) DNA extraction on different device | Sample and buffer application Reagent application Sample collection Mix with buffer Sample transfer Sample applicatoin FL measurement | Yes NaOH | Humid petri dish Heater FL meter | 20 µL | Paper-based nucleic acid testing system for simple and early diagnosis of mosquito-borne RNA viruses from human serum | B. S. Batule, Y. Seok and M.-G. Kim, Biosens. Bioelectron., 2020, 151, 111998
SARS-CoV-2 | N.A. (aerosol generated by breathing simulator) | Spiked into artificial biofluid sample | 1 | 50–5×106 copies | Colorimetric | N.A. | 17 aM | N.A. | 30 min (sample collection) ~1.5 h (on chip) | No Automated sequential reaction | No | Sample collection Button push | No | No | No (mask) | Wearable materials with embedded synthetic biology sensors for biomolecule detection | P. Q. Nguyen, L. R. Soenksen, N. M. Donghia, N. M. Angenent-Mari, H. De Puig, A. Huang, R. Lee, S. Slomovic, T. Galbersanini, G. Lansberry, H. M. Sallum, E. M. Zhao, J. B. Niemi and J. J. Collins, Nat. Biotechnol., 2021, 39, 1366-1374
Pathogens DNA (E. coli, S. aureus, P. aeruginosa, K. pneumoniae, A. baumannii) | Swab (mimicking bacterial contamination on bench) | Spiked into artificial biofluid sample | 1 | 103–106 CFU/mL | FL | 1–106 copies/ μL (PBS) | 1 copy/μL (PBS) | N.A. | <60 min | Yes Multiple operations | No | Sample and buffer application Putting on a heater Putting on a detecter FL imaging | No | Heater Plastic case Digital fluorescence detection module | 20 µL | An integrated and paper-based microfluidic system employing LAMP-CRISPR and equipped with a portable device for simultaneous detection of pathogens | S. Cui, K. Wang, Y. Yang, X. Lv and X. Li, Anal. Bioanal. Chem., 2024, 417, 785–797
HIV-1 | Human whole blood 6.25-fold diluted with buffer | Spiked into diluted biofluid sample | 1 | 3 × 105 copies | Colorimetric | N.A. | 3 × 105 copies (whole blood) | In bags with silica gel desiccant Room temperature >3 weeks | 90 min | No Custom-designed electronic device controlled | (RNA extraction) | Sample application Device assemble Buffer application Heating | No | Portable heater and circuit | 75 µL | Microfluidic rapid and autonomous analytical device (microRAAD) to detect HIV from whole blood samples | E. A. Phillips, T. J. Moehling, K. F. K. Ejendal, O. S. Hoilett, K. M. Byers, L. A. Basing, L. A. Jankowski, J. B. Bennett, L.-K. Lin, L. A. Stanciu and J. C. Linnes, Lab Chip, 2019, 19, 3375-3386
Mpox virus | Artificial urine or saliva (25-fold diluted) | Spiked into diluted biofluid sample | 1 | 1–100 pM | Colorimetric | 10–1000 copies/mL (urine, saliva) | ~5 copies/mL (urine, saliva) | 55°C >30 days >95% | <30 min | Yes Multiple operations | DNA extraction Anealing Pre-incubation | Pre-mixing sample and reagent Pre-mixing two reagents separately Pre-mix above two solutions Sample application Imaging | Yes AgNO3 Urease | Incubator Microtubes Colorimeter | 200 µL | Intended Mismatch DNA Nanocomplexes for Ultrasensitive, Enzymatic Colorimetric Detection of MPOX Virus in Minutes Using a Paper-Based Fast-Flow Micromixing Device | L. Hongtanee, S. Chaiyo, H. Wu and A. Yakoh, Adv. Healthc. Mater., 2025, 14, 2404320
HPV16 | N.A. | N.A. | - | N.A. | Echem (SWV) | N.A. | 1 copy/µL | N.A. | 130 min | Yes Multiple operations | (DNA extraction) Fragmentation Pre-RPA amplification Pre-CRISPR reaction | Buffer application Pre-mixing sample and reagent Sample, buffer and reagent application Echem emasurement | Yes (no reagent storage) | Potentiostat | 10 µL | Paper-Based Laser-Pyrolyzed Electrofluidics: An Electrochemical Platform for Capillary-Driven Diagnostic Bioassays | L. Bezinge, J. M. Lesinski, A. Suea-Ngam, D. A. Richards, A. J. deMello and C.-J. Shih, Adv. Mater., 2023, 35, 2302893
Human papilloma virus (HPV) | Buffer | N.A. | - | N.A. | Echem (personal glucose meter) | N.A. | 57 pM | Aluminium–coated pouches with a silica gel -20°C >4 weeks | 120 min | Yes Multiple operations | (DNA extraction) | Sample application Film removal Buffer application Echem measurement | No | Personal glucose meter | 10 µL | Paper-based analytical device for point-of-care nucleic acid quantification combining CRISPR/Cas12a and a personal glucose meter | Y. Tanifuji, G. Tong, Y. Hiruta and D. Citterio, Analyst, 2024, 149, 4932-4939
Ebola virus | Water | N.A. | - | N.A. | Colored length | N.A. | 11 aM (water) | N.A. | 285 min | Yes Multiple operations | Preincubation of CRISPR reaction | Pre–incubation Dilution Sample application Resistance measurement | Yes Amplification-related reagent CRISPR-related reagent | Incubator Digital Multimeter | 5 µL | Programmable CRISPR-responsive smart materials | M. A. English, L. R. Soenksen, R. V. Gayet, H. de Puig, N. M. Angenent-Mari, A. S. Mao, P. Q. Nguyen and J. J. Collins, Science, 2019, 365, 780-78

## Table S3 (Other targets)
Table S3 Recent examples of µPADs (2018-2025) for medical diagnostic applications (other targets)
This is an extended version of Table 3 in the main text
Columns not shown in the main text tables are indicated with light blue colour column headings
Analyte | Sample matrix | Real sample test | Sample type category | Concentration of measured clinical or spiked sample | Detection technique | Detection range | Detection limit | Storage stability | Assay time | Necessity of time control | Off-device pretreatment | Operation steps | Reagent handling | External equipment | Sample volume control | Mechanism | Article title | Ref
Uric Acid | Saliva | Clinical sample | 4 | 10 – 80 ppm | Colorimetric | 0–214.8 ppm (artificial saliva) | 0.1 ppm (artificial saliva) | Air–tight sacks 4°C >3 months | <10 min | Yes Multiple operations | Pretreatment (surfactant additon, vortex and heating) | Sample application IPA application Water application Imaging | No | Smartphone Black box | 10 µL | Ox/Red | Digital Quantification Method for Sensitive Point-of-Care Detection of Salivary Uric Acid Using Smartphone-Assisted μPADs | K. Fan, J. Zeng, C. Yang, G. Wang, K. Lian, X. Zhou, Y. Deng and G. Liu, ACS Sens., 2022, 7, 2049-2057
Hemoglobin | Human saliva | Clinical sample | 5 | 0.01–107.40 μg/mL | Colorimetric intensity | 1.57–2000 µg/mL | 1.57 μg/mL | Sealed Room temperature >3 weeks | 3 min | Yes Single reaction | No | Sample application Imaging | No | Mobile phone | 38 µL | Enzyme | A Machine Learning-Driven Cyclic Optimizing Strategy for the Construction of Paper-Based Microfluidic Devices in the Early Diagnosis of Periodontitis | K. Lv, Y. Zhang, K. Tang, W. Huang, F. Chen, M. Chen, Y. Wang and J. Zhang, ACS Sens., 2025, 10, 7002–7013
Hemoglobin | Human saliva | Clinical sample | 5 | 0.01–107.40 μg/mL | Colored distance | 3–200, 200–2000 µg/mL | 3 μg/mL | Sealed Room temperature >3 weeks | 3 min | Yes Single reaction | No | Sample application Imaging | No | Mobile phone | 34 µL | Enzyme | A Machine Learning-Driven Cyclic Optimizing Strategy for the Construction of Paper-Based Microfluidic Devices in the Early Diagnosis of Periodontitis | K. Lv, Y. Zhang, K. Tang, W. Huang, F. Chen, M. Chen, Y. Wang and J. Zhang, ACS Sens., 2025, 10, 7002–7013
Albumin | Human urine | Clinical sample | 5 | 3.4–1025.4, 4246 mg/L | Colored distance | 0–1000 mg/L (Artificial urine) | N.A. | In a desiccator Room temperature >2 months | 15 min | No 15–45 min | No | Dipping | No | No | >3 mL (Dip) | Chelation | Paper-Based Device for Naked Eye Urinary Albumin/Creatinine Ratio Evaluation | R. Hiraoka, K. Kuwahara, Y.-C. Wen, T.-H. Yen, Y. Hiruta, C.-M. Cheng and D. Citterio, ACS Sens., 2020, 5, 1110-1118
Creatinine | Human urine | Clinical sample | 5 | 239–2609 mg/L | Colored distance | 0–3000 mg/L (Artificial urine) | N.A. | In a desiccator Room temperature >2 months | 15 min | No 15–45 min | No | Dipping | No | No | >3 mL (Dip) | Chelation | Paper-Based Device for Naked Eye Urinary Albumin/Creatinine Ratio Evaluation | R. Hiraoka, K. Kuwahara, Y.-C. Wen, T.-H. Yen, Y. Hiruta, C.-M. Cheng and D. Citterio, ACS Sens., 2020, 5, 1110-1118
Glucose | Human whole blood | Clinical sample | 5 | 5.4–12.4 mM | Colorimetric | 2.5–50 mM (whole blood) | 0.3 mM (whole blood) | N.A. | 10 min | Yes Single reaction (without kinetic data) | No | Sample application Imaging | No | Smartphone | 100 µL | Enzyme | Three-Dimensional Paper-Based Microfluidic Analytical Devices Integrated with a Plasma Separation Membrane for the Detection of Biomarkers in Whole Blood | C. Park, H.-R. Kim, S.-K. Kim, I.-K. Jeong, J.-C. Pyun and S. Park, ACS Appl. Mater. Interfaces., 2019, 11, 36428-36434
Glucose | Human whole blood | Clinical sample | 5 | 63–146 mg/dL | Colorimetric | 50–365 mg/dL (PBS) | N.A. | N.A. | 30 min | Yes Multiple operations | No | Sample and buffer application Imaging | No | Scanner | 1 µL | Enzyme | An innovative blood plasma separation method for a paper-based analytical device using chitosan functionalization | D. Kim, S. Kim and S. Kim, Analyst, 2020, 145, 5491-5499
Glucose | Human whole blood | Clinical sample | 5 | 5.1–6.4 mM | Colorimetric | 2.20–15.00 mM (whole blood) | 0.12 mM (whole blood) | Sealed with a plastic film -4°C <40 days 89% | 12 min | Yes Single reaction | No | Sample application Imaging | No | Home–made readout device (smartphone) | 20 µL | Enzyme | Hydrogel Paper-Based Analytical Devices: Separation-Free In Situ Assay of Small-Molecule Targets in Whole Blood | J. Xu, H. Khan and L. Yang, Anal. Chem., 2021, 93, 14755-14763
Urea | Human whole blood | Clinical sample | 5 | 2.46–38.14 mM | Colorimetric | 2.46–38.14 mM (blood) | 0.03 mM (blood) | 35°C >8 days 96% | <2 min | Yes Single reaction | No | Sample application | No | Plastic cartridge Portable optizal reader | 20 µL | Enzyme | A low-cost paper-based blood urea nitrogen optical biosensor for renal surveillance in fingertip blood | J. Cheng, Y. Fu, J. Guo and J. Guo, Sens. Actuators B: Chem., 2023, 387, 133795
Uric acid | Human whole blood | Clinical sample | 5 | 0.1198–0.6401 mM | Colorimetric | 0–1.5 mM (whole blood) | 0.1127 mmol/L (whole blood) | Sealing and light avoidance 4°C <28 days | >10 min | No 10–24 min | No | Sample application Buffer application Imaging | No | Smartphone Box with light source | 30 µL | Enzyme | Microfluidic Chip with Integrated Whole Blood Separation for Simultaneous Accurate Quantification of Multiple Chronic Kidney Disease Markers and Long-Term Monitoring | J. Yang, X. Feng, H. Li, P. Chang, X. Ren and X. Li, ACS Sens., 2025, 10, 5623–5632
Creatinine | Human whole blood | Clinical sample | 5 | 58.8–1092.4 µM | Colorimetric | 0–100 μM (whole blood) | 0.2978 μmol/L (whole blood) | Sealing and light avoidance 4°C <28 days | >13 min | No 13–24 min | No | Sample application Buffer application Imaging | No | Smartphone Box with light source | 30 µL | Dye | Microfluidic Chip with Integrated Whole Blood Separation for Simultaneous Accurate Quantification of Multiple Chronic Kidney Disease Markers and Long-Term Monitoring | J. Yang, X. Feng, H. Li, P. Chang, X. Ren and X. Li, ACS Sens., 2025, 10, 5623–5632
Albumin | Human whole blood | Clinical sample | 5 | 33–50.9 mg/mL | Colorimetric | 0–100 mg/mL (whole blood) | 0.7696 mg/mL (whole blood) | Sealing and light avoidance 4°C <28 days | >6 min | No 6–24 min | No | Sample application Buffer application Imaging | No | Smartphone Box with light source | 30 µL | Dye | Microfluidic Chip with Integrated Whole Blood Separation for Simultaneous Accurate Quantification of Multiple Chronic Kidney Disease Markers and Long-Term Monitoring | J. Yang, X. Feng, H. Li, P. Chang, X. Ren and X. Li, ACS Sens., 2025, 10, 5623–5632
Hemoglobin | Human whole blood | Clinical sample | 5 | N.A. | Colorimetric | 2.5–15 g/dL | 2.5 g/dL | Room temperature >8 hrs | <5 min | No Automated smartphone app | No | Sample application Auto–Imaging with smartphone | No | Plastic box Smartphone | around 10 µL (finger–pricked) | Enzyme | Smartphone-Enabled Paper-Based Hemoglobin Sensor for Extreme Point-of-Care Diagnostics | S. K. Biswas, S. Chatterjee, S. Bandyopadhyay, S. Kar, N. K. Som, S. Saha and S. Chakraborty, ACS Sens., 2021, 6, 1077-1085
Cardiac troponin I | Human whole blood | Clinical sample | 4 | <0.5 ~ 0.83 ng/mL | Colored distance | 0.025–2.5 ng/mL (whole blood) | 0.025 ng/mL (whole blood) | Sealed brown bag Room temperature >10 weeks >95% | 15–30 min | No 15–30 min | Sample sterilization | Sample and buffer application | No | No | 3 µL | Dye | Rapid Distance-Based Cardiac Troponin Quantification Using Paper Analytical Devices for the Screening and the Follow-Up of Acute Myocardial Infarction, Using a Single Drop of Human Whole Blood | K. Khachornsakkul and W. Dungchai, ACS Sens., 2021, 6, 1339-1347
Urea | Human saliva | Diluted clinical sample | 3 | <0.212–2.02 mM | Colorimetric | 0.10–5.0 mM | 0.032 mM | Under vacuum Room temperature <1 months 96% | 35 min | No 35–120 min | 10-fold dilution with synthetic saliva | Sample application Sealing Imaging | No | Scanner | 15–25 µL | Enzyme | On-hand tool for ammonium and urea determination in saliva to monitor chronic kidney disease – Design of a couple of microfluidic paper-based devices | F. T. S. M. Ferreira, R. B. R. Mesquita and A. O. S. S. Rangel, Microchem. J., 2023, 193, 109102
NHx | Human saliva | Diluted clinical sample | 3 | <0.384–3.01 mM | Colorimetric | 0.16–5.0 mM | 0.049 mM | Under vacuum Room temperature >2 months 105% | 15 min | No 15–60 min | 10-fold dilution with synthetic saliva | Sample application NaOH application Sealing Imaging | Yes NaOH | Scanner | 15–25 µL | pH change | On-hand tool for ammonium and urea determination in saliva to monitor chronic kidney disease – Design of a couple of microfluidic paper-based devices | F. T. S. M. Ferreira, R. B. R. Mesquita and A. O. S. S. Rangel, Microchem. J., 2023, 193, 109102
Zinc(II) | Human urine | Spiked into biofluid sample | 2 | 0.115–1.419 mg/L | Colorimetric | 50.0–750 μg/L (urine) | 10 μg /L (urine) | Room temperature <7 days >90% | >5 min | No 5 min–2 hrs | 2-fold dilution | Sample application Imaging | No | Scanner | 15 µL | Dye | A microfluidic paper-based analytical device for the determination of zinc(II) in children’s urine samples and food zinc supplement solutions | M. M. P. Melo, A. O. S. S. Rangel and R. B. R. Mesquita, Talanta, 2026, 298, 128865
Histidine | Human urine | Spiked into biofluid sample | 2 | 1–5 mM | FL–pixel | 0.5–20 mM | 0.021 mM | N.A. | >10 min (drying time not included) | Yes Multiple operations | No | Sample application Covering Peeling off Reagent application FL imaging | Yes BSA-AuCNs-Cu2+ | Oven or hair dryer Imaging box Excitation light Smartphone | 7 µL | Coffee–ring | Dual-Mode Undistorted Visual Fluorescent Sensing Strategy through Manipulating the Coffee-Ring Effect on Microfluidic Paper-Based Chip | J. Qi, B. Li, D. Lin, Z. Du, L. Fu, X. Wang, Z. Zhang, L. Luo and L. Chen, Anal. Chem., 2023, 95, 10486-10491
Histidine | Human urine | Spiked into biofluid sample | 2 | 1–5 mM | FL–distance | 0.5–20 mM | 0.5 mM | N.A. | >10 min (drying time not included) | Yes Multiple operations | No | Sample application Covering Peeling off Reagent application FL imaging | Yes BSA-AuCNs-Cu2+ | Oven or hair dryer Imaging box Excitation light Smartphone | 7 µL | Coffee–ring | Dual-Mode Undistorted Visual Fluorescent Sensing Strategy through Manipulating the Coffee-Ring Effect on Microfluidic Paper-Based Chip | J. Qi, B. Li, D. Lin, Z. Du, L. Fu, X. Wang, Z. Zhang, L. Luo and L. Chen, Anal. Chem., 2023, 95, 10486-10491
Glucose | Human serum | Spiked into biofluid sample | 2 | 5–11 mM | Colorimetric | 5–11 mM (serum) | 0.3 mM (PBS) | N.A. | 5 min (buffer) 8 min (serum) | Yes Single reaction (without kinetic data) | (Blood separation) | Sample application Imaging | No | Camera | 500 µL | Enzyme | Double-sided 3D printing on paper towards mass production of three-dimensional paper-based microfluidic analytical devices (3D-μPADs) | C. Park, Y. D. Han, H. V. Kim, J. Lee, H. C. Yoon and S. Park, Lab Chip, 2018, 18, 1533-1538
Cholesterol | Human serum | Spiked into biofluid sample | 2 | 2.6–6.7 mM | Colorimetric | 2.6–6.7 mM (serum) | 0.2 mM (PBS) | N.A. | 5 min (buffer) 8 min (serum) | Yes Single reaction (without kinetic data) | (Blood separation) | Sample application Imaging | No | Camera | 500 µL | Enzyme | Double-sided 3D printing on paper towards mass production of three-dimensional paper-based microfluidic analytical devices (3D-μPADs) | C. Park, Y. D. Han, H. V. Kim, J. Lee, H. C. Yoon and S. Park, Lab Chip, 2018, 18, 1533-1538
Triglyceride | Human serum | Spiked into biofluid sample | 2 | 1–2.3 mM | Colorimetric | 1–2.3 mM (serum) | 0.3 mM (PBS) | N.A. | 5 min (buffer) 8 min (serum) | Yes Single reaction (without kinetic data) | (Blood separation) | Sample application Imaging | No | Camera | 500 µL | Enzyme | Double-sided 3D printing on paper towards mass production of three-dimensional paper-based microfluidic analytical devices (3D-μPADs) | C. Park, Y. D. Han, H. V. Kim, J. Lee, H. C. Yoon and S. Park, Lab Chip, 2018, 18, 1533-1538
Secnidazole | Serum | Spiked into biofluid sample | 2 | 0.1–300 μM | FL | 0.1–300 μM | N.A. (117 nM in bulk) | N.A. | N.A. | N.A. | (Blood separation) | Sample application Drying FL measurement | No | Smartphone Excitation light | 50 µL | FL quenching | Microfluidic Paper-Based Fluorescent Sensing Device Using Citric Acid Derived Carbon Dots-Tungsten Sulfide Spheres for On-site Detection of Secnidazole | S. Arif, M. Hassan, S. Rasheed, N. Ahmad, B. Fatima, S. G. Musharraf and D. Hussain, Biosens. Bioelectron., 2026, 291, 118025
Alkaline phosphatase (ALP) | Calf serum Human serum Goat serum | Spiked into biofluid sample | 2 | 35, 105, 8 U/L | Echem (CA) | 1–1000 U/L | 1.4 U/L (buffer) | With desiccant Room temperature >4 weeks >70% | 20 min (pre–incubation) 20 min (24 samples flow measurement) | Yes Single reaction | Centrifugation Pre–incubation | Pre-mixing sample and reagent Sample application Echem measurement | Yes (no reagent storage) | Potentiostat | 10 µL | Enzyme | Paper-Based Laser-Pyrolyzed Electrofluidics: An Electrochemical Platform for Capillary-Driven Diagnostic Bioassays | L. Bezinge, J. M. Lesinski, A. Suea-Ngam, D. A. Richards, A. J. deMello and C.-J. Shih, Adv. Mater., 2023, 35, 2302893
H2O2 | Artificial saliva | Spiked into artificial biofluid sample | 1 | 5, 10 μM | Echem (CA) | 0.25 μM to 5 mM (PBS) | 0.15 μM (PBS) | >4 weeks 82.9% | 110 s | No Immediate measurement | No | Sample application Echem measurement | No | Potentiostat | ~400 μL | Enzyme | Development of an origami microfluidic paper-based analytical device for highly sensitive detection of dual biomarkers in artificial saliva | C. Chen, B. Ran, B. Liu, Y. Xiao, G. Tang, T. Shen, J. Liang and Z. Liang, Microchem. J., 2025, 210, 112966
Glucose | Artificial saliva | Spiked into artificial biofluid sample | 1 | 100, 200 μM | Echem (CA) | 2.5 μM to 3 mM | 1.5 μM | In PBS 4°C <4 weeks 78.4% | 110 s (only Echem measurement time) | No Immediate measurement | No | Sample application Echem measurement | No | Potentiostat | ~400 μL | Enzyme | Development of an origami microfluidic paper-based analytical device for highly sensitive detection of dual biomarkers in artificial saliva | C. Chen, B. Ran, B. Liu, Y. Xiao, G. Tang, T. Shen, J. Liang and Z. Liang, Microchem. J., 2025, 210, 112966
Glucose | Artificial saliva | Spiked into artificial biofluid sample | 1 | 15–200 mg/dL | Coloried distance (barcode) | 15–200 mg/dL (artificial saliva) | 15 mg/dL (artificial saliva) | N.A. | 90 min | Yes Multiple operations | No | Sample application Drying Smartphone readout | No | Smartphone | 80 µL | Enzyme | Development of a Semiquantitative Barcode Readout Approach for Paper-Based Analytical Devices (PADs) for Enzymatic H2O2 and Glucose Detection | Y. Manmana, S. Kinugasa, Y. Hiruta and D. Citterio, Anal. Chem., 2025, 97, 1500-1506
Glucose | Artificial urine | Spiked into artificial biofluid sample | 1 | 1–100 mM | Colorimetric | 0–20 mM (aritificial urine) | 1 mM (aritificial urine) | N.A. | 10–30 min | No No color degredation | No | Sample application | No | No | 40 µL or 2 drops | Polymer degration | Traffic light type paper-based analytical device for intuitive and semi-quantitative naked-eye signal readout | S. Ohta, R. Hiraoka, Y. Hiruta and D. Citterio, Lab Chip, 2022, 22, 717-726
Ascorbic acid | Artificial urine | Spiked into artificial biofluid sample | 1 | 0.1135 mM | Echem (CA) | 0.15–0.8 mM (PBS, carriar buffer) | 92.8 μM (PBS, carriar buffer) | N.A. | 250 s | No Immediate measurement | No | Sample application Echem measurement | No | Potentiostat | 1 µL | Ox/Red | 3D Capillary-Driven Paper-Based Sequential Microfluidic Device for Electrochemical Sensing Applications | A. Yakoh, S. Chaiyo, W. Siangproh and O. Chailapakul, ACS Sens., 2019, 4, 1211-1221
Lithium ion | Human whole Blood (1.25-fold diluted) | Spiked into diluted biofluid sample | 2 | 0–2.0 mM | Colorimetric | 0–2.0 mM* (whole blood) | 0.281 mM (whole blood) | N.A. | 10 s (does not include blood wicking time) | No Immediate measurement | No | Sample application Placement of detection unit Imaging | No | (Camera) | 10–20 µL | Dye | Paper-Based Device for the Facile Colorimetric Determination of Lithium Ions in Human Whole Blood | T. Komatsu, M. Maeki, A. Ishida, H. Tani and M. Tokeshi, ACS Sens., 2020, 5, 1287-1294
DNA adenine methylation Mtase (Dam MTase) | Human serum (5-fold diluted) | Spiked into diluted biofluid sample | 2 | 30–70 U/mL | FL | 1.0–90 U/mL (buffer) | 0.11 U/mL (buffer) | >10 days | 35 min | Yes Multiple operations | (Blood separation) | Pre-mixing of sample and reagent Sample application Reagent addition FL measurement | Yes Fluorescent dye | FL meter | ~7 µL (no details) | Enzyme | Paper-based fluorescence chip with umbellately anchored dimeric G-quadruplex coupled DNA tetrahedron-structured probe for sensing DNA methyltransferase | M. Wang, J. Qi, F. Li, Y. Liu, J. Xiang, B. Li, X. Zhuang, L. Chen and X. Fu, Sens. Actuators B: Chem., 2026, 446, 138643
Nicotinamide adenine dinucleotide phosphate (NADPH) | Simulated urie Human serum (10-fold diluted) | Spiked into diluted biofluid sample | 2 | 0.01−1.0 mM | Echem (CA) | 0.01–1.0 mM, 1.0–4.0 mM (urine and serum) | 3.75 μM (buffer) | Sealed in aluminum bags 55°C >7 days (ambient condition for 56 days) 99.8% | <10 min (5 samples with flow measurement) | No Immediate measurement | No | Buffer application Sample application Echem measurement | No | Potentiostat | 10 µL | Ox/Red | Paper-Based Fast-Flow Electrochemical Device Enabling Self-Calibrated and High-Throughput NADPH Measurements | N. Thueankhum, L. Hongtanee, U. Boonyuen, A. Yakoh and K. Charoenkitamorn, Anal. Chem., 2025, 97, 20771–20779
L-Tryptophan (Trp) | Human serum (depotenized) Human urine (centrifuged and 10-fold diluted) | Spiked into diluted biofluid sample | 2 | 0.075–0.125 mM | Colorimetric | 0.25–1.5 mM | 220.7 μM | N.A. | 15 min | Yes Multiple operations | (Blood separation) | Sample application Centrifugation twice Reagent application Imaging | Yes ZIF-8-His-MIP-Gel Au-Nos-D,L-Ala | Centrifuge Smartphone | N.A. | AuNP aggregation | Paper-based microfluidic system and chiroptical functionalized gold nano-oval for colorimetric detection of L-Tryptophan | M. Karimian, K. Dashtian, R. Zare-Dorabei and S. Norouzi, Anal. Cheim. Acta, 2024, 1285, 342022
Albumin | Human serum (2-fold diluted) | Spiked into diluted biofluid sample | 2 | 1.5–18.0 g/L | Colored distance, angle | 1–25 g/L (PBS) | 0.8 g/L (PBS) | N.A. | 30 min | No 30–40 min | (Blood separation) | Sample and buffer application | No | No | 5 µL | Dye enzyme | Distance-Based Paper Device for a Naked-Eye Albumin-to-Alkaline Phosphatase Ratio Assay | K. Phoonsawat, K. Khachornsakkul, N. Ratnarathorn, C. S. Henry and W. Dungchai, ACS Sens., 2021, 6, 3047-3055
Alkaline phosphatase (ALP) | Human serum (2-fold diluted) | Spiked into diluted biofluid sample | 2 | 12.5–140.0 U/L | Colored distance, angle | 5–50 U/L and 50–200 U/L (PBS) | 5 U/L (PBS) | N.A. | 30 min | No 30–40 min | (Blood separation) | Sample and buffer application | No | No | 5 µL | Dye enzyme | Distance-Based Paper Device for a Naked-Eye Albumin-to-Alkaline Phosphatase Ratio Assay | K. Phoonsawat, K. Khachornsakkul, N. Ratnarathorn, C. S. Henry and W. Dungchai, ACS Sens., 2021, 6, 3047-3055
Serotonin (5-HT) | Buffer | N.A. | - | N.A. | Echem (CA) | 1–20 mM (PBS) | 0.15 mM (PBS) | N.A. | 13 min | Yes Multiple operations | N.A. | Reagent application Washing Sample application Moving PAD Echem measurement | Yes K[AuCl4] | Potentiostat | 1 µL | Ox/Red | 3D Capillary-Driven Paper-Based Sequential Microfluidic Device for Electrochemical Sensing Applications | A. Yakoh, S. Chaiyo, W. Siangproh and O. Chailapakul, ACS Sens., 2019, 4, 1211-1221
Cardiac troponin I (cTnI) | PBS | N.A. | - | N.A. | SERS | 0–0.5 ng/ mL (PBS) | 0.02 ng/mL (PBS) | N.A. | <90min | No Automated sequential reagent addition | N.A. | Sample and buffer application SERS measurement | No | SERS spectrophotometer | 75 µL | Aptamer | Paper Microfluidic Device with a Horizontal Motion Valve and a Localized Delay for Automatic Control of a Multistep Assay | D. Tu, A. Holderby, J. Dean, S. Mabbott and G. L. Coté, Anal. Chem., 2021, 93, 4497-4505
*estimated from figures